## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2014-56

Agency: IOS, Ocean Sciences Division, Sidney, B.C.

Location:  Arctic (McKenzie Shelf, Amundsen Bank, Beaufort Sea)

Project: BREA

Party Chief: Dempsey M.

Platform: F/V Frosti

Date: 2 August 2014 – 9 September 2014

Processed by: Germaine Gatien

Date of Processing: 2 March 2015 – 13 May 2014

Number of original HEX files: 87 (1 aborted; 1 saved under 2 different names)

Number of CTD files: 59	Number of bottle casts:	85

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (Arctic #1189) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#SCST1666DR), 2 SBE 43 DO sensors (#1202 up to cast #42 and #2599 from cast #50 to the end), a SeaPoint Fluorometer (#SCF3651), a Biospherical PAR sensor (#70501), 2 surface PAR sensors (#20317 for first 2 casts and 20279 thereafter), a SeaPoint turbidity meter (#14014) and an altimeter (#41098).

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log book and rosette log sheets were available. No cruise report was available at the time of processing. There was a lot of confusion over event numbers and sample numbers and there was no single place where the most accurate information could be obtained. After much of the processing had been done more information was received that required repeating some steps.

There were frequently 2 casts at a single station because more than 12 bottles were needed. Separate event numbers were used. Spreadsheet 2014-56-bottles-final.xlsx has the data for each station arranged by sample number so it is easy to see the full bottle profile for each site.

For the CTD files only the first cast of the paired casts was fully processed. The second was always shallower and of poorer quality than the first.

The fluorometer performance was typical of these sensors with the fluorometer reading much higher than extracted chlorophyll at low CHL values. As CHL rises the ratio drops and the average is about 1 for CHL between 1 and 5ug/L. At high CHL (only a few values) the fluorometer ranges from 20% to 50% of the CHL.

There were some patches where the fluorometer was obviously off-scale in the full cast files, so values >4.9ug/L were replaced with pad values. Similarly, there were some sections in the full files where transmissivity appeared to have no signal for extended periods, so values <0.02% were replaced with pad values. The same steps were applied to the bottle files as similar features were seen in the upcast sections of the full files.

There was no dissolved oxygen sampling but post-cruise calibrations were available for the 2 sensors used. A comparison was made between files converted using pre-cruise and post-cruise calibrations. These indicate that there was a drift downwards of ~1% for sensor #1202 and ~5% for sensor #2599. Recalibration was applied by assuming all the change had occurred by the time of this cruise. For sensor #1202 there was some damage when the sensor reached the factory, so the results may not be reliable. For sensor #2599 there is no way to know when the drift occurred, but there was no evidence of drift through the cruise.

The dissolved oxygen sensor signal was very noisy during events #41 to #51 despite a change of sensor before event #50. The spikes are 2-sided and the average values look ok, so the data were kept in the files.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

Many files had names with an extra “.hex” in them, so program Bulk Rename Utility was used to change them to standard format.

There are 87 HEX files but 3 were later deleted since they were copies of files that had been renamed. One XMLCON file is missing but since the files acquired before and after use the same configuration, this should not be a problem.

There are 3 distinct XMLCON files – all the others are copies of 1 of those 3. The versions were saved as:

2014-56-ctd1.xmlcon – for casts #1 and #8

2014-56-ctd2.xmlcon – for casts #17 to #42 – change to Surface PAR sensor

2014-56-ctd3.xmlcon – for casts #50 to the end of the cruise – change of DO sensor.

The calibration parameters were all checked and no errors were found.

The rosette log sheets were obtained as well as analysis spreadsheets.

The cruise summary sheet was completed.

The histories of the pressure, temperature and conductivity sensors were found and they had only been used once between the most recent factory visit and this cruise. There was no history found for the DO sensor.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2014-56-ctd*.xmlcon to create CNV files. (*=1, 2 or 3)

A few casts were examined and all expected channels are present.

The temperature and conductivity channels track well on downcasts; on the upcasts there is more noise. The altimetry had a clear signal at the bottom.

The fluorescence and PAR look normal.

The SPAR traces have a lot of one-sided spikes.

The transmissivity has some odd excursions; for 2 casts there were sections between roughly 100 and 200m depth with low values. This was below the fluorescence maximum and dissolved oxygen was fairly low, so perhaps this is a real feature, but it is not seen in the upcast.

During cast #41-50 the DO signal is bad, full of spikes. The sensor was changed for cast #50 but that obviously did not fix the problem. However, by cast #71 the trace looks ok.

For cast #197 it was noted that downcast and upcast data are quite different.

The descent rates are quite steady at about 1m/s.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2014-56-ctd*.xmlcon. (*=1, 2 or 3)

The files were converted to IOS format.

They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files. A little noise was found in the primary salinity at about 30m during event #450. CTDEDIT was used to remove a few bad points. The resulting file was then copied to 2014-56-0450.BOT

A preliminary header check turned up no problems.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged including standard deviations. Some of the standard deviation channels are not useful such as those for the pump status or bottle number, so REMOVE was run to remove them. Output was SAMAVGREM

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2014-56-bot-hdr.txt; it may need further editing to reflect problems found during processing.

EXTRACTED CHLOROPHYLL

The CHL data were obtained in spreadsheet “BREA 2014-56 total CHL for Jane”.xlsx. Event numbers and sample numbers were missing so those were added and the file was saved as 2014-56chl.csv. There were often several bottles at a single depth and only depth is given in the spreadsheet, not a sample number. So it was assumed the sample came from the first bottle closed at the indicated depth. The file was sorted on sample numbers. There were no quality flags; a channel was created for flags. The file was sorted on event number and then sample number; it was then converted to individual CHL files.

SALINITY

Salinity analysis was obtained in Frosti 2014-56 salinity data v2015-01.xlsx. The analysis was done about 4 months after collection. The files were simplified and saved as 2014-56sal.csv. There were no event numbers so those were added. There were 6 samples with labels that do not match any sampling indicated on the rosette sheets. There were also 6 missing samples. Making rough estimates of the values expected for the 6 missing samples suggests a correspondence as follows:

| Missing Sample # | Samples with labels that do not correspond to log records of sampling, with rough matches based on salinity value |
| --- | --- |
| 465 | possibly sample labelled 454d |
| 466 | possibly sample labelled 455d |
| 467 | possibly sample labelled 456d |
| 468 | possibly sample labelled 457 |
| 469 | possibly sample labelled 458 |
| 470 | possibly sample labelled 459d |

The salinity analysis sheets were found later and they confirm that sample numbers 454-470 were used on two stations, so the correspondence shown above does appear to be correct. So the sample numbers were changed to enable merging with the other chemistry. File 2014-56-0349.sal was created with samples #465-470. The merge process was repeated for this cast. Note that this was done after the salinity comparison had been run, so these 6 samples are not included in the compare file.

There were many comments about bottle samples but no quality flag entries, though there were some pad values. It was decided to assign a flag 3 where the analyst noted leaking bottles or salt on the rims of bottles. These can be adjusted later if COMPARE suggests the problems are more serious than a 3 would suggest, or a 2 if there is clearly no effect. The salt was not thought to have affected the seal – probably got on the bottles after sealing but salt flakes might accidentally fall into the samples during analysis despite attempts to avoid that. Flag 5 was added where the analyst rejected the values.

The spreadsheet was then converted to individual SAL files.

The SAL and CHL files were merged with CST files in 2 steps.

After the 2nd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMREM files choosing the Bottle_Number from the SAMAVG files, to produce MRG files.

The MRG files were put through CLEAN to produce MRGCLN2 files; 0s were entered into any empty flag channels.

In the course of building the bottle files some problems were found. It is difficult when there is not a single source with event #, sample # and either Niskin bottle or depth information. The salinity analysis data came with sample numbers and station names, but no event numbers, depths or Niskin bottle #s. The CHL data came with depth and station names only, no sample numbers. There was a sampling spreadsheet for part of the cruise which did contain the information required.

- For files #475/476 at station DOL-05 a bottle was fired around 80m during both sections but only one sample number was assigned. It was assumed to belong to Niskin 1 of file #476. (Mike Dempsey later confirmed that was correct.)

- There were SAL and CHL sampling on a rosette sheet with event # given as 506/507 but there were no CTD raw data files for 507. While there was a plan to have 2 files (506 and 507) for station DOL-01, only 12 bottles were fired so only cast #506 was created. The SAL and CHL values labelled as form file #506.

- The casts at station FRK07 and SMO were originally planned as events 240 and 247 but the numbers were changed on the rosette sheets to 222 and 229. However, the CTD files were saved as 240 and 247. Since the casts were not run in the order originally planned, it was assumed that the file names should be changed to the numbers indicated in the sampling log. So they were renamed as 2014-56-0222 and 2014-56-0229, respectively.

There are 26 file pairs. These will not be joined as the Niskin numbers would repeat and might cause confusion. A spreadsheet will be prepared at the end which will display the data so the full profile is clear for the pairs.

##### Compare

Salinity

Compare was run with pressure as reference channel.

The plot of differences against pressure has two major outliers (samples 350 & 409) which both had standard deviations in the CTD salinity of ~0.004; they were from 48m and 20m. Typically, a standard deviation of 0.001 is used as a cut-off in such comparisons. .

- But for sample #350at 49m the error is in the opposite direction to what we would expect with the bottle value being much lower than the CTD salinity. Poor flushing won’t explain this and the local gradient was not high enough for the difference in depths between the CTD sensors and the bottle to explain difference of 2psu. The bottle value was 29.97 and all other values <30psu in this region are found above 30m. The value is out of line in profile and does not look like it was drawn form some other Niskin. Sample #350 was flagged “4” with the comment    “Sample_Number 473: Salinity: A major outlier in comparison of bottles and CTD salinity..”

- For Sample #549 from cast #409 at about 21m., the bottle is higher than the CTD by 2.5 salinity units. The vertical salinity gradient is high here, so it is possible that the Niskin just didn’t flush very well. Arguing against that, the CHL is in reasonable agreement with the fluorometer. It is possible the sample came from the 30m Niskin. Again there are no comments from the analyst about this sample. It was flagged 4 with the comment:    Sample_Number 473: Salinity: A major outlier in comparison of bottles and CTD salinity, but came from a depth with high salinity gradient.

All cases with differences >0.1 were examined. Only one of the cases with standard deviation in the CTD salinity <0.002 was from below 50m. Only 3 cases with large differences and low standard deviation had been flagged:

Sample 201 – 30m - salt on rim

Sample 202 – 20m - leaking

Sample 498 – 101m -no insert – flagged 5 by analyst – value replaced with pad value.

Most samples with the comment “salt on rim” did not stand out as outliers or had high standard deviation in the CTD salinity. So leaving 3 flags for samples with that comment looks appropriate. They are probably ok, but the evidence is too weak to change the flags to 2.

When bottles are excluded that have differences >0.1 and standard deviation in the CTD salinity >0.001 the primary salinity is low by an average of 0.018 and the secondary by 0.014 with standard deviations in the fits of 0.018 and 0.0189. If data from the top 50m are also excluded there is little difference. The variability is high. Excluding the two major outliers and then gradually removing data based on residuals produces a somewhat lower difference with the primary salinity low by 0.0144 and the secondary by 0.0119, with standard deviations of 0.012 for both.

Plots of differences against file pair numbers show no obvious drift with time and no difference between the samples run on different Autosals, but the variability in these data is too great to conclude much. The post-cruise calibration show that the CTD read a little low, but not as low as these results suggest. Thus there is likely either some evaporation of samples or poor flushing. The plots against pressure do suggest flushing is a problem as the differences are much higher in the top 50m where salinity gradients were high. So that is likely the main issue.

For full details for the COMPARE run see file 2014-56-sal-comp1.xls.

Extracted CHL vs Fluorescence sensor

COMPARE was run with pressure as the reference channel.

A plot of the ratio CTD Fluorescence / Extracted CHL versus CHL shows the usual pattern of the fluorometer reading much higher at low CHL values. As CHL rises the ratio drops and the average is about 1 for CHL between 1 and 5ug/L. At high CHL (only a few values) the fluorometer ranges from 20% to 50% of the CHL.

For full details see file 2014-56-fl-chl-comp1.xlsx

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on 2 casts for DO sensor #1202 and 3 for sensor #2599.

- For #1202 there are few useful casts to examine; most were shallow, some had bad DO signals and there are frequent stops for bottles so that upcasts look quite different from downcasts. But overall, an advance of +2s appears to produce better alignment with the temperature traces.

- For #2599 a setting of +3s looks best overall.

ALIGNCTD was run using a setting of +2s to the dissolved oxygen channel for casts #1-42 and +3s for casts #50-506.

##### CELLTM

Tests were run on a few casts to see which parameters worked best for the thermal mass correction. Results were judged by how well they made the upcast and downcast data look closer on a T-S plot. The high variability and frequent stops for bottles make it difficult to judge which were best except at a few “quiet” sections of the casts, but all parameters tested improved the correspondence. Overall, the best choice appears to be (α = 0.03, β=9) for both the primary and secondary sensors.

CELLTM was run using (α = 0.03, β = 9) for the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

The differences between sensor pairs were studied from a few casts, but most were too shallow and too noisy to tell us much about sensor performance.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-56-0061 | 250 | +0.0002 VN | +0.00010 | +0.0012 | Mostly High, Noisy |
|  | 490 | +0.00004 N | +0.00010 | +0.0010 | High, steady |
| 2014-56-0110 | 240 | +0.0005 VN | +0.00013 | +0.0015 | High, very steady |
| 2014-56-0469 | 240 | +0.0004 | +0.00012 | +0.0019 | High, Noisy |
|  | 380 | +0.0001 | +0.00014 | +0.0016 | High, Noisy |

The differences were often very noisy even though the descent rate was steady. The differences are small. There may be some drift towards higher salinity differences.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number. An initial examination showed that there were many cases of fluorescence being off-scale, so CLEAN will be rerun later to set to replace values of fluorescence >4.9 with pad values to make it clear that the maximum value found, 4.915ug/L, is not reliable. There are also large sections of very low transmissivity that appear to be due to malfunction so CLEAN may be deal with that too.

##### Oxygen calibration study

Since there was no DO sampling but there was a post-cruise calibration for both sensors used, a study was made of the differences between the data found using the pre-cruise and post-cruise parameters. All hex files were converted using the post-cruise calibration and the data were put through steps 6, 7, 8 and 9; the word POST was included in the file names. Those files were then converted to IOS header format. Those files and the CLN files were then put through REMOVE to remove unnecessary channels. The output files were named REM1 for pre-cruise and REM2 for post-cruise. The REM2 file names were adjusted to remove the word POST from the file names. Both REM1 and REM2 files were then put through steps DELETE, BIN-AVERAGE and THIN. COMPARE was run using files THN1 and THN2. Mistakes were made in a few of the post-cruise files so they were removed from the comparison. There were a lot of data and COMPARE ran very slowly.

The comparison for sensor #1202 shows a lot of variability which is likely due to the very noisy signal in some of the casts. However, despite that there was a fairly tight fit that showed that the pre-cruise calibration produced results that were ~99.0% of the post-cruise. The noise appears to be 2-sided and the fit looks convincing. So the correction:

DO corrected = 1.010 * DO using pre-cruise calibration

However, repairs were done when the sensor was returned to the factory and comments were made about poor response. The noise observed during the last few casts during which the sensor was used turned out to be due to something other than the sensor, so that does not show when the problem arose. There was no apparent drift through the cruise.

The comparison for sensor #2599 has a very tight fit with the pre-cruise files having values that are about 95% of the post-cruise values, so the correction would be:

DO corrected = 1.049 * DO using pre-cruise calibration

There was no damage noted in the factory report. There is no apparent drift through the cruise.

##### Checking Headers

A cross-reference list was checked against the log book. The only discrepancies found were the station names for casts 459 and 460 and 476. Those errors were fixed in the CNV, IOS and CLN files.

A header check was run. There are a few values that are odd, such as a negative pressure and conductivity values; these may be due to near-surface spikes. No other problems were noted.

The cruise track was plotted with event #s and station names and both were added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 1.5db. The last time this pressure sensor was used (2014-28) the offset was set to -0.69db. It was thought that this must have been due to a test during the cruise, but later it was apparent that the pressures were too low, so pressure was recalibrated by adding 0.9db. For this cruise there are many readings at about -0.4db that appear to be in the water with pumps on, but conductivity is low enough that it must have been near the surface. So adding 0.9db is a reasonable adjustment. This step will be taken now so that not too much data will be removed by DELETE.

File 2014-56-recal1.ccf was prepared to add 0.9db to the pressure channel. (Note: For some casts the wrong calibration control file was selected and 1.25db was added; a later correction was made by subtracting 0.35db to get a net +0.9db change.)

The water depth readings were included in most of the Sea-Bird file headers but the format was wrong and the information did not get added to the LOCATION section of the files converted to IOS Header format. A spreadsheet was prepared with the water depths and file names and that was used to merge the information with the COR1 files and the MRGCLN2 files with output extensions *.MRH and *.MRGMRH.

The altimeter readings from the headers of the MRH and MRGMRH files were exported to spreadsheets. Plots were made of altimetry for a selection of casts and they showed the readings in the MRH files to be appropriate. The entries in the bottle files are always a little lower than in the MRH files. Perhaps the altimeter changed its angle so it was reading straight down after a little while at the bottom.

##### Shift

Fluorescence

SHIFT was run on the fluorescence channel in all casts using the usual advance of +24 records. Plots made after this step were hard to interpret due to noisy fluorescence and many stops for bottles and many cases of off-scale fluorescence. There is some improvement in matching the fluorescence offset to the temperature offset, but it not impressive.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel, but the alignment looked good.

Conductivity

Tests were run on 5 casts using a variety of shifts and the best results were found with the settings -0.3s for the primary and -0.4s for the secondary. SHIFT was run twice using those settings.

##### DELETE

Before running DELETE plots were made to look for casts where acquisition started before the soak period. None of those were found, but there were 2 files (#75 and 85) which had been accidentally lowered with the pumps off and then brought back to the surface to run a complete cast with pumps on. A text editor was used to remove the data from the first drop so that DELETE will work properly. Notes were made about this in the relevant headers.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings..

##### Other Comparisons

Previous experience with these sensors –

Salinity: The temperature and conductivity sensors were new in April or early May. They were used for 2014-28 but there was no salinity calibration sampling.

Dissolved Oxygen: The sensors were both calibrated at the factory in October 2013 and this was likely their first use since then.

Pressure: The sensor was new in December 2013 It was used during 2014-28 before this cruise. An offset had been added so that was used in processing, but it was later found to be inappropriate so that amount was subtracted from the pressure later.

Historic ranges – Local climatology was not available.

Repeat Casts – There were many repeat casts as two were needed in order to fire the required bottles, but one of the two was always shallow where variability is too high to enable a good comparison.

Post-Cruise Calibration – There were post-cruise calibrations in mid-December 2014 for the temperature and conductivity channels. The factory reported drift since the previous calibration of -0.0009C°/year and -0.0008C°/year for the primary and secondary temperature channels, and -0.0011psu/month and -0.0007psu/month for the primary and secondary conductivity channels. This implies that in December the primary and secondary salinity were low by about 0.008 and 0.005. The T and C sensors were used on a previous cruise so some of the drift likely occurred before this cruise and there may have been some drift on the shelf and in transport to the factory after this cruise.

There were also post-cruise calibrations from the 2 oxygen sensors but no drift estimate was included in the reports, but a comparison of results using pre-cruise and post-cruise parameters for sensor #2599 is described in section 11.

##### DETAILED EDITING

The secondary temperature and salinity channels were edited since they looked a little smoother in T-S space than the primary channels.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and records corrupted by shed wakes including some records from near the top and bottom of the casts.

All files were edited lightly.

All EDU files were copied to EDT.

##### Initial Recalibration

The pressure was recalibrated earlier, but it was discovered that for most files the wrong calibration control file had been selected. Going back to fix that would require re-editing the data which is impractical, so a second calibration will be applied to correct the first for casts 1-40 and 51-506 by subtracting 0.35db.

The comparison of CTD salinity with bottles was sensitive to the choice of outliers, but suggests that both sensors were lower than bottles. The primary was likely lower by about 0.0144 and the secondary lower by 0.0110, but incomplete flushing of bottles is likely making the CTD salinity look lower than it really is. The post-cruise calibration shows that the CTD primary and secondary sensors were reading low by about 0.008 and 0.005, so the differences between the two sensors match those found between them in the bottle comparison. There was one cruise between the pre-cruise calibration and this one. Assuming there was a little drift through this cruise, we might make an estimate that the primary was low by ~0.0060 and the secondary low by ~0.0035 mid-way through the cruise. The differences found in section 9 are smaller than 0.0025 but that is the average difference found in the bottle comparison; however, it may be that alignment of sensors affect that study since the CTD was in motion.

There was no dissolved oxygen sampling. The comparison described in section 11 provides a means of recalibrating the CTD DO assuming no significant drift after this cruise. Since there is no obvious sign of drift through the cruise this is the best we can do.

CALIBRATE was run using file 2014-56-recal2.ccf to apply the following corrections in the MRGCLN2 files:

CTD DO corrected = 1.010 * CTD DO (casts 1-42)

CTD DO corrected = 1.049 * CTD DO (casts 50-506)

Pressure corrected = Pressure – 0.35db (casts 1-40 & 51-506)

CTD Salinity:T0:C0 corrected = 1* Salinity:T0:C0 +0.0060 (all casts)

CTD Salinity:T1:C1 corrected = 1* Salinity:T1:C1 +0.0035 (all casts)

The pressure calibration had not been done for the bottle files at this point so file 2014-56-recal2-che.ccf was prepared to correct only DO and Salinity with no adjustment to the pressure.

CALIBRATE was run on the SAM files in 2 steps using 2014-56-recal1.ccf and 2014-56-recal2-che.ccf. COMPARE was rerun to ensure the changes were made properly to salinity and they were. A few casts were compared for DO values before and after the correction and the changes were found to be correct.

CALIBRATE was then run on the EDT files using 2014-56-recal2.ccf.

At this stage it was discovered that an error had been made in the downcast files 41 and 51 so that all fluorescence data were replaced with pad values, so the steps between conversion and this stage were repeated for those casts.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### Plot Checks

Plots were made of all variables to check for off-scale values or signal drop-out. Problems were found with fluorescence appearing to be off-scale. The maximum values found were 4.91, a run of CLEAN was used to replace all transmissivity values >4.9 with pad values so that users would not assume the values are correct.

There was at least section of 0 transmissivity that looks like a loss of signal rather than real values, so a second run of CLEAN was used to replace values <0.02% with pad values.

PAR and Surface PAR have more noise than usual and since it is not always clear whether there is no signal or no light, no attempt was made to edit those channels.

T-S plots turned up a few surface salinity points from before the pumps were turned on, so those were replaced with pad values.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some unstable features, but those are expected in this region.

Two sets of files were prepared from this point onwards, one with standard deviations for the use of the Arctic group and the other without them for the OSD Data Archive.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

For casts #1 and 3 channel PAR:Reference was also removed as it had no signal in those casts.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

CLEAN was run twice to replace transmissivity values <0.2 and fluorescence values >4.9ug/L with pad values.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

---------------------------------------------------------------------------------

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, Turbidity, PAR and Surface PAR data are nominal

and unedited, except that some records were removed in editing

temperature and salinity and some off-scale fluorescence values and

transmissivity with no signal were replaced with pad values.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was recalibrated based on a post-cruise calibration. The

DO signal was bad for casts #41-51 despite a change of sensor before

event #50. The spikes are two-sided and average values look ok, so the

data were kept.

Salinity data were recalibrated based on post-cruise calibrations plus a study

of differences between salinity channels. It is likely that not all the

drift had occurred at the time of this cruise, so an estimate was used.

For details on the processing see the report: 2014-56_Processing_Report.doc.

---------------------------------------------------------------------------------

The Standards Check routine was run and no problems were found.

The Header Check was run and no further problems were found.

Profile plots were made and look ok.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Most values were close to 100% with only 2 lower than 95% and 9 slightly higher than 105%.

##### Final Bottle Files

The MRGCLN2 files were put through the MERGE HEADER routine to add water depths. Those files were then put through CALIBRATE twice, first to fix pressure using 2014-56-recal1.ccf and next to correct the DO and salinity using 2014-56-recal2-che.ccf.

The MRGCOR2 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

For events #1 and #8 PAR:Reference was also removed.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

For the special files for Jane Eert, Depth was derived.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check was run on all files and no errors were found.

A header check was run on the final files.

No further errors were found.

The track plot looks ok.

Plots of each file were examined and no further problems were found.

Data from the CHEA files were exported to a spreadsheet and ordered on sample numbers so that it is easy to see the full profile for each site. The file was saved as 2014-56-bottles-final.xlsx.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# NOTE: In the final stages of processing some changes were made to salinity flags and in a few cases pad values were changed to original values. In case there is a need to repeat some processing, be aware that the late corrections were made to the MRG1, MRG2 and MRG files. Files between MRG and CHE will have some errors. The samples affected are #250, 350, 360, 351, 473, 549 and 695.

# Particulars

1 & 8. No SPAR signal.

17. SPAR on.

31. Same data as cast #32 – cast #31 deleted.

32. CTD hit bottom.

41. Partway down noise starts in DO signal and continues during 42.

42. DO noise continues

50. Change of DO sensor but noise continues.

51. Connector cleaned and DO improves slightly.

61. New DO/ALT cable installed and DO noise disappears.

75. Return to surface from 35m then full cast; initial downcast removed.

85. Fluorometer noisy with offset on upcast – clean connectors.

100. File has only partial upcast in it. The full cast was named as cast #102. Deleted #100.

149. Station name should be NTC01

150. Station name should be NTC01

166. DO dropped out. Connectors cleaned. No problem with next cast.

204. Fluorometer and Transmissometer failed – noisy, then dead.

206. Cable replaced – FL/TR ok.

230. Renamed 212

231. Renamed 213

240/222 – rosette log indicates file was saved with wrong name. Converted files renamed as 222.

247/229 – rosette log indicates file was saved with wrong name. Converted files renamed as 229.

264. Should be 265 – CTD cast only – no bottles

322. touched bottom.

349. Salinity sample numbers from cast #341 were apparently reused on the sample bottles of cast #349, but the rosette sheet shows different #s. There is a chl sample with a sample # that matches rosette log. Salinity sample #s were changed to match rosette log.

498. Should be 499

499. Should be 500

475. Station name should be DOL_05. Sample 642 was not used for this cast. Was used for 476. So line for that sample was removed from bottle file.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 1189 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5830 | 1May2014 | Factory | 18Dec. 2014 | Factory |
| Conductivity | 4327 | 10April2014 | Factory | 17Dec. 2014 | Factory |
| Secondary Temp. | 5831 | 29April2014 | Factory | 18Dec. 2014 | Factory |
| Secondary Cond. | 4339 | 9April2014 | Factory | 23Dec. 2014 | Factory |
| Transmissometer | CST1666DR | 30June2014 | IOS |  |  |
| SBE 43 DO sensor | 1202 | 23Oct2013 | Factory | 19Dec2014 | Factory |
| SBE 43 DO sensor | 2599 | 23Oct2013 | Factory | 23Dec2014 | Factory |
| PAR | 70501 | 8Aug2013 | IOS |  |  |
| Surface PAR | 20317 | 24Mar2008 |  |  |  |
| Surface PAR | 20279 | 13Mar2007 |  |  |  |
| pH | 0691 | 24Dec2010 | Factory |  |  |
| Seapoint Fluor. | 3651 | June2014 |  |  |  |
| Turbidity Meter | 14014 | 25Mar2013 | Factory |  |  |
| Pressure Sensor | 1189 | 14May2014 | IOS |  |  |
| Altimeter | 41098 | n/a |  |  |  |