## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27 Mar 2015 | 1. Post-cruise calibration reports showed that the recalibration applied to temperature and recalculation of salinity were inappropriate. That step was reversed. For details see note at end of this report after §24. 2. Header notes about salinity bottle inserts were corrected in CHE files. |
| 24 Feb 2015 | Data for this cruise were collected at the end of 2014-19 and beginning of 2014-22; TSG files 2014-19-0005.tob and 2014-22-0001.tob include data for 2014-60. The data can be found in the “Underway” folders for cruises 2014-19 and 2014-22 in the OSD data archive. |

## PROCESSING NOTES

Cruise: 2014-60

Agency: OSD

Location: WCVI

Project: Strait of Georgia/Juan de Fuca Water Properties Survey.

Party Chief: Yelland D.

Platform: John P. Tully

Date: 5 September 2014 – 15 September 2014

Processed by: Germaine Gatien

Date of Processing: 18 December 2014 – 27 January 2015

Number of original HEX files:	70	Number of CTD files: 70

Number of bottle files: 24	Number of bottle casts processed: 22 (2 files had no sampling)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (# 0443) was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), an SBE 43 DO sensor (#0997) on the secondary pump, a SeaPoint Fluorometer (#3642) with a 3X cable on the primary pump, a Biospherical QSP-200 PAR sensor (#4615) and an altimeter (#1204).

The data logging computer was #3. Seasave version V7.22.4 was used for acquisition.

The deck unit was a Seabird model 11, serial number 0471.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book and rosette sampling log sheets were in good order.

Due to a lack of ship time, this cruise was fit into the programs of 2014-19 and 2014-22, but the data were saved with a separate cruise number. Most casts were run between September 2 and September 6, but the last 2 were run on September 14 and 15.

During 2014-19 there was a large difference between the two CTD temperature channels, and this continued through this cruise. An initial examination of data from the cruise which followed this one also shows an unusually large difference. Calibration drift in the primary temperature was believed to be the main cause of differences between the derived salinity channels for 2014-19. The recalibration method applied to 2014-19 was repeated for this cruise as well, namely, primary temperature was recalibrated and primary salinity recalculated. NOTE: It was later found that the problem was with the secondary temperature, so the later processing steps were rerun. See notes in after §24.

The fit of the Oxygen:Dissolved:SBE data against titrated dissolved oxygen samples was unusual, but similar to those found for the latter part of the cruise that preceded it and all of the cruise that followed it. The same DO sensor was used for all 3 cruises.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±1mL/L from 0db to 20db

±0.25mL/L from 20db to 125db

±0.10mL/L from 125db to 200db

±0.06mL/L below 200db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets as well as various analysis logs were obtained.

Notes were added to the particulars section at the end of the report based on log book notes.

Extracted chlorophyll, nutrients, dissolved oxygen and salinity data were obtained in spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were obtained. They had been used for 1 other cruise since the last factory calibration.

The calibration constants were checked for all instruments. A few small errors in spelling and the date of the pressure sensor calibration were fixed in the configuration file, which was then saved as 2014-60-ctd.xmlcon.

##### BOTTLE FILE PREPARATION

The ROS files were converted using files 2014-60-ctd.xmlcon with Tau corrections selected but not the hysteresis correction. Those files were put through CLEAN to add event numbers (*.BOT).

Temperature and salinity were plotted for all BOT files. The only bad points found were in cast #7 Niskin #1. These were removed using CTDEDIT. The resulting file was copied to *.BOT.

Header Check was run on the BOT files and no problems were noted.

The BOT files were then averaged to enable an ADDSAMP file to be prepared. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

Sort was used to ensure that the ADDSAMP file was in sample number order.

There were 24 bottles fired at the surface of cast #48 but there is no rosette sheet and no log entry, so this was assumed to be accidental, so a rosette file will not be prepared for this cast.

SAM files were created using the Add Sample Number routine and those files were then bin-averaged.

The water depth headers was fixed in one SAM file based on study described in section 12.

Bin-average was then run using bottle numbers for bins to produce SAMAVG files.

The addsamp.csv file was ordered on sample number and converted to CST files. The CST files will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analyst wanted included in the header file. These were used to create file 2014-60-bot-hdr.txt.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2014-60CHL.xls. A simplified version of the spreadsheet was saved as 2014-60chl.csv which was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2014-60oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified by removing a few unnecessary columns and the file was then saved as 2014-60oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was provided in spreadsheet QF2014-60SAL.xls. The file was simplified and saved as 2014-60sal.csv. That file was converted to individual SAL files. The salinity data were analyzed on 3 days, 10 September, 22 October and 23 October, 4 to 37 days after collection. As well as possible differences due to increased evaporation for those samples that waited 5 weeks, there was also a problem with the salinometer on October 23rd that may affect results from the last two casts.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2014-60nuts.xls which included a report on precision.  The loop samples were copied to file 2014-60-loop.csv. The file was simplified, reordered on sample numbers and the file was saved as 2014-60nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The MRG files were exported to a spreadsheet and compared to the rosette sheets. No inconsistencies were found.

An initial header check did not turn up any problems.

##### COMPARE

Dissolved Oxygen

COMPARE was run using DO from the MRG and SAM files.

Differences were plotted against CTD DO for all casts and outliers were removed first by large differences and then using residuals to get a reasonably tight fit which was:

DOX_BOT = 1.0329 * DOX_CTD -0.1963

This is very close to the fit for 2014-22:

DOX_BOT = 1.0322 * DOX_CTD -0.1949

And the slope is similar to that found for 2nd half of 2014-19 when the fit changed from a normal looking slope/offset to:

DOX_BOT = 1.0326 * DOX_CTD -0.1050

The offset was larger than usual for 2014-19 and got even larger for 2014-60 and 2014-22. There are two things that distinguish this cruise from 2014-19 and might be relevant to the fit. First, the descent rate of the CTD was remarkably steady, so flushing of Niskin bottles is likely to have been less efficient than usual. Unfortunately, we don’t have a lot of salinity bottles, so we can’t assess flushing by using them. If flushing were poor, then the CTD would look higher than usual relative to bottles. Second, there were many casts with small dissolved oxygen ranges; this would limit the effect of poor flushing, but would also lead to better response of the sensor.

There were 2 other cruises during 2014 that had large offsets, 2013-15 (WC/DC) and 2013-06 (SoG/JdF). The DO sensor was different for those cruises and in subsequent use, fits with normal offsets were found.  Both were cases where flushing was likely to be incomplete, so the large offset might be due to poor flushing in those cases.

With that in mind, it is possible that the changes in offset between the latter part of 2014-19 and 2014-60 may be due to poor flushing. So we may have 2 effects – a DO calibration shift plus poor flushing.

For 2014-19 there was no evidence of analysis or sampling problems.

While there are a few values that look out of line none look like they are due to analysis or sampling problems. Most are associated with high standard deviation in the CTD DO channel. A few outliers that were well below the surface were investigated because the standard deviation was relatively low. None of these outliers had been flagged by the analyst. Plots show that for each of them there was a lot of variability during the bottle stop despite very steady pressure. For 1 bottle stop at 75db the temperature changed by 0.4C° and dissolved oxygen by 0.2mL/L though there was little change during the 10s window used in building the bottle files. No quality flags are justified.

For more details see 2014-60-dox-comp1.xls.

Salinity

Compare was run with pressure as reference channel. Salinity samples were analyzed in 3 groups, casts 1-40 were run within 8 days of collection so we expect no problem with evaporation. Samples from casts 46-73 were run about 37 days after collection. Samples from cast #74 and 75 were run 1 day later but on a day when problems were noted with the salinometer and with other samples run from another cruise.

An initial comparison showed a lot of scatter in the fit. A fit against file pair number makes it clear that there was a problem with the bottles from the final day of analysis – the values for casts #74 and 75 are all low despite the fairly long wait which might lead to evaporation of samples. This fits the results found for a group of 2014-22 samples that were analyzed on the same day. A major problem was found with the salinometer soon after these analyses. When those samples are excluded from the comparison, there is still a lot of scatter. Three factors may account for this scatter: sampling in areas of rapidly varying salinity, poor flushing of Niskin bottles in areas of large vertical gradients and evaporation of samples. Variability is a factor in these data as evidenced by high standard deviation in the CTD salinity during many of the bottle stops; this could lead to the CTD looking either higher or lower than bottles. Poor flushing usually leads to salinity in bottles being too high but could occasionally do the opposite in a complex profile. Evaporation would lead to bottles being high in a random way but the effect is expected to be small since all samples were analyzed within 5 ½ weeks of collection.

For this program the best comparisons are usually from the deeper parts of the central Strait of Georgia. There are 3 such samples from this cruise that were analyzed quickly. The primary salinity was found to be low by 0.0015, 0.0047 and 0.0045, while the secondary was high by 0.0008, low by 0.0032 and low by 0.0039. The first of the 3 casts had a low vertical gradient just below the level of the bottle firing, so is probably more reliable than the other 2.

When we compare the differences for the samples analyzed quickly with those that were run 5 weeks after the first set, the scatter was higher with the delayed analysis; that is something we would expect from evaporation as it will not affect all bottles equally. However, the differences themselves are larger than expected from a little evaporation and there is so much variation in salinity gradients at the sampling levels, that it is impossible to justify general conclusions from so few data.

In summary, if we restrict the data included to those with very low standard deviation in the CTD salinity (including data from both quick and delayed analysis) the primary salinity is found to be low by an average of 0.0032 and the secondary are low by 0.0010. Removing 1 outlier based on differences between CTD and bottles gives an average of -0.0037 and -0.0016. If we use only deep bottles with quick analysis the primary is found to be low by 0.002 and the secondary very close to bottles.

Results from the 2 cruises that bracket this one are clearer and more consistent with each other once we allow for the quick analysis for the first and delayed analysis for 2014-22. For 2014-22 when shallow bottles were excluded as well as some data that were analyzed when we know there were problems with the salinometer, both CTD channels look lower than during 2014-19 by about 0.0025. There may be a small amount of drift of primary relative to secondary salinity, but most of the change looks like it is due to either evaporation of samples or poor flushing of Niskin bottles. The latter is not so likely in deep waters offshore, but the CTD descent rate was very steady so it may account for some of the difference.

| Cruise | Sal 0 – Sal Bottle | Sal1 – Sal Bottle |
| --- | --- | --- |
| 2014-19 | -0.0033 | -0.0001 |
| 2014-22 | -0.0057 | -0.0028 |

The 2014-19 data are likely the more reliable for salinity calibration.

A separate run of COMPARE was used to check the differences between temperature channels in the bottle files. The average difference was 0.0034C° with the primary reading higher. This is similar to the result found during 2014-19 when it was 0.0032C° below 1500db. (See 2014-60-temperature-comop1.xlsx.)

When outliers are removed based on high standard deviation in the CTD salinity during the 10s window around firing time, as well as the cases where bottle analyses were run on a malfunctioning salinometer, a lot of scatter remains. The two most significant outliers that remain were investigated:

- Cast #12 – 300db – The salinity is increasing with depth near the bottom, so poor flushing would lead to the CTD reading higher than the bottle which it does in this case.

- Cast #16 – 175db – A large shed wake passed through this depth at the beginning of the stop with salinity high enough to explain the difference. The local gradient is very complex during the downcast. So, anything looks possible!

There is insufficient evidence to justify changing any quality flags except for events #74 and 75 when the salinometer was malfunctioning. Those flags were changed to 4 for cast #74 since it not clear the value is bad and 5 and pad values for cast #75 when it is clear.

For full details for the COMPARE run see file 2014-60-sal-comp1.xls.

Fluorescence

COMPARE was run to see how the SeaPoint fluorescence channel compared with extracted chlorophyll.

The SeaPoint Fluorescence is close to or higher than CHL at low values of extracted CHL. The values fall off to lower ratios as CHL gets higher, though there appear to be 2 groups with higher and lower ratios at fairly low CHL values. The higher ratio of FL/CHL occurs in the northern and central Strait of Georgia with lower ratios in the southern SoG and Juan de Fuca Strait. This may be due to higher gradients to the north so that poor flushing of Niskin bottles is more significant; the CTD fluorescence is being compared with samples from deeper water with lower CHL. In Juan de Fuca the waters are better mixed. When all data are included the SeaPoint Fluorescence is ~57% of Extracted CHL.

For more detail see file 2014-60-fl-chl-comp.xlsx.

##### Conversion of Full Files from Raw Data

All files were converted using con files 2014-60-ctd.xmlcon with hysteresis correction turned off. The Tau correction will be applied.

A few casts were examined and all expected channels are present.

The two temperature and channels are sometimes quite close but at other times far apart. In general, the upcasts have larger differences, but in low-gradient zones they are similar. This is presumably due to minor differences in sensor plumbing or alignment that decreases differences on the way down and increases them on the way up in higher gradient zones.

The conductivity traces are close during downcasts and further apart during upcasts. The differences between downcast and upcast are larger for the primary sensor but it has smoother traces.

The transmissivity looks fine. Fluorescence values are quite high and traces look fine.

PAR and dissolved oxygen profiles look normal.

The altimetry looks useful though there are sometimes spikes at the bottom, so the header values will need to be checked.

Station 42 was occupied twice about 11 days apart and the traces are very different with higher temperatures and lower DO below 100m for the 2nd cast; the largest changes are around 150m.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

High variability in temperature makes upcasts and downcasts hard to compare. It is thus difficult to determine the best alignment setting to use, but the choice of +2.5s that was used for 2014-19 produces reasonable results.

ALIGNCTD was used to advance the DO Voltage by 2.5s relative to the pressure for all casts.

##### CELLTM

The usual test for CELLTM settings is to check which ones produce the best correspondence between downcast and upcast traces on a T-S plot. For CTD #0443, the values used for 2014-19 were tested and produce as good results as are likely to be achieved given the noisy temperature traces.

CELLTM was run using (α = 0.03, β=9) for the primary and secondary conductivity for all casts.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences were often very noisy and the casts shallow. For comparison, readings from one of the late casts from 2014-19 are included.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-60-0007 | 290 | -0.0042 | -0.00025 | 0.0010 | High, V. Steady |
| 2014-60-0015 | 320 | -0.0047 | -0.00026 | 0.0014 | “ |
| 2014-60-0024 | 300 | -0.0045 | -0.00028 | 0.0010 | “ |
| “ | 340 | -0.0044 | -0.00022 | 0.0015 | “ |
| “ | 400 | -0.0043 | -0.00022 | 0.0017 | “ |
| 2014-60-0031 | 300 | -0.0047 | -0.00033 | 0.0005 | “ |
|  | 350 | -0.0045 | -0.00029 | 0.0010 | “ |
| 2014-60-0074 | 300 | -0.0039 | -0.00027 | 0.0008 | “ |
|  | 340 | -0.0042 | -0.00029 | 0.0010 | “ |
| 2014-19-0109 | 300 | -0.0037 | -0.00044 | -0.0013 | High, X Noisy |
|  | 400 | -0.0038 | -0.00037 | -0.0005 | “ |
|  | 600 | -0.0038 VN | -0.00023 VN | 0.0007 | “ |
|  | 1000 | -0.0033 XN | -0.0002 XN | 0.0016 | “ |
|  | 1980 | -0.0034 XN | -0.0001 XN | 0.0021 | “ |
|  | 2980 | -0.0035 | ~0 | 0.0033 | “ |
|  | 3900 | -0.0037 | ~0 | 0.0035 | “ |

Despite a very steady CTD descent rate there is a lot of variability in the traces.

The temperature differences are fairly steady and slightly higher than those seen at similar depths during 2014-19.

The conductivity differences are small, and are a little lower than for the earlier cruise.

Salinity differences always vary with pressure (perhaps due to the decreasing effect of minor misalignments as vertical gradients decrease, or differential pressure effects on the flow rates of the 2 pumps) and for the 2014-19 cast you can see that the differences change sign around 500m. For 2014-60 the differences show the same trend of Sal1 increasing relative to Sal0 as pressure increases but (S1 – S0) is positive at 300m for the current cruise while it was negative for cast 2014-19-0109. This could imply a drift in calibration of one or both sensors, or perhaps just a difference in plumbing performance or alignment. Conditions were definitely very different for the 2 cruises.

In any case the temperature differences have not gone away, so we are still faced with the fact that either one sensor has drifted significantly, or both could be off. For 2014-19 the excellent correspondence between the secondary salinity and the bottles in deep water suggested that the problem was with the primary temperature. Unfortunately, for that cruise the secondary channels were very noisy and hard to edit, so it was decided to recalibrate the temperature and recalculate salinity. The best temperature correction was found by testing so as to make the primary temperature move close to the secondary and make the primary salinity channel move close to bottles. The ideal correction for one was not idea for the other, so a compromise correction was found.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

The header check was run.  The minimum pressure was negative so some casts were checked and it was found that there were many records with in-water values and pumps on with negative pressure. At the end of casts, the CTD data suggests that the sensors are out of the water when pressure is about -0.5db.

The surface check gave an average value of 1.6db which is a little low.

CALIBRATE was run using file 2014-60-recal1.ccf to add 0.6db to all casts.

The cross-reference check was compared with the log book. There were 2 discrepancies:

- The log times differ by 1 hour for event #6. The time given in the log does not allow enough time to travel between events 5 and 6. Adding 1 hour to those times is an approximate match to the header.

- The station names for events 13 and 14 were identified as 9 and 8 in the log, but as 8 and 9 in the headers. Checking other cruises for positions for each of these sites showed that the header position was consistent with the station name, but the log station names are incorrect.

The altimeter, maximum pressure and water depth were exported from the headers of the CLN and MRG files to spreadsheets.

The depths were checked against the log book; the difference between maximum pressure (changed to depth) + altimeter reading at the bottom were also compared to the bottom depth entry. There were 6 cases where the difference between log and headers was significant and the log entries were more believable; in one case it was clearly a typo in the header. The log values were entered in the headers for those cases.

Plots were made to check altimetry entries and all look appropriate when the CTD got within 20m of the bottom. No problems were found; even though there were frequent spikes at the bottom the algorithm worked well.

The cruise track was plotted and added to the end of this report. No problems were found.

##### 13. Shift

Fluorescence

Normally a shift of +24 records is used to align SeaPoint fluorescence with temperature. Tests were run

to see if that setting improves the alignment and it does.

SHIFT was run on all casts to advance the SeaPoint fluorescence by +24 records.

Conductivity

Tests were run on a few casts using a variety of settings to see if the alignment settings used for 2014-19 suit these data. The effect of applying slightly larger or smaller shifts was examined by checking the stability of features in T-S plots. The 2014-19 shifts look appropriate for these data.

SHIFT was run twice on all casts applying a shift of -0.7 records to the primary conductivity and +0.7records to the secondary conductivity.

SALINITY was recalculated after each of the conductivity shifts.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel. The alignment looks as good as we are likely to achieve.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –These sensors were used during 2014-19.

1. PRESSURE The pressure was not recalibrated. Tests suggested it was within 0.2db.

2. TEMPERATURE/SALINITY There was a much larger difference between the temperature channels than usual, ~0.0035C°. Because the secondary salinity was very close to bottles it was presumed that the primary temperature needed recalibration. It was recalibrated by subtracting 0.003C° which brought the primary salinity into good correspondence with the bottle salinity.

3. DISSOLVED OXYGEN The calibration of the DO sensor changed abruptly during the cruise, so 2 different corrections were applied:

DOX_BOT = 1.0245 * DOX_CTD -0.0036

DOX_BOT = 1.0326 * DOX_CTD -0.1050

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. These are not really appropriate for casts close to shore. The temperatures were all within the climatology except for some high values in the top 20m in the northern part of the Strait of Georgia. Salinity values were low at depth for some of the northern casts. In the middle section near-surface salinity was low, but near the bottom it was high. In the southern part of the Strait of Georgia high salinity values were found below 100m. In Haro and Juan de Fuca Strait there were no significant excursions from the climatology. Given that the salinity excursions are not systematic and it is a region noted for rapid change, the results are not considered indicative of a problem with calibration.

Repeat Casts – There were repeat casts but the casts were shallow and in an area known for rapid change.

Post-Cruise Calibration - There were no post-cruise calibrations available at the time of processing.

##### DETAILED EDITING

The first issue is to decide which sensor pair to edit. As noted during 2014-19 it appears that the calibration of the secondary temperature and conductivity are likely better than for the primary. However, there is a lot of noise in the secondary data. From this cruise it is quite clear that when the vertical temperature gradient is high the secondary sensor is delayed in responding. This leads to salinity that is extremely noisy. So, as for 2014-19, it looks best to edit the primary channels.

CTDEDIT was used to clean salinity, and to remove some bad records near the top and bottom of many casts and some records corrupted by shed wakes though this was not a big problem due to steady descent rates except at the top and bottom of casts. All casts required some editing. For cast #69 there was a stop at mid-depth of the downcast to improve the wire angle. This led to an unstable profile presumably because the CTD has moved to an area of similar T/S characteristics but at different depths. No editing was applied.

T-S plots were examined and while there are a few small unstable features, in this region they look believable. No further editing was applied.

The edited files were copied to *.EDT.

##### Initial Recalibration

The pressure correction has already been applied to the full files, but not to the bottle files.

The MRGCLN2 and SAM files were recalibrated using file 2014-60-recal1.ccf to add 0.6db to Pressure.

Then file 2014-60-recal2.ccf was prepared to correct dissolved oxygen using

DOX_Corrected = 1.0329 * DOX_CTD -0.1963

and to subtract 0.003 to Temperature:Primary followed by a recalculation of Salinity:Primary.

NOTE: The corrections to the temperature and salinity were later reversed based on post-cruise reports. Steps from here to the end were rerun. See end of report (after §24) for details.

COMPARE was rerun to see if the corrections were appropriate. The decrease in temperature brought the two salinity channels into good correspondence, within 0.001 of each other. The temperatures are within 0.0011C°. The dissolved oxygen fit is quite flat with a minimal offset.

(See files 2014-60-dox-comp2.xls, 2014-60-sal-comp2.xlsx and 2014-60-temp-comp1.xlsx.)

The second recalibration was applied to the EDT files.

##### Final Calibration of DO

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles.

When the differences were plotted against DO concentration, the fit had a lot of scatter but looked reasonable. When a few outliers were removed the SBE DO was high by an average of 0.05mL/L but the standard deviation was 0.18mL/L. The SBE values were higher than bottles near the surface and a little low for bottles fired at the bottom of the cast; both these differences could be the result of incomplete flushing of Niskin bottles.

Given the high variability, possible issues with flushing and the large scatter in the comparisons, no further recalibration is justified.

A plot of differences between bottles and CTD DO against pressure was used to make a final study of outliers and to make a rough estimate of the accuracy of downcast CTD dissolved oxygen.

##### Special Fluorometer Processing

A median filter, fixed size=11, was applied to reduce spikiness in the fluorescence channel. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was considered appropriate.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove channels: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

Profile and T-S plots were produced at this point to check for errors. No problems were found.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, correct the header entry for chief scientist and platform and add “Mid-ship” to the instrument location section and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence data are nominal and unedited except that some

records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that an offset

in the fit was allowed.

Post-cruise calibration reports indicate that the primary temperature and

conductivity sensors had insignificant drifts.

The Oxygen:Dissolved:SBE data had an unusual fit against titrated samples, but

it was similar to that found during 2014-19 and 2014-22 during which the

same equipment was used.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±1mL/L from 0db to 20db

±0.25mL/L from 20db to 125db

±0.10mL/L from 125db to 200db

±0.06mL/L below 200db

For details on the processing see the report: 2014-20_Processing_Report.doc.

For details on the processing see the report: 2014-20_Processing_Report.doc.

The final files were named CTD.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

Profile plots were made and look ok.

The track plot looks ok.

The sensor history files were updated.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values were low in the southern Strait of Georgia, Haro Strait and most of Juan de Fuca Strait (70%-90%). This is normal due to vertical mixing. In the centre of the Strait of Georgia values were in the 90%-110% range. The northern part of the Strait of Georgia had high saturations, between 110% and 140% (with the exception of one cast which was well mixed) and there were a few high values in amongst the low ones in the southern part of the Strait. The values are not out of line for this region.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, fix a few headers, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data.

A header check was run and a problem was found with the draw temperature. In one case there was a value of -44 – this was changed to -99. Header check was rerun and no further problems were found.

For a final check the CHE bottle data were exported to a spreadsheet and compared with the rosette log sheets. No problems were found.

Plots were made of CTD Salinity versus SBE Dissolved Oxygen and bottle DO and no further outliers were identified.

Standards check was run on all files and no problems were found.

A cross-reference list turned up no errors.

The track plot was produced on screen and no errors were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

CORRECTIONS – March 27, 2015: A post-cruise report on the drift in conductivity and temperature sensors showed that it was the secondary temperature sensor that was out of line, not the primary. So the corrections made in the original processing (subtracting 0.003 from the temperature and recalculating salinity) were inappropriate. The steps from the 2nd calibration to the end were rerun to produce new CTD and CHE files. The 2nd calibration step was changed to include only a correction for the Oxygen:Dissolved:SBE channel.

At the same time errors in the header comments about salinity bottle inserts were corrected in the CHE files.

# Particulars

48. Bottles fired at surface for storage; no sampling.

50. Depth missing in header – fixed in raw files.

54. Stop at 80dm in downcast to fix wire angle.

63. Stop at 167m in upcast to fix wire angle.

65. Stop at 15m & 87m in downcast to fix wire angle.

69. Stop at 107m to adjust wire angle.

# CRUISE SUMMARY  2014-60   CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1. | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information CTD #0443 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 7May2014 | Factory |  |  |
| Conductivity | 2280 | 7May2014 | Factory |  |  |
| Secondary Temp. | 2663 | 4Jun2014 | Factory |  |  |
| Secondary Cond. | 2754 | 7May2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 997 | 3May2014 | Factory |  |  |
| PAR | 4615 | 16Mar2011 | IOS |  |  |
| SeaPoint Fluor | 3642 | Jan. 2014 |  |  |  |
| Pressure Sensor | 77511 | 14Jul2014 | Factory |  |  |
| Altimeter | 1204 | n/a |  |  |  |