## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27 Mar 2025 | Updated channel names & formats in TOB files. G.G. |
| Jan. 2019 | Bottle spreadsheet converted to searchable BOT files. |

## PROCESSING NOTES

Cruise: 2014-66

Agency: SAFE

Location: B.C. Coast

Project: High Seas Salmon

Party Chief: Thiess M.

Platform: W.E. Ricker

Date: 9 October 2014 – 31 October 2014

Processed by: Germaine Gatien

Date of Processing: 25 March 2015 – 15 April 2015

Number of original CTD HEX files: 55			Number of CTD files: 55

Number of original TSG HEX files: 1			Number of TOB files: 1

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. Attached were an SBE 43 DO sensor (#1176) and a WET Labs ECO_AFL Fluorometer (#2215) on the primary pump and an unpumped SBE18 pH sensor (#0852).

The deck unit was a Seabird model 11.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

An SBE21 Thermosalinograph (serial #2488) sampled every 30s.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log was in good order.

During downcasts the descent rate of the CTD was often very high with frequent sudden decelerations and reversals in direction. This has not been noted previously on Ricker cruises, but is often seen on Tully cruises. There are associated excursions in salinity values that are not seen in temperature, so are assumed to be due to significant variations in alignment of T and C sensors. Corruption of data due to shed wakes are also associated with these decelerations, but show up a little later in the record. It seems likely that these decelerations and reversals in direction are also hard on the cable.

Salinity calibration sampling was not done in the usual way with the depth and length of bottle stops varying more than usual. There was a post-cruise calibration available and a shallow bottle comparison is of limited value in areas where vertical salinity gradients are high, so only a cursory comparison was carried out. The primary salinity was recalibrated by adding 0.002.

Only a cursory comparison was made of the CTD fluorescence and extracted chlorophyll samples for the same reasons as for salinity.  For low CHL (<2ug/L) the fluorescence is 1 to 2 times the CHL, and was about 0.6 times CHL for higher values. This is similar to the results found for 2014-36 when a fuller comparison was made.

The fluorometry traces show little spiking so the sampling rate was presumed to have been set low. This rate is adjusted within the sensor itself and is set before the cruise. Normally we have a lot more detail in this channel and use a filter to reduce it. No filter was applied to these data.

There was dissolved oxygen sampling for only the 1st of the 7 cruises that have used this sensor between factory visits. The results of cruise 2014-01 were used to recalibrate DO. There was a long gap between the two cruises but the post-cruise factory calibration suggests a drift that is reasonable close to the correction applied. The surface oxygen saturation values are fairly low (60% - 100%). The lowest values are associated with low vertical gradients in dissolved oxygen near the surface, so vertical mixing is likely the cause rather than calibration problems. There are also few data in the top 4m, so “real” surface saturation rates are likely to be a little higher at sites with less well-mixed surface waters.

Thermosalinograph data do not include flow rate, intake temperature or fluorometry and there were no loop samples available. Because there was no intake temperature available, a proxy was created, called Temperature:Primary, which was derived by recalibrating the lab temperature by subtracting an estimate of heating in the loop. Recalibration was based on comparisons to CTD salinity and temperature, and for temperature there was some history of previous use on the Ricker that was helpful.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book was obtained. The only note that might affect data processing was that for cast #32 the bottom depth changed during deployment from 192m to 183m. The log entry is 183m, and the header will be changed to 183m after conversion to IOS header format.

The same configuration file was used throughout the cruise.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity, fluorometer and DO sensors were obtained.

The extracted CHL, nutrient and salinity analysis spreadsheets were obtained and the data from each were combined in spreadsheet 2014-66_bottle_data.csv.

The calibration constants were checked for all instruments. The only change made was to the dissolved oxygen factor E (from 3.6 to 3.75) to reflect tests run using 2014-01 data to fine-tune the hysteresis setting. This will not affect this cruise since the hysteresis correction will not be needed, but it was changed to avoid confusion in the future.

##### Conversion of Full Files from Raw Data

All files were converted using con file 2014-66-ctd.xmlcon with hysteresis correction turned off.

A few casts were examined and all expected channels are present.

The two temperature and conductivity channels are fairly close to each other during the downcasts but are noisier in the upcasts. There are a few more spikes in the secondary conductivity than in the primary.

The fluorometer, pH and dissolved oxygen traces look normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A setting of +4s was used during 2014-01 to advance the DO voltage. Checks of the offset between distinctive features in T and DO suggest that +4s is right for this cruise too.

ALIGNCTD was used to advance the DO Voltage by 4s relative to the pressure.

Plots were examined before and after that step and the results with +4s look good.

##### CELLTM

Tests were run on 4 casts to see if the settings that were found best when this equipment was used during 2014-36 were suitable for these data and they were. There are places where they did not improve the data but that was because the upcast and downcast were very different. Where they are similar, these settings brought them into better correspondence in T-S space.

CELLTM was run using (α = 0.03, β=9) for both the primary and secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

When these sensors were used during 2014-01 there was so much noise in the data that no reasonable comparison could be made between channels. For cruises 2014-16, 2014-17, 2016-36 and this cruise, the casts are all shallow, making comparison less reliable than if we were looking at very low gradient zones as found in deeper water. The deepest sampling was about 250m, so 2 casts each from the 2 previous cruises were compared with 3 from this cruise at 240db. The differences were:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-16-0224 | 240 | +0.0003 | +0.00013 | +0.0011 | Fairly steady |
| 2014-16-0087 | 240 | +0.0004 | +0.0002 | +0.0017 | Fairly steady |
| 2014-17-0155 | 240 | -0.0001 | +0.0004 | +0.0006 |  |
| 2014-17-0170 | 240 | -0.0002 | -0.00005 | -0.0004 |  |
| 2014-36-0016 | 240 | -0.0003 | -0.00003 | ~0 noisy | Fairly steady |
| 2014-36-0141 | 240 | -0.0002 | -0.0001 | -0.0013 | Fairly steady |
| 2014-36-0162 | 240 | -0.0004 | -0.0001 | -0.0011 | Fairly steady |
| 2014-36-0174 | 240 | -0.0002 | -0.0002 | -0.0017 N | Fairly steady |
| 2014-66-0045 | 240 | ~0 N | -0.00043 | -0.0044 | Fairly steady |
| 2014-66-0159 | 240 | -0.0001 | -0.00058 | -0.0062 | Fairly steady |
| 2014-66-0181 | 240 | ~0 | -0.00064 | -0.0072 | Fairly noisy |

The differences look quite different from those of 2014-36. There is some suggestion that the differences between the conductivity and salinity are growing. It will be seen later posts-cruise calibrations show that drift was smaller than this study shows. There may be differences in the CTD channels while moving that are due to alignment variations rather than calibration drift.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number. As is typical of this type of fluorometer, there were some negative values, so Fluorescence values <0ug/L were replaced with pad values.

##### Checking Headers

The header check was run and some negative values were found in pressure.

The surface check gave an average value of 3.7db which is not out of line for the Ricker. There were a few slightly negative pressures with in-water salinity values, but they were isolated and seen only at the end of casts after the pumps were turned off. There are examples of pressure ~0.2db with in-water salinity values. The pressure seems reasonable.

The cross-reference check was compared with the log book and some errors were found:

- The CTD cast at station JS07 was saved as 2014-66-0010 but in the log book and on the analyst’s records it was event #11. The file name was changed to 2014-66-0011.

- Station names were changed in the log book for events #3, 7, 80, 115, 118, 121, 124, 127, 130 & 141. Those were fixed in the headers of the CLN files.

The cruise track was plotted and added to the end of this report. No problems were found.

##### Shift

Fluorescence

A shift of +12 is the usual setting used to align the Wetlabs ECO fluorescence with temperature and it was used for 2014-36 when the same sensor was used. Tests were run on a few casts and the results look generally satisfactory.

SHIFT was run on all casts to advance the ECO fluorometer by +12 records.

pH

Because of hysteresis in the pH signal, tests are best run on casts with distinctive features in the pH traces. This was done by finding the distance between such features in the upcast and downcast and comparing that with the temperature offset after applying a variety of shifts. Settings between +50 and +60 records produced the best results overall during 2014-36 with the same equipment, and that setting looks good for these data as well.

SHIFT was run on all casts with the setting +55 records.

Conductivity

Tests were run on 4 casts using a variety of settings to see which shift produces reasonably stable T-S plots. The optimal choice varied slightly from feature to feature, but overall the best results were found by using -0.7 records for the primary conductivity and -1.3 records for the secondary.

Two runs of SHIFT were used to apply advance the primary by -0.7 records and the secondary by -1.3 records.

Dissolve Oxygen

Checks were made of the alignment of DO and temperature and the shift applied earlier (§5) appears to have done a good job though, as usual, for some features it was too little or too much.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds (taken from header)

COMMENTS ON WARNINGS: None of the warnings will affect the data to be archived as they concern pressure spikes during the upcast, mostly right at the surface.

##### Comparisons of CTD and bottle samples

Normally during the High Seas Salmon cruises the CTD is stopped at about 15m during the upcast and after a short wait the Niskin bottle mounted 5m above that is fired. The depth of the stop always varies a little but over a fairly small range so an average level is picked and CTD data from that level are chosen to be compared with the bottle samples. For this cruise the variability is much higher with many stops being as deep as 18m. And some of the stops at 15m were shorter than the 30s recommended.

Salinity

Near-surface salinity comparisons are not very useful for calibration purposes unless surface waters are very well mixed to at least 15m; they were not for this cruise. There are also some limitations because a small mismatch in depth can lead to large errors in high salinity gradients and if flushing of Niskin bottles is not excellent there are further errors in large vertical gradient zones. There was a post-cruise calibration available, so the usual comprehensive study comparing bottles and CTD salinity was not done since it would require much more effort than usual to determine the appropriate depth from which to select CTD data. A few casts were examined where the near-surface gradient was relatively low.

| Event | Station | Salinity:Bottle | CTD Sal0 | CTD Sal1 | CTD sal0-bottle | CTD sal1-bottle |
| --- | --- | --- | --- | --- | --- | --- |
| 85 | EP 07 | 32.0381 | 32.0246 | 32.0212 | -0.0135 | -0.0169 |
| 108 | VI 2 | 31.7856 | 31.7775 | 31.7744 | -0.0081 | -0.0112 |
| 141 | T 08 | 31.9462 | 31.9406 | 31.9366 | -0.0056 | -0.0096 |

The best-mixed of the three was event #141. In all cases the CTD salinity was lower than the bottle CTD with the primary closer to bottles than the secondary.

Fluorescence

Comparison of CTD fluorescence with extracted CHL samples is even more difficult as there is a lot of noise in the fluorescence near the surface. It would require a lot of data to reach any conclusion. However, the 3 casts examined show results similar to those found in cruise 2014-36 when the same fluorometer was used. The fluorometer reads between 1 and 2 times the CHL value for CHL<2ug/L, and for CHL>2 it reads relatively low.

| Event | Station | CHL | CTD Fluor | FL/CHL |
| --- | --- | --- | --- | --- |
| 85 | EP 07 | 0.74 | 1.4 | 1.9 |
| 108 | VI 2 | 2.07 | 1.2 | 0.6 |
| 141 | T 08 | 0.82 | 1.4 | 1.7 |

- Other Comparisons

Previous experience with these sensors –

Salinity:

The conductivity and temperature sensors were both recalibrated in late January 2013 and they were used for 6 cruises since then prior to this cruise. Of those, 2014-01, 2014-16, 2014-17 and 2014-36 have been processed. For 2014-01 the primary was used for only 2 casts and was recalibrated to match the secondary salinity which was not recalibrated. There was either no sampling or inadequate sampling for recalibration purposes from the 3 cruises that followed. The post-cruise calibration was used to make an estimate of drift for 2014-36.

Dissolved Oxygen:

The DO sensor was used for 6 cruises prior to this one since it was recalibrated in February 2013. Of the 4 cruises that have been processed only 2014-01 included DO sampling so the fit from that cruise was used for all the others (slope/offset = 1.0502/+0.0142). During 2014-01 data hysteresis tests were done and factor E was fine-tuned.  The calibration drift seen in the post-cruise calibration was roughly 5%.

Pressure:

The pressure sensor was recalibrated in May 2013 and the factory offset was found appropriate for the 4 cruises since then that have been processed.

Historic ranges – Many casts only have climatology covering a large area, so it is not useful close to shore. Where the climatology looks appropriate there were a few excursions at the bottom of casts with temperature being slightly high and salinity slightly low. Similar results were found during 2014-36.  For 2014-36 investigations were made to see if there might be a pressure error causing these excursions, the pattern of where they occurred did not match that theory. These outliers look like real variations not the result of systematic instrumental problems or calibration drift. High temperatures and low salinity in the upper ocean have been reported from many parts of the north-east Pacific in 2014.

Repeat Casts –

There were no repeat casts.

Post-Cruise Calibration

Post-cruise calibrations were obtained from March 2015 for the primary conductivity and December 2014 for the other temperature and conductivity sensors. They show that the primary and secondary temperature sensors were reading slightly low, by 0.0008C° and 0.0007°, respectively. The conductivity sensors had drifted so that the primary salinity was low by 0.0026 and the secondary low by 0.0046 salinity units. The combined effect produces primary and secondary salinity low by about 0.002 and 0.004psu. Since this was the final use of this equipment before the factory service, it is assumed that all the error had occurred by the time of this cruise.

- DETAILED EDITING

The first issue is to decide which sensor pair to edit. For most casts there is little to distinguish between them, but the primary salinity looks slightly less noisy so the primary channels were selected for archiving.

CTDEDIT was used to remove spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records.

During downcasts there were some cases of very high descent rate with frequent sudden deceleration. Associated with the sudden slowing we see excursions in salinity values that are not seen in temperature. This is not the usual problem of low descent rates leading to corruption by shed wakes which would also affect temperature, though shed wake trouble usually occurs as well, just a little later. It is presumed that the salinity excursions are due to variations in alignment of the T and C sensors. The most extreme cases were for events at exposed sites such as #66, #68, #118, #121 and along the EP and T lines with descent rates getting up to 2.5m/s. In one case the CTD was going down at 2.5m/s when it suddenly slowed and within 3s reversed to -1m/s. As well as causing problems with salinity, this may be a problem for the cable. The sections with high descent rate also lead to a reduction in the amount of data that goes into bin averages.

All casts required some editing.

- Initial Recalibration

The pressure does not need recalibration.

Various limitations in the comparison of CTD salinity with bottles make it unreliable for determining what recalibration should be applied. The post-cruise calibration indicates that the primary salinity was low by about 0.002 in March 2015. There were no cruises between this one and the factory visit, so the primary salinity will be recalibrated by adding 0.002.

The only information available about the SBE dissolved oxygen sensor since it was last recalibrated comes from 2014-01, so the correction applied to that data set will also be used for this cruise:

DOX_CTD Corrected = DOX_BOT = 1.0502* DOX_CTD +0.0142

File 2014-66-recal1.ccf was prepared to apply the above correction to Oxygen:Dissolved:SBE and to add 0.002 to the primary salinity.

CALIBRATE was run on the EDT files.

##### Fluorometer Processing

At this stage a median filter, fixed size=11, is usually applied to reduce spikiness in fluorescence channels, but examination of all casts shows little spikiness, so this step was skipped. It looks like the sampling rate was set low on the sensor.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate and Flag

Profile plots were produced at this point to check for errors. No problems were found.

T-S plots were produced; there are some small unstable features that likely reflect real conditions.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

----------------------------------------------------------------------------------

Data Processing Notes:

----------------------

Fluorescence and pH data are nominal and unedited except that some

records were removed in editing temperature and salinity.

Dissolved oxygen was calibrated using the results of cruise 2014-01 when the

same sensor was used. There was no calibration sampling from this cruise.

Recalibration of the salinity channel was based on a report from the factory on

drift in the temperature and conductivity sensors as of 23 December 2014 and

3 March 2015, respectively. This was the last cruise before the factory checks.

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were collected.  Calibration is required for each cast to get

absolute values, although general trends within a cast are likely real.

For details on the processing see the report: 2014-66_Processing_Report.doc.

---------------------------------------------------------------------------------------------

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

The track plot looks fine.

The sensor history files were updated.

##### Dissolved Oxygen Surface Saturation

Dissolved Oxygen saturation was derived and plotted. The surface saturations varied from about 70% to 110%. Values <70% were found in Johnstone Strait and in Juan de Fuca Strait, both areas where DO tends to be low due to vertical mixing; DO values vary little in the top 20m in those areas. In Queen Charlotte Strait values are a little higher ranging from 75% to 100% with the highest saturations to the north near Queen Charlotte Sound. In the more open waters along the T and EP lines most values were around 95% to 100%. Saturation rates were highly variable in inlets but for Rivers Inlet they ranged from 70-90%. In general the saturation rates reflect the DO gradient near the surface. There are few data above 4m so it is expected that right at the surface the values would be a little higher for those casts that are not extremely well-mixed.

These values look reasonable but we have no DO sampling to confirm them.

##### Thermosalinograph

Data were provided in 1 hex file.

There was no external thermistor or flow meter.

There were no loop salinity samples taken.

a.) Checking calibrations

The configuration file was checked and the parameters are correct.

The file was saved as 2014-66-tsg.con.

b.) Conversion of Files

The hex file was converted to CNV files using configuration file 2014-66-tsg.con. It was then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the header.

ADD TIME CHANNEL was used to add Time and Date channels based on the Julian time.

A track plot shows that the TSG was not in use until 25 October at 23:41, so the first overlapping CTD cast will be event #139.

Time-series plots were produced. No obvious problems were found.

The track plot looks fine and was added to the end of this report.

c.)  Checking Time Channel

The CTD data were exported to a spreadsheet which was saved as 2014-66-ctd-tsg-comp.xls.  Data from events #139 to 190 at 3m and 4m only were selected.

18 CTD casts with data from 4m overlapped with TSG data and, of those, 16 also had data from 3m.

The TSG file was opened in EXCEL, median and standard deviations (over 5 records) were calculated for temperature and salinity and the file was reduced to the times of CTD files. Those data were added to the CTD data.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude and longitude were all <0.0004° and the median differences in both are 0.0001°. This shows both the times and positions are reliable for both systems.

This spreadsheet will also be used in step (d) to compare temperature and salinity.

d.) Comparison of T and S from TSG and CTD data

Salinity

The salinity comparison shows huge variability in differences in either direction ranging from the CTD being low by 5psu to high by 0.3psu. The data fall into 2 groups: when the TSG salinity is lower than CTD salinity, it is usually much lower; when it is higher, it usually within 0.005 of the CTD. The high cases are almost all in inlets while the low ones are in more exposed areas. There are higher near-surface salinity gradients in the inlets and salinity always increases with depth, so the water in the loop is likely coming from higher in the water column than we are sampling with the CTD. In the more exposed areas the gradients are very low below 3m and likely quite low above that. So getting the depth wrong likely matters less.

When only 8 casts with low vertical gradients are used (mostly offshore and 2 from QC Strait) we find that the TSG salinity is higher than the CTD at 4m by about 0.004 which either means the loop water is coming from deeper than 4m or that the TSG salinity is actually reading high; the latter is much more likely given the results found in the inlets. There are only 6 of these CTD casts with values at 3m and the TSG salinity is higher than those by about 0.0045.

It is likely that the salinity is high by at least 0.0045. The CTD salinity was recalibrated based on a factory calibration, so it likely good to at least 0.001. Recalibration of salinity by subtracting 0.0045 looks justified.

Lab Temperature

Like the salinity comparison the differences between the TSG and CTD temperatures mostly fall into 2 groups. There are 10 cases that show the TSG reading too high by between 0.016C° and 0.088C° with most values between 0.05 C° and 0.08 C°. Where the TSG reads low the range is from 0.004C° to 0.30C°.  So the average of all values is not very useful. If only the 8 casts from exposed areas chosen for the salinity comparison are included the TSG temperature is found to be higher than the CTD by a median of 0.065C° based on 3m and 0.062C° based on 4m CTD.

The temperature gradient is more variable near the surface than that of salinity, sometimes increasing with depth and sometimes decreasing. There are 5 cases where the CTD profiles show that near-surface T and S gradients are both low and the standard deviation in the TSG T and S are also low. Using those casts the TSG temperature is found to be high by a median of 0.06C° at either 3 or 4m.

The most recent information available from other uses of the Ricker TSG comes from 2 cruises in 2012. Cruise 2012-67 was in October but there were only 3 CTD casts for comparison. Cruise 2012-03 was in March and showed the temperature to be high by 0.065C° with 40 casts for comparison. The near-surface water temperatures were lower for 2012-03 (mostly 7 to 8C°) than 2014-66 (around 10 to 13C°) for the data included in the comparison) and the heating in the loop is temperature dependent with more warming as the difference between the intake temperature and the ambient ship temperature increases. So it makes sense that the heating in the loop would be lower for 2014-66 though one might expect it to be even less than 0.06C°.

While the information available is very limited it is expected that the temperature will be too high in the lab due to heating in the loop and the best guess is that it is high by 0.06C°.

Conclusions

1. The TSG clock appears to have worked well.

2. The TSG temperature is likely higher than the CTD temperature by about 0.6Cº. This may include error due to calibration drift and heating in the loop but the former is likely to be much smaller than the latter.

3. The TSG salinity appears to be higher than the CTD salinity by at least 0.0045.

e.) Editing

CTDEDIT was used to examine the ATC file but no editing was found necessary.

g.) Recalibration

Because there is no intake temperature available we want to create a proxy for that. So channel Temperature:Lab was created,  set equal to Temperature:Primary and placed after Temperature:Primary.

File 2014-66-tsg-recal1.ccf was used to subtract 0.0045 from salinity and 0.060 from Temperature:Primary which will serve as a proxy for Temperature:Intake.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Conductivity:Primary, Flag and Position:New channels.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to ensure channel names have standard names and formats. Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | No | Yes |

| Calibration Information CTD #443 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2710 | 31Jan2013 | Factory | Mar2015 | Factory |
| Conductivity | 2128 | 29Jan2013 | Factory | Dec2014 | Factory |
| Secondary Temp. | 2374 | 31Jan2013 | Factory | Dec2014 | Factory |
| Secondary Cond. | 3396 | 30Jan2013 | Factory | Dec2014 | Factory |
| SBE 43 DO sensor | 1176 | 5Mar2013 | Factory |  |  |
| WetLabs ECO-AFL/FL | 2215 |  |  |  |  |
| SBE18 pH | 0852 | 21April2011 | Factory |  |  |
| Pressure Sensor | 75636 | 25Mar2013 | Factory |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/2488       Cruise ID#:	2014-66

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 7Jan14 | Factory |  |  |
| Conductivity | 2488 | 7Jan14 | “ |  |  |