## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: CFMETR Nanoose

Agency: Canadian Forces

Project: Canadian Forces Maritime Experimental Test Range

Area: Nanoose

Date: 12 January 2014 – 16 December 2014

Platform: CFAV Stikine (and others)

Processed by: Germaine Gatien

Date of Processing: 25 February 2015 – 2 March 2015

Number of original CTD casts: 78

Number of casts processed: 76 (2 files contained bad data)

INSTRUMENT SUMMARY

Six Sea Bird Model SBE 19 SEACAT CTDs were used – serial numbers #1449, 1450, 1451, 1452 1510, 1511, 1512 and 1513. Each CTD had an altimeter attached.

SUMMARY OF QUALITY AND CONCERNS

Six different CTDs were used. The calibration information could not be confirmed.

To assist in analysis event numbers are helpful and the software in use will only allow 4 digits which suits most cruises. However a 7-character system is used for this project. A 4-character format was derived from those IDs based on the month and day of each cast. This means the same event # will appear if there is more than one cast on one day as frequently occurred.

The soak period for the CTD varied greatly from about 20s to over 2 minutes, with most lasting a minute or more. It is recommended that the SBE19+ be soaked for at least 2 minutes to improve salinity quality, though in the absence of large differences between air and water temperatures it is likely that 1 minute is sufficient.

There was no salinity calibration sampling. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors are expected to be as high as 0.05units in areas of high gradients. Away from high gradients errors in salinity are expected to be on the order of 0.005.

The times given in the SeaBird headers do not include a time zone, so it was assumed that they are in PST. All times were converted to UTC by adding 8 hours.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

78 hex and configuration files and 77 hdr files were delivered.

- Conversion of Raw Data

- The raw data were converted by matching the configuration file to the input file name.

- Positions are missing from the headers of the files but were provided in a spreadsheet. The information was saved as “NanooseCtdStations2014.csv” and will be used later to add the information to headers.

The format of latitude and longitude in the spreadsheet needed to be changed to DD MM.MM N and DDD MM.MM W. Concatenate was used to do this. For example for column LOC:Latitude the entries are found using:  =CONCATENATE(INT(C2),"  ",TRUNC(((C2)-INT(C2))*60,2)," N ") where C2 is latitude. There needs to be a space before the latitude and around the N and W. (There is no need to add “! (deg min)” – that will be handled later in the MERGE routine.)

A column LOC:Station was added with every row having “Nanoose” entered.

A column LOC:Event Number was added with a 4-digit id based on month and day of cast. If there are more than 2 casts on a day they will have the same event number; the event number will not be used to rename the files in standard format because of that.

- There are some differences between the spreadsheet and the raw files available, such as no position information being available for some files and some naming discrepancies. The actions required to resolve the discrepancies are as follows:

- n14012201 – position information missing - use position from n14012202

- n1402120* - only 6 casts in position file – there were actually 7 casts that day, all at different times. Used position from 6th file as position for 7th

- n14022501 – position information missing - use position from n14022502

- n14061101 – position information missing - use position from n14061001

- n14070901b – some files saved as 1b and some as 2 – used con file from n14070901b and data file from n14070902 – converted file named n14070902

- n14071402 – only one cast file for that day, so would expect n14071401, but possibly other cast was taken and discarded, so name was left unchanged.

- n14071503 - not standard name since only 1 cast file available for that day, but perhaps previous casts taken and discarded, so name left unchanged

- n14082701b / n14082702– some files saved with this name and some with n14082702 – look like all from same cast. Used con file from 1b and hex from 2 and saved converted file as n14082702.cnv.

- n14091103 – only cast from that day but perhaps earlier casts were aborted – name not changed

- n14091602 – only cast from that day but perhaps earlier cast was aborted – name not changed

Where position information was taken from another cast, a note to that effect will be added to the comments in the final files.

The spreadsheet did not contain the initial “n” found in the file names and was missing the final digit which was either a 1 or 2 since on some days there were 2 casts. Those items were added to the file names in the spreadsheet and the file was saved as a csv file.

- The soak period was typically 1.25minutes. A 2-minute soak is generally advised for the best salinity data, but in most cases the conductivity seems to have reached equilibrium before the downcasts started.

- Temperature and conductivity generally look ok.

- The altimetry looks useful near the bottom when the CTD got close to the bottom.

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature.  Tests were run on a few casts to ensure this was the best choice.

The pressure was filtered with a time constant of 0.2s to increase the pressure resolution.

Plots of a few casts before and after this step show that reversals in pressure were removed without oversmoothing. A setting of 2s is often used for SBE19 CTDs but that looks too high for these data. The pressure is generally very smooth with just a few examples found that show signs of needing a filter.

- ALIGNCTD

An initial run with no alignment applied to the temperature channel produced poor results with instabilities in T-S plots. Tests were run on a few casts with different CTDs and the best results were found when temperature was advance by 0.73s relative to pressure. That is close to the setting used in 2012 (0.7s). There are small instabilities left in many casts. Alignment could be fine-tuned later using SHIFT if further adjustment is found useful.

- DERIVE

Program DERIVE was run to calculate salinity

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers.
The program MERGE CSV FILES TO HEADERS was then used to add that information to the headers.

8. Checking Headers

Track Plot was used to ensure all positions were close and they were. The plot was added to the end of this report.

The times are said to be UTC but are clearly local time, so ADD TIME CHANNEL was used to add 8 hours. It is not known if PST was used throughout the year, so the times may be off by one hour when PST was in effect. It is more likely that the computer time was not updated to PDT.

Header Check was run and makes it clear that there were some large spikes, with file n14071402 being most notable. There were 7 different equipment configurations with 6 different CTDs (S/N 1449, 1450, 1451, 1511, 1512, 1513) and a change of altimeter for one CTD (1511) in December.

Surface check was run and the average surface pressure was 0.6db.There were some negative values, but the only one that is significantly so is file #n14042401, with a pressure of -5.95db throughout the soak period. Comparison of temperature and salinity with casts before and after show a very large difference in temperature and salinity, so clearly this cast data is bad and will not be processed further.

- An attempt was made to study how each pressure sensor behaved, but this is very time consuming. The pressure from #1513 looked slightly higher than the others early in the year but very close to average later in the year; so variations among the readings are more likely due to how the CTDs were deployed rather than calibration issues.

- The pressure sensors are considered ±1db.

- The altimetry header information from each file was exported to a spreadsheet. In most cases the CTD did not get within 15m of the bottom, so there is no header entry. A few files with no altimetry header were checked and either the altimetry was very spiky with no clear signal or there was a near-constant value. All cases with an entry were checked to ensure that spikes were not interpreted as a signal. There was just 1 such case; the header entry was removed from file n14021204.mrh as it was clearly just a misinterpretation of a spike.

9. SHIFT

These data were aligned earlier in the process, but tests were run to see if a further shift might reduce unstable features but nothing helped.

10. CLIP

CLIP was used to remove the first 120 scans.

11. DELETE

Then DELETE was run on the clipped files using the following parameters:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

There were warnings for events #0619 and 0831 during the first run. The MRH files for those casts were edited and then CLIP and DELETE were rerun. There were no warnings on the final run of DELETE.

12. CTDEDIT

The DEL files were copied to EDT.

In the course of editing a few cases were found where CLIP had removed too much data because the soak period was a very short. In those cases, DELETE was run on the MRH files instead.

DEL files were opened in CTDEDIT. Most of the casts required light editing, mostly to smooth cases where unstable features appeared to be due to minor misalignments of T and C in the presence of temperature gradients. There were also a few cases of shed wake corruption near the bottom and poor data at the surface.

For file n14071402 there was severe corruption with many off-scale spikes. The file would not open properly in the editor. An attempt was made to remove the largest spikes using a text editor to put in pad values. That did enable display in CTDEDIT where more bad records were removed. However, plots made together with other casts from about the same data show that these data are out of line, so the file will not be processed further.

The edited files were copied to EDT.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure      	Averaging interval = 1.000

Minimum bin value = .000	Average value will be used.

Interpolated values are NOT used for empty bins

Profile plots were examined on-screen and no problems were found.

Page plots were examined and look ok with only a few very slightly unstable features.

14. Intercomparisons

Historic ranges – There were some excursions from the local climatology. Temperature and salinity were both low around 200db in mid-February. In March to early June temperature values were a little below the climatology minimum somewhat deeper. In the autumn there are some low temperature values around 150db. There were a few low near-surface salinity values and in September salinity was low near the bottom. The 3-standard deviation limits are too severe for this area and the excursions are small. There is no indication of calibration problems.

Repeat casts – Casts were all close in space, and on a few occasions there were a group of casts on one day. Ten casts were examined that occurred on Feb. 11 and 12th. Two different CTDs were used. From 200db down the variations are fairly small but there was clearly a shift to cooler, saltier water between the 2 casts on Feb. 11. There is little variability after that shift.

15. CALIBRATION

No recalibration was applied as there was no salinity calibration sampling.

16. REMOVE and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity, Descent_Rate Altimetry and Flag.

The HEADEDIT routine was used to fix formats and to add the following comment to the headers:

SeaCat salinity is prone to error due to a mismatch of conductivity and

temperature response times, especially when the descent rate is non-uniform.

Such errors are likely to be larger than calibration errors. SeaCat salinity

errors are expected to be as high as 0.05 units in areas of high gradients,

but on the order of 0.005 where gradients are low and descent rate steady.

There were no salinity calibration data available for this cruise.

The standards check routine was run and no problems were found.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced. No errors were found.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	Nanoose Test Range

Dates: 	Start:	15 January 2014	    End:	16 December 2014

Location: Nanoose Test Range 	Vessel:    Stikine (and others)

NOTE: Six SEACATs were used for this cruise.

The calibration information used in conversion can be found in the headers of the CTD files.