## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 18 Jan. 2016 | Correction to station name in file 2015-03-0049.CTD |
| 11 Dec. 2015 | Corrections made to CHL in file 2015-03-0087.CHE and to spreadsheet QF2010-03 CHL.xlsx |

## PROCESSING NOTES

Cruise: 2015-03

Agency: Ocean Sciences Division

Location: Douglas Channel

Project: WC-DC

Party Chief: Vagle S.

Platform: Vector

Date: March 5, 2015 – March 12, 2015

Processed by: Germaine Gatien

Date of Processing: 29 May 2015 – 9 June 2015

Number of original HEX files: 40 (+5 test files and 1 accidentally created – no useful data)

Number of CTD files: 39	Number of bottle casts:	19

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0506) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#1438), a SeaPoint Fluorometer (#3642), a Biospherical QSP-400 PAR sensor (#4565), a surface PAR (#16504) and an altimeter (#1252).

The CTD deck unit was an SBF model 11+, serial number 0619.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log and rosette log sheets were in fairly good order with many comments about problems encountered. There was confusion about which pump was which so that the equipment list has some known errors and some doubtful entries. Watch keepers initials were missing early in the cruise, but from event #24 onwards they were entered.

The deployment method did not include a soak at 10m.

The file name for event #62 was accidentally entered as 2015-03-0074 in acquisition, so event #74 was not used to avoid overwriting of the file. In processing the file name was corrected.

The cast type was mistakenly entered as CTD when there were no samples taken; unless the CTD was removed from the rosette it should be entered as ROS. The distinction is important since it indicates whether the fish was mounted vertically or horizontally. If there are problems such as irregular flow to the pumped channels, this information may help determine the source of trouble.

There was an error in the headers: because the entry for station name was entered using “Stn” instead of “Station” that information would be lost in conversion to IOS HEADERS. This was corrected in processing. There have recently been a number of errors in the formats for information in the SeaBird headers for Station Names and in Water Depths. The correct formats are “Station:” and “Water Depth:”.

A problem was found in the secondary pump after cruise 2015-17. This could be a problem for the external sensor attached to the pump but there was a problem figuring out which one it was since the entries in the log book equipment list contained a number of errors. Fortunately, there were photos taken which established that the dissolved oxygen sensor was on the secondary pump. A close examination of data shows that the temperature was noisier than usual and the secondary conductivity took longer to equilibrate when the pumps were turned on and also after shed wakes passed through at the beginning of CTD bottle stops.  For most casts there was no obvious effect on the dissolved oxygen sensor data, but this may not be significant because DO is always slow to respond. For a few casts at the heads of inlets the very large temperature gradient does lead to spikes in DO in the top 10m that may be partly due to the pump problem. The primary temperature and salinity were selected for editing because of the problem.

A second problem that affects bottle data is that most of these casts were in protected areas where the descent and ascent rates were very steady. It is likely that this will lead to poor flushing of the Niskin bottles.

The Oxygen:Dissolved:SBE data are considered of lower quality than usual due to poor flushing of Niskin bottles and the problem with the secondary pump that may cause poor response of the SBE sensor. Re-calibration was based on the results of cruise 2015-01 when the same sensors were used, but there remain some doubts due to the pump problem. These data, especially in high gradient zones, should be used with caution.

Salinity:T0:C0 data were also recalibrated based on 2015-01 results due to concerns about the flushing of the Niskin bottles during this cruise. While this is the best available comparison, it included a lot of scatter.

There was a constant signal in the PAR:Reference channel with a very low value for the downcast of event #1, but it dropped to 0 during the upcast and had no signal for the rest of the cruise. The channel was removed from all casts. PAR is generally low but most casts were at night.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX. All files had an error in the event number section of the file names. Those were corrected before conversion.

There is an error in the header entries: STN was used instead of Station so the information would not be picked up in conversion to IOS header format. So this was corrected in the HEX files.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets for dissolved oxygen, extracted chlorophyll and salinity.

There are errors in the equipment list in the log book. A note indicates that mentions that they were unsure which were the primary T/S sensors and which secondary, but they also reversed temperature with conductivity sensors. The configuration file makes the details clear for T/S. Given the confusion over which pump was which, it is likely that the entry about which pumps had the oxygen sensor and fluorometer attached is likely also wrong. The arrangement on cruises before and after this one with the same equipment had the reverse of what was entered in the log for this one. There are no photos. Perhaps there could be a clearer indication of which pump is which.

There were many notes in the log book about problems encountered. Early on there were no initials of watch keepers, but later they were entered.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked.

The XMLCON files did not change through the cruise.

The calibration constants were checked for all instruments and the only corrections needed were to correct spelling of the date for the transmissivity. The corrected file was saved as 2015-03-ctd.xmlcon.

Because other recent uses of this CTD have found that the pressure is too high by roughly 1.25db, initial conversions were done to see if the offset should be updated for this cruise. There was a single deck pressure measurement of –0.47db and surface pressure 1db. The pressure sensor is considered ±1db so a single measurement is not reliable, but it suggests that the pressure is low by ~0.5db.

Yet, examination of a few converted files shows many instances of pressures as low as -1.26db with transmissivity values that show the CTD was in-water. For cast #81 with pumps on, the upcast has a section of pressures of -1.4db with transmissivity of 77%, conductivity ~3.1 and fluorescence of 1.1ug/L. The negative pressures mostly occur with pumps off, so conductivity and fluorescence are not very useful since they are all pumped. Nonetheless, even for the pumped channels we expect those values to eventually drop if the CTD is held out of water for some time. For these data there was no sign of conductivity values falling while held at -1.1db. Other evidence of pressures being too low is that the surface bottles for casts #29, 32, 57 and 81 have pressures 0.9db, 0.5db, 0.6db and 0.4db with pumps on; if those readings were correct the tops of the Niskin bottles would have been out of the water at the time the bottles were fired. This is unlikely.

Similar results were found during 2014-19, 2014-50 and 2015-01 and an initial examination of 2015-17 shows evidence of pressures being too low as well.

File 2015-03-ctd.xmlcon was revised to change the pressure offset from -1.50646 to -0.25.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2015-03-ctd.xmlcon to create CNV files.

Cast #17 did not convert properly. Examination of the raw files turned up no reason for this, but when 2015-03-0017.xmlcon was used conversion was successful. The same corrections applied to 2015-03-ctd.xmlcon, were applied to the xmlcon file for event #17 which was then saved as 2015-03-ctd17.xmlcon. Conversion was then repeated using that file.

A few casts were examined and all expected channels are present, but there was no Surface PAR signal except for a tiny constant signal in the first cast. The PAR is mostly very low because most of the casts were at night. The temperature and conductivity channels are close on downcasts. As usual the upcasts are noisier. The altimetry looks useful and fluorescence, transmissivity and PAR look ok.

The descent rate of the CTD is mostly high and steady.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2015-03-ctd.xmlcon and 2015-03-ctd17.xmlcon.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files and no significant outliers were found.

A preliminary header check was done and no problems were found. Fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

A few problems were found:

- Cast #14, Niskin bottle #13 was fired at 19.6db but was given no sample # and there was no sampling, so this was dropped from the ADDSAMP list. (The line for this record was later removed from the SAMAVG file.)

- Cast #17 had 3 bottles fired but only 1 has a sample number and only 1 was sampled, so the other 2 lines were removed from the ADDSAMP list.

- Cast #57 shows 16 sample #s including sample #247 but #247 is not shown on the rosette log sheet and there were only 15 bottles fired.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine.

Those files were then bin-averaged and called SAMAVG.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2015-03-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2015-03chl*.xls. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2015-03chl.csv which was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2015-03oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2015-03oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2015-03SAL.xls. The analysis was done within 14-19 days of collection. The files were simplified and saved as 2015-03sal.csv. The analyst noted that there was only 1 pair of duplicates and they were far apart.

There were some missing samples flagged 9 but samples #214 and #215 from cast #43 were also missing so those were added to the CSV spreadsheet with 9 flags and pad values.

That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2015-03nuts.xls. This includes a precision study.

Then the file was simplified, reordered on sample numbers and saved as 2015-03-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. A few discrepancies were found:

- Cast #10 – There are no samples. Bottles were fired as a test. This bottle file will not be processed further.

- Cast #17 – As noted earlier 3 bottles were fired but only 1 sampled. The unsampled bottles were removed from the MRG file.

- Cast #34 – On the rosette sheet there are samples 172, 174, 176 but the labels were 173, 174, 176. The sample #173 value looks appropriate for the depth so this is assumed to be an error in the rosette log.

- Cast #87 – The log shows 4 CHL samples but there are only 3 analyzed with the 30m sample missing. This is likely an error in the rosette log. No other casts in this area had sampling at 30m.

##### Compare

Salinity

Compare was run with pressure as reference channel.

When bottles were excluded based on differences >0.02 and then based on residuals the primary salinity was found to be low by 0.0036 and the secondary low by 0.0029; both had a standard deviation of 0.004.

Many of the larger outliers are on the low side in the top 100m. This is suggestive of incomplete flushing of Niskin bottles, which is quite likely for this cruise as sea conditions were quite flat. Given that the vertical salinity gradients are relatively high, this could lead to a fairly large error. Analysis was done 6 to 10 weeks after collection which could also lead to the CTD looking low. While the type of bottle inserts used should minimize evaporation, there could well be some and it might explain some of the scatter in the results. There were 4 casts with bottles at 21m and they showed steadily increasing differences with time. When the vertical salinity gradients were examined, it was found that the first had a very low gradient, the 2nd had a slightly higher gradient, the 3rd had a fairly low gradient but a lot of variability while the last had the largest gradient. Incomplete flushing could explain the pattern; bottles containing water from a little lower in the water column would make the CTD look like it is reading low.

The CTD technician discovered after cruise 2015-17 that the secondary pump was not working properly and there are some clear signs of that in the secondary salinity traces. The problem may not affect the bottle comparison a lot, but there was clear trouble with downcast secondary data. So the emphasis was put on the primary salinity since that will have to be chosen for the archive.

The few bottles that were flagged by the analyst do not look notably out of line, so no changes to the flags are suggested. However, it is noted that the deepest bottle looks close to the trendline, but the 2 bottles were sufficiently different that the analyst flagged them 4. When the CTD salinity is compared with the two bottles, the CTD looks low by 0.0064 in one case and low by 0.0008 in the other, so having the average look close to the trendline is a matter of chance. It was removed from the comparison.

There were usually 2 bottles fired at each depth. They were analyzed in order of collection, so a comparison was done to see if there is a pattern of the 1st reading higher or lower than the second. No pattern emerged but significant differences are very common. Since COMPARE was run with pressure as the reference channel, the 2 bottles are compared with identical CTD salinity. COMPARE was rerun with Bottle number as reference. This is far harder to work with because pressure is so important in detecting flushing problems. Near-surface bottles were removed by referring to rosette sheets but still left a lot of scatter. The differences were a little smaller than found in the original comparison using pressure, but not enough to explain the differences. Some of the difference might be due to variations in flushing from bottle to bottle.  (See 2015-03-sal-comp1b.xlsx.)

One more approach is to look at the full profiles to see whether flushing is likely a problem below 100m. It does look like the local gradients could easily lead to the CTD salinity looking lower than it really is, but there is no way to quantify this. Casts with salinity samples at the bottom were examined and many of those have high standard deviations in the CTD salinity. The 5 examples found that have quiet CTD data have the CTD high by 0.029, 0.035, 0.002, 0.003 and low by 0.003. If flushing is an issue we would expect the CTD salinity to be higher than bottle salinity at the bottom and it certainly looks the case for at least 2 of the 5.

Based on the comparison the secondary salinity is higher than the primary by ~+0.0007 which is reasonably close to the differences described in section 9. The difference has changed sign from that of 2015-01 when the secondary was lower than the primary by ~-0.0015.

There are many problems with this comparison. There clearly has been a change in the secondary salinity between 2015-01 and the 2 later cruises both in the poor response of the secondary conductivity sensor in areas of large vertical salinity gradients and the fact that the secondary salinity is now slightly higher than the primary whereas it was lower for 2015-01. Given what we know about the secondary pump, it is wise to assume the secondary salinity is the one that has shifted most. But there is also a possibility that there could be some drift in any of the sensors. The results of 2015-17 do suggest that the primary salinity is reading at least a little high, but there are many limitations in that comparison as well since there were not many bottles and flushing is likely an issue for most of them.

For full details for the COMPARE run see file 2015-03-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The scatter is higher than usual though there is a fairly clear “average” fit. When outliers are excluded based on residuals the fit with offset forced to 0 has a slope of 1.0273. However, that fit does not look good with few points falling on the trendline. When there are no bottle values <2.6mL/L a zero offset would usually be chosen, but for some cruises, especially in inland waters, it often looks necessary to allow an offset. When that is done for these data and outliers are selected based on residuals, the fit is:

CTD DO Corrected = CTD DO * 1.0147 + 0.0647

For the last cruise on which this sensor was used the fit was:

CTD DO Corrected = CTD DO * 1.0187 + 0.0569 (2015-01)

That cruise had many deep casts with many low DO samples and good flushing, so the fits are reasonably close.

The minor outliers were examined and those that are above the trendline are mostly associated with fairly high DO gradients. In those areas the effects of poor flushing would be larger. The CTD almost always reads lower than bottles, but if the bottles contain water from deeper in the water column that would bring the 2 into better correspondence. A rough fit to some of those bottles showed a slope of ~1.01. A plot of differences for event #1 which had very well mixed waters shows a fit with slope 1.03. So it is possible the correction is underestimated when flushing is incomplete, though offsets in fits with poor flushing tend to be higher. The offsets may reflect flushing.

Since the fit for this cruise is close to that for 2015-01 but 2015-01 was very likely to have had better flushing of bottles and sampled a wider range of values, using that fit may be the appropriate choice for 2015-03.

The most severe outlier came from the surface and the standard deviation in the CTD DO was high. There is no reason to suspect the analysis or sampling. None of the other outliers are large enough to justify assigning quality flags in light of concerns about flushing.

Plots were made of CTD DO and bottle DO versus CTD salinity and no significant outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor. Plots of the ratio Fluor/CHL versus CHL show the usual shape for this type of fluorometer. For very low CHL the fluorescence was high by about a factor of 2 and between 0.3<CHL<2ug/L the ratio gradually falls from about 1.3 to 0.8. There were a few outliers with a ratio ~0.6 when CHL~1ug/L; those were from casts in Johnstone Strait and Hecate Strait. There are only 2 bottles with CHL >2ug/L and they are both ~2.75ug/L. For those the fluorescence is only 0.3*CHL for one and ~1*CHL for the other. The latter case has a high standard deviation in the fluorescence, so may not be reliable.

For full details of the comparison see file 2015-03-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

An initial estimate was made that the DO leads the temperature traces by about 2.5s.  ALIGNCTD was run using a setting of +2.5s to advance the DO signal. The results look good for some features, not for others. Tests using +3s and +3.5s show that +2.5 is generally better.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. The default setting of (α = 0.0245, β=9.5) has been found appropriate for other cruises using this equipment. Tests on one cast showed the default setting looked reasonable, though there was little difference among the choices tested.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. While none are very deep, it does allow us to check for sudden changes.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-50-0019 | 330 | ~0 N | -0.00008N | -0.0006 N | Mod, V.Noisy |
| 2014-50-0037 | 325 | ~0 VN | -0.00004 | -0.0004 | Mod, Steady |
| 2015-01-0040 | 325 | -0.0027 | -0.00002 | +0.002 XN | High, Very Noisy |
| “ | 1000 | -0.0013 | -0.00013 | -0.0001 | “ |
| “ | 3000 | -0.0007 | -0.00013 | -0.0011 | “ |
| 2015-01-0093 | 325 | -0.0008 | -0.00011 | -0.0005 | High, Very Noisy |
| “ | 1000 | -0.0008 | -0.00015 | -0.0010 | “ |
| “ | 3000 | -0.0006 | -0.00016 | -0.0014 | “ |
| 2015-03-0017 | 600 | +0.0001 | +0.00003 | +0.0004 | Moderate, steady |
| 2015-03-0027 | 450 | +0.0004 | +0.00005 | +0.0003 | “ |
| 2015-03-0075 | 350 | +0.0004 | +0.00004 | +0.0003 | “ |
| 2015-03-0087 | 450 | +0.0003 | +0.00004 | +0.0002 | “ |
| 2015-17-0058 | 325 | +0.0008 | +0.00005 | -0.0001 | High, Steady |
| 2015-17-0068 | 300 | +0.0006 | +0.00004 | +0.0001 | High, F.Noisy |
| 2015-17-0083 | 300 | +0.0010 | +0.00006 | -0.0002 | High, Steady |

The conductivity and salinity differences are very small but quite noisy; the values shown are rough averages for a particular depth. The results for 2015-03 are different from the observations of 2014-50 and 2015-01, but fairly close to 2015-17. The differences are slightly smaller than those found in the bottle comparison, but the values in the table are from the downcast, while the bottle comparison is gathered while stopped so alignment may be an issue.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book and there were a few minor discrepancies between the log and headers, but since the log entry formats varied from one watch keeper to another it is not clear what is correct. No changes were made.

A header check was run. The only problem found is that there are still negative pressures as low as -0.58 even after the correction to the offset in the pressure calibration.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 0.85db which is lower than usual, but this cruise was in quiet waters and there was a desire to measure near the surface, so it may be ok. So given the sensor is considered ±1db, a few negative values are not necessarily a problem. But this at least confirms that the addition of 1.25db to the pressure offset was justified.

The altimeter and water depth readings from the headers of the CLN and SAMAVG were exported to a spreadsheet.

- Plots were made of altimetry near the bottom for most casts and the headers look appropriate. Where there were no entries it is usually because the CTD did not get within 15m of the bottom, but in some cases it probably did but the altimetry was spiky and no estimate was made.

- Some water depths differed from the log entries. A decision was made based on which entry made most sense in light of the altimetry. Changes were made to 5 casts and the missing entry for cast #1 was added. Water depths are bound to be approximate due to ship movement.

- Cast #21 has almost no data; it was not a CTD or ROS cast, so should not be processed further.

- For the bottle files there were only 2 water depths that needed adjusting or adding.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and overall it did. No further was applied to the DO concentration channel,

Conductivity

Tests were run on 4 casts using a variety of shifts and the best results were found with -0.6 records for the primary and -0.4 records for the secondary.

SHIFT was run on all casts using -0.6 records for the primary conductivity and -0.4 records for the secondary conductivity.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

Conductivity, pressure and dissolved oxygen sensors were all recalibrated in late 2013 or early 2014. They were used for 1 cast during 2014-19 and all of 2014-50 and 2015-01. Both T/C sensor pairs produced salinity within 0.001 of bottles for 2014-50, while from 2015-01 it appeared that the primary salinity was high by ~0.003 and the secondary high by ~0.0015.

The pressure sensor was found to have drifted lower by about 1.25db during 2014-50 and 2015-01 and an initial examination of 2015-17 which followed this cruise shows pressures too low as well.

For 2014-50 the dissolved oxygen sensor was corrected using a linear fit of slope 1.0281 since there was too little sampling of waters with low DO values to estimate an offset. For 2015-01 the fit used had a slope of 1.0187 and offset of +0.056mL/L.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Local climatology was only available for a few of the open water casts. For those the temperatures were high between 60 and 100m. High temperatures have been widely reported off the west coast of B.C. and the wider north-east Pacific, so these excursions from the historic ranges are considered real. Salinity was within the climatology. There is no evidence of calibration problems.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available, but the secondary pump was found to be in need of repair after 2015-17.

SUMMARY OF SALINITY COMPARE RESULTS FOR CTD #506

| Cruise | Sal0-Bottle | Sal1-Bottle | Comments |
| --- | --- | --- | --- |
| 2014-50 | +0.0003 | -0.0004 | Flushing likely fairly good; limited sampling; quick analysis |
| 2015-01 | +0.003 | +0.0015 | Flushing likely good; many deep samples; quick analysis |
| 2015-03 | -0.0036 | -0.0029 | Flushing likely poor; problem found with secondary pump; analysis delayed |
| 2015-17 | +0.0007 | +0.0012 | Flushing likely fairly poor for most casts; limited sampling; quick analysis |

Something happened between cruises 2015-01 and 2015-03 that led to the secondary sensor reading higher than the primary sensor. This could be due to the malfunctioning of the secondary pump. While this would not likely have a large effect on the salinity while stopped, it might lead to slightly higher secondary salinity if full equilibration had not been reached. It is also possible that one or more sensor calibrations drifted.

##### DETAILED EDITING

A first attempt at editing the secondary temperature and salinity was abandoned when news was received that there was a problem with the secondary pump. Many casts were examined to see if there was a problem with the secondary data and the evidence included:

- Secondary pair have smoother T-S curves - descent rate was steady.

- Repairs were done to the secondary pump after 2015-17.

- Odd behaviour in secondary conductivity during 2015-03 for casts with very low salinity at the surface – keeps low value for a while then “catches up” with different shape from T. Primary salinity has shape similar to both T sensors. Similar behaviour is seen during bottle stops when a shed wake passes through – the secondary conductivity is slow to recover while the primary snaps back pretty quickly.

- In COMPARE the primary CTD salinity has higher standard deviations. This could mean the secondary is not sensing variations quickly.

The primary sensors were chosen for editing and eventual archiving because of the problems with the secondary pump.

The DEL files for many of the casts include an opening section for which the pumps were not on.

CTDEDIT was used to remove sections with the pumps off and to remove or clean spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes mostly near the top and bottom of the casts. All files required some editing.

A return was made to this stage later to fix file #43 which had one very shallow record left with pumps off.  A text editor was used to remove 1 record from the EDU file and CLEAN was run to fix the headers with output EDT.

##### Initial Recalibration

The best evidence available concerning salinity calibration is from the 2015-01 comparison with bottles and that suggests that the primary salinity is high by about 0.003; the 22 bottles fired at 2000m were the most convincing evidence. There was some hint from 2015-01 that the differences were a little lower near the end of the cruise but it was not strong enough to consider a time-dependent recalibration. The few bottles considered most reliable from the 2015-17 comparison also suggest a slightly lower correction for the primary, but even slight problems with flushing could explain the difference in light of the local salinity gradients.

CALIBRATE was run using file 2015-03-recal1.ccf to subtract 0.003 from the primary salinity and to correct the Oxygen:Dissolved:SBE channel in the SAM and MRGCLN2 files using:

CTD DO Corrected = CTD DO * 1.0187 + 0.0569

COMPARE was rerun for salinity and dissolved oxygen using the recalibrated values. The results confirm that the recalibrations were applied properly. The salinity comparison looks worse, as expected.

The average of differences in the DO fit once the same outliers were removed as in the previous comparison was -0.0126mL/L. The differences are greatest for high DO. This could be due to incomplete flushing off bottles since that effect is greatest in the high DO gradients which tend to be at the upper end. There is also a possibility that the SBE dissolved oxygen data are slow to respond due to the problem with the secondary pump. Neither the bottle data nor the sensor data are as reliable as usual. (See file 2015-03-DO-comp2.xlsx for details.)

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data. For this cruise we expect the downcast DO to read higher due to inefficient flushing.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed, the CTD DO was higher than the bottles by an average of ~0.023mL/L and standard deviation of 0.035mL/L; the differences vary with depth with even deep DO values generally higher than the bottles. This suggests that flushing is an issue even at depth. The comparison does not support further recalibration. See 2015-03-dox-comp3.xlsx for details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some unstable features, but that is normal for this region in near-shore and well-mixed casts. No problems were found.

Profile plots look ok except that the PAR:Reference channel appears to contain no signal.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, PAR:Reference, Altimeter, Status:Pump, Descent_Rate and Flag.

(There was a constant signal from the PAR:Reference with a very low value for the downcast of event #1, but it dropped to 0 during the upcast and had no signal for the rest of the cruise.)

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and unedited except that

some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

The Oxygen:Dissolved:SBE data are considered of lower quality than usual due to

likely incomplete flushing of Niskin bottles and a problem with the secondary

pump that may cause poor response of the SBE sensor. Re-calibration was based

on the results of cruise 2015-01 when the same sensors were used, but there

remain some doubts due to the pump problem. These data should be used with

caution especially in areas of large vertical DO gradients.

Salinity:T0:C0 data were also recalibrated based on 2015-01 results due to

concerns about the flushing of the Niskin bottles during this cruise. While

this is the best available comparison, it included a lot of scatter.

For details on the processing see the report: 2015-03_Processing_Report.doc.

----------------------------------------------------------------------------------------------

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values varied from about 50% to 105%. The lowest values were in Gardner Canal and in the well-mixed waters of Johnstone Strait. Along Douglas Channel the surface saturation falls steadily from about 101% at the head to 90% near the open waters. Most values were between 90% and 98%. The highest values were for events #5 and #29. The former was the most exposed site where we would expect values of at least 100% and the latter is one with a fairly high surface DO gradient where we also expect higher saturation rates. These 2 cases suggest that the DO values are reasonable, but higher values would also be reasonable.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, PAR:Reference, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check and a header check were run on all files and no errors were found.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

August 12, 2015 – Based on information from C. Wright, changed station names for events 49 & 51 from DC-26 to DeC-26 & DC-31 to GC-31.

# Particulars

1. There was confusion about bottle firing, but notes on the rosette log sheet make it clear what happened.

5. Deck pressure -0.47db, surface press=1.0. First attempt pumps were off, so redid it overwriting 1st file.

6. Labels wrong – should be HEC1, not H-Deep. According to the log Bottle 9 did not fire; in fact it did fire, but did not close. All chemistry for 30m was taken from bottle #10.

10. Bottle 9 trigger cleaned - appears to have fixed problem. This was not a rosette cast – firings just to work on trigger.

14. Sample #s mistakenly skipped – numbers not used during cruise.

19 & 20. Log note that Oxygen trace looked noisy. Seen in other casts as well but also in 2015-01.

21. Contains only a little data –not a CTD cast at all, so presumed accidental acquisition.

29. Winch started to make rattling sound while under load. Altimetry bad.

32. Winch noise worse.

33. New winch used.

42. Pumps off – return to surface for restart.

49. First drop with syringes on – redid.

62. Accidentally labelled as 74. So event 74 was skipped to avoid overwrite.

77. Stopped while CTD at 200m to reposition ship in strong current.

82. Bottom lanyard/end cap closed during cast but bottle itself didn’t trip. No sampling.

87. Niskin 3 fired late and on the move.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 31Jan2013 | Factory |  |  |
| Conductivity | 1763 | 1Jan2014 | Factory |  |  |
| Secondary Temp. | 5013 | 27Feb2013 | Factory |  |  |
| Secondary Cond. | 3394 | 3Jan2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 1438 | 3Jan2014 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| SeaPoint Fluor. | 3642 | n/a |  |  |  |
| Pressure Sensor | 506 | 30Dec2013 | Factory |  |  |
| Altimeter | 1204 | 1252 |  |  |  |