## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 27 Mar 2025 | ### Updated channel names & formats in TOB & Loop files. |
| ### 1 December 2021 | ### Corrected Salinity:Bottle precision lost during the addition of HPLC, added DIC, TA, and pH to casts 5 & 63, dsmp comments to #23. S.H> |
| ### 12 Aug 2021 | ### Added DIC, TA and pH data. SH. |
| ### 1 April 2021 | ### Added dmsp analysis methods to comments for event 23. SH |
| ### 11 Dec 2020 | ### Re-ran addition of HPLC data. S.H. |
| 16 July 2020 | Added HPLC data & corrected truncated flag channel name for DMSP_D |
| 19 June 2019 | Added DMSP-D and DMSP-T – only pad values due to contamination. |

## PROCESSING NOTES

Cruise: 2015-09

Agency: OSD

Location: North-East Pacific

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: 7 June 2015 – 23 June 2015

Processed by: Germaine Gatien

Date of Processing: 30 July 2015 – 27 October 2015

Number of original HEX files:	71	Number of CTD files: 69 (1 cast aborted /1 split cast)

Number of bottle files: 	57		Number of bottle casts processed: 57

Number of original TSG files:  	7	Number of processed TSG files:	7

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0506) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#1438) on the primary pump, a SeaPoint Fluorometer (#3640) on the secondary pump, a Biospherical QSP-400 PAR sensor (#4565), and an altimeter (#62354).

A thermosalinograph (Seacat 21 S/N 3363) was mounted with a SeaPoint fluorometer (#2356), remote temperature sensor and a flow meter.

The data logging computer was #2. Seasave version V7 22.4 was used for acquisition.

The deck unit was a Seabird model 11, serial number 0508.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

The oxygen kit was IOS kit #2.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log, rosette log sheets and analysis logs were in good order.

Sampling notes from the chief scientist were very useful.

Event #48 was aborted due to a computer crash and a full cast was run as event #49.

Event #89 was interrupted by a computer crash during the upcast, so a 2nd file was started with the event #90. No bottles were sampled from the upcast portion of the first file and the 2nd file was renamed as 2015-09-0089.CHE to match the downcast file.

The salinity data were corrupted by excursions that are not seen in the temperature channels; both salinity channels were affected but in opposite directions. These excursions are usually seen where the CTD is decelerating. This is assumed to be due to varying alignment perhaps due to small flow rate variations. The excursions are most significant in areas where there are large temperature gradients combined with low salinity gradients, but are not seen in the shallow casts shoreward of P4. This is likely because interleaving leads to frequent areas where there is little variation in salinity while temperature is decreasing. Sometimes salinity data were removed from such excursions but temperature and other variables were left in the record. The errors that remain due to this problem are expected to be small.

The comparison between salinity bottles and CTD salinity had a little more scatter than expected. The analysis precision was a little poorer than usual and there is some evidence that there may have been intermittent analysis problems. Yet, tests of the salinometer turned up no problems. The analysis was done within 13-37 days of collection and plastic inserts were used in the bottles, so we would not expect significant evaporation. When the bottle comparison was used to compare the results from the two days on which analysis was done, the results from the later analysis show less noise, further indication that evaporation was not large. There are reports in the literature of salinity increasing in soft glass (not tempered) bottles at a rate of 0.0012 per month during the first 6 months. The scatter is likely due to combined errors from the salinometer, sample collection, sample storage and noise in the CTD data.

Prompt salinity analysis and continued monitoring of the salinometer performance are recommended.

The primary salinity was found to be high by an average of 0.0005 but the standard deviation is 0.0016, so the range is -0.001 to +0.002. If evaporation is an issue then the most reliable estimate would be near the top of that range. Since all these estimates are within ±0.002 and there is no reliable way to determine the cause of the scatter, the salinity data were not recalibrated.

The CTD fluorescence values compare well with the extracted chlorophyll overall, except for low CHL values (<1.5ug/L) when the fluorometer generally reads too high.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 150db

±0.2 mL/L from 150db to 400db

±0.08 mL/L from 400db to 600db

±0.04 mL/L below 600db

The Thermosalinograph performed well with few large spikes in salinity. TSG Salinity was found to read lower than CTD and loop salinity by about 0.03 in the off-shore area, and by more than that inshore of station P6. These differences are likely due to a combination of calibration drift and the effect of bubbles. Salinity was recalibrated by adding 0.03 which looks suitable for the most of the cruise, but is likely an under-correction for near-shore areas. The fluorometer malfunctioned and recorded unbelievably high values for most of the cruise; that channel was removed from all files.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

There were sampling notes provided by the chief scientist.

Nutrients, extracted chlorophyll, dissolved oxygen, DMS and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked.

The XMLCON files did not change through the cruise.

The calibration constants were checked for all instruments and the only correction needed was to fix a spelling error. The corrected file was saved as 2015-09-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2015-09-ctd.xmlcon to create CNV files.

There were 2 casts that had computer crashes.

- For cast #48 the cast was aborted at about 500m. The CTD was returned to the surface and a full cast run as event #49. The data from cast #48 looks ok above 300m but bad below that. It will not be processed. A preliminary derivation of salinity shows that while both casts had many spiky features the ones from #48 were often unstable in T-S space. File #48 will not be processed.

- File #89 crashed during the upcast before any bottles had been fired. A new file, #90, was started and bottles were all fired during that cast. This file was renamed as #89 to match the downcast file.

A few casts were examined and all expected channels are present.

The temperature and conductivity channels are close at depth. Near the surface there is a lot of variability especially during upcasts.

Early in the cruise the descent rate was often extremely noisy with complete reversals of the CTD. Later in the cruise the descent rate is steadier.

The altimetry has frequent spikes at the bottom but an estimate is usually possible.

Fluorescence, transmissivity and PAR look ok.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2015-09-ctd.xmlcon.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. There were many more than usual. There appears to be more motion during stops than usual and where that occurs in fairly high gradients, more outliers are seen. We usually just see a few distinct outliers that are easily removed, but this won’t be practical here. So no attempt will be made to edit these data.

A preliminary header check and a cross-reference check were run and no problems were found. CTD fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

Sort was used to ensure that the ADDSAMP file was in sample number order.

One bottle was removed from the list where there was no sampling and no sample number.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

The SAM files were bin-averaged on bottle # and called SAMAVG.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2015-09-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2015-09chl*.xls. The file included comments and flags and a precision study. The loop samples were removed to a separate spreadsheet, 2015-09_loop.csv. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2015-09chl.csv. The csv file was then converted to individual CHL files.  Note that loop sampling was complex with some samples having been through the filter and others not, as efforts were made to track a problem with the TSG fluorescence.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2015-09oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2015-09oxy.csv.

That file was converted into individual *.OXY files.

SALINITY make sure cast 89/90 is identified as 89

Salinity analysis was obtained in 2015-09SAL.xls. The event number 90 was changed to 89 as this was a split cast which will be named #89. The analysis was done between 15 and 42 days of collection. The files were simplified and saved as 2015-09sal.csv. The loop samples were removed to a separate spreadsheet, 2015-09_loop_data.csv. File 2015-09sal.csv was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2015-09_nutrients*.xls. This includes a precision study. The loop samples were removed to a separate spreadsheet, 2015-09_loop.csv. Event #90 was changed to (#89 to match the downcast file (files split due to computer crash.)Then the file was simplified, reordered on sample numbers and saved as 2015-09-nuts.csv. The file was converted to individual NUT files.

DMS

DMS data were obtained in file DMS summary (2015-09).xls. Values given as < were changed to 0 and the comments that will go into the header will explain that 0 means below detectable level. (There was a separate report on analysis techniques and problems.) A comment was added to one sample that had been flagged 5; there was a relevant comment on the raw data page. The file was then saved as 2015-09DMS.csv and converted to individual DMS files.

The SAL, CHL, OXY, NUT and DMS files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. Errors found or things to be adjusted were:

- 51. DMS was missing – had been identified as event #57 – understandable as that is what it looks like on the sampling log sheet. Event 57 is at P21, not P20.

- 89. This was missing from the list but the file was found and was correct. The list was corrected.

- 106. There are 2 CHL samples from the surface but only 1 shown on Sampling Log. One of the samples looks wrong – assume this was mis-labelled and should be 108

- 108. There are 2 CHL samples on the sampling log but only 1 sample. The extra sample from 106 belongs in 108.

- 116. Only Niskin 1 was sampled. Niskin 2 has no sample number so was removed from the SAMAVG file.

After these corrections the merge process was repeated and CLEAN was rerun.

##### Compare

Salinity

Compare was run with pressure as reference channel.

When bottles above 120m are excluded as well as a few outliers based on residuals the primary CTD salinity was found to be high by an average of 0.0005 and the secondary to be low by an average of 0.0014; the standard deviations were 0.0014 and 0.0015, respectively.

The analyst reported poor precision though it doesn’t seem particularly bad with Sp = 0.0012 after 1 outlier was removed (based on the Chauvenet criterion). This is near top of the range of Sp values found for this project in the past, but not high enough to limit the number of digits reported. Analysis was done within 6 weeks and plastic liners were used in the bottles so we do not expect a lot of evaporation. Evaporation could explain the scatter in the comparison and precision study. Ky and Hinrichsen (1986) reported that salinity increases in “soft” glass bottles (i.e. not tempered glass) at a rate of ~0.0012 per month for the first 6 months of storage. If evaporation or desorption was significant then the bottle values would be too high so that the CTD would actually be reading higher than the comparison shows. For the preceding cruise, 2015-20, there was poor precision in analysis, more scatter in the CTD-bottle comparison and a longer delay before analysis. Yet the comparison between CTD and bottles produced very similar results for average differences to the one for this cruise. Evaporation cannot explain that unless this is just a chance result.

One bottle from the previous cruise was found to have an odd film on the inside of the bottle. Might there have been similar effects on other bottles that were not so well developed that they were noticed? The analyst was looking for that and found no evidence of any other bottles containing a film. The analyst reported stability seemed fine during analysis, though comparison of substandard results over a few weeks shows more variability than expected.

There were 22 bottles fired at 2000m during cast #63. Pressure varied by ±0.5db during the bottle firing period, which corresponds to salinity variability of about ±0.0002 based on the downcast gradient just above 2000m. There were a few salinity spikes during that time so that the range of values in the SAM files is ~0.0018, but after averaging it is 0.0002. So variability is low in the CTD salinity. The bottles show much higher variability, with a range of 34.5801 to 34.5862 and standard deviation of 0.0015.  Precision in the analysis was considered a little worse than usual and may account for the scatter in the differences between bottles and CTD salinity. However, reasons for the lower precision have not been found. The salinometer was tested and performed well and the analyst did not notice problems with sample bottles and inserts.

The average differences at 2000m were:

| Sensor Pair | Average (CTD-Bottle) | Std. Dev. (CTD-Bottle) |
| --- | --- | --- |
| Primary | 0.0000 | 0.0015 |
| Secondary | -0.0015 | 0.0015 |

These differences are slightly smaller, but close to, those found for the whole data set. The differences between primary and secondary CTD salinity, as deduced from these, is about 0.0016, which is close to those seen in the deep parts of the downcasts as described in section 9.

If evaporation or desorption caused the scatter in the bottle comparison, then an analysis with time as the significant factor might make this clear. Analysis was done on 2 days: July 6 and July 20. Samples were collected between June 8 and June 21, so we have a storage time of from 15 days to 42 days.  Evidence examined was based on primary CTD salinity and included:

- Most of the samples collected from the first 5 casts were analyzed on the later date, so those were examined to see if they stood out in the comparison. Only cast #10 has bottles deep enough to be reliable in a comparison. The CTD was found to read higher than the bottles by 0.0003, 0.0005 and 0.0019 for 3 bottles and low by 0.0014 for the 4th; the average is close to that for all data.

- The average is very close to the average for the whole data set. Most of the samples collected late in the cruise were analyzed quickly; again, many of those are shallow, but the deeper samples mostly seem close to the average.

- 9 of the 22 bottles fired at 2000m of a single cast were analyzed on July 20th and 13 on July 6th. The average difference for those analyzed quickly showed the CTD to be low by 0.0007 and for those analyzed later it was high by an average of 0.0007. If we had evaporation we would expect the July 20th bottles to read higher and hence the CTD should look lower for the slower analysis. So this suggests that is not the problem. When the differences for these bottles were plotted against the order of analysis, it is quite clear that variability was much higher on July 20th (from 60 to 69) than on July 6th.

- COMPARE was run using sample number instead of pressure for cast #63 to see if that shed any further light on the issue. It makes it clear that both the CTD salinity values and their standard deviations were very stable. Samples #391, 403, 404 and 406 are out of line, with the 401-409 distribution looking markedly noisy. Samples 400 to 408 were run on July 20th. When arranged according to Sample ID (order of analysis) it is clear that there are very consistent results from #41 to #55 (July 20) with higher variability for #60-69 (July 6). Evaporation should affect the July 20 analysis more than July 6. The pattern is the same for the secondary CTD salinity.

- The CAL C samples run on the 2 days were analyzed and the standard deviation was higher on July 6th, but if we exclude 2 outliers with salinity higher than the average by ~0.003, the results are similar to July 20th both in average value and standard deviation. This suggests there might have been a sporadic problem on July 6th.

- Most of the duplicates were analyzed on July 20th. One was split between the 2 days. There are no outliers from July 6th but there were only 4 samples. There are 12 samples from July 20th and a few outliers and mostly differences are higher than on July 6. This suggests there might be a problem on July 20th, but there are too few samples from July 6th to say it was better.

- There was one case where duplicates were analyzed on different days with results 34.3629 and 34.3654 on July 6 and July 20, respectively. This might suggest evaporation of sample. When compared with the CTD salinity (34.3646) the July 6 sample is lower than the CTD by 0.0017 while the other is higher by 0.0008. This is insufficient evidence to conclude there is evaporation, but if there is, it is probably not significant.

- Finally, the comparison of bottles and CTD versus pressure was run so that all samples analyzed early were a different colour from those run late. Since some that were analyzed late were also collected late, this is a very rough way to judge if there had been a change. The overall impression is that there is no pattern except that perhaps the later analysis has less scatter than the early one.

If there was evaporation or desorption of samples it seems to have been slight and there is no evidence that samples that waited longest for analysis stand out from the rest. There was no evidence of poor seals. If the bottle values are a little higher than they should be, then the CTD salinity might be somewhat higher than indicated by the comparison, but not by much since many of the samples were analyzed in less than a month. So the primary salinity may be high by roughly 0.001, but given the weak evidence, no recalibration is justified.

Was there a significant problem with analysis? There is some evidence of a problem on July 20th from the 2000m bottles and from the duplicate analysis though there were very few of those from July 6th. But in the CAL C study only the July 6th data have significant outliers. We have to expect some noise from the salinometer, but there does seem to be some indication of a few bottles that are out of line occurring close together. There may be an intermittent problem, but it is not large.

There was only 1 outlier with a difference > 0.1; this came from event #9, sample #36. There is no such sample noted on the sampling log. The sample looks more like it came from 5m. This is likely a case of mis-labelling. Since the source is in doubt it was flagged ”5” with a pad value.

Other outliers with differences >0.006 were examined. Most came from near the surface where large gradients would mean that any inefficiency in Niskin bottle flushing would lead to large mismatches with the CTD. But one outlier was not easily explained:

- Sample 289, event #41 – There were two bottles at 3500m. One was higher than the CTD salinity by about 0.005 which is a bit off, but not terrible; the other was higher by about 0.014, so definitely an outlier. Both CTD salinity channels had very close values for both bottles. The dissolved oxygen samples from the same 2 bottles compared very well with each other, so there does not appear to be a problem with early firing of a Niskin bottle. There is no CTD salinity that high in the full cast, so that is not a good explanation anyway. It is possible there was some evaporation. The sample was flagged 4.

For full details for the COMPARE run see file 2015-09-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There was only one significant outlier, with most of the data falling into a reasonably tight group.

When outliers were removed based on residuals the fit was:

CTD DO Corrected = CTD DO * 1.0246 + 0.0452

To test for hysteresis points below 1200m were also excluded and the fit was then:

CTD DO Corrected = CTD DO * 1.0248 + 0.0385

The corrections using the 2 fits are similar with differences ~0.006 at DO=0mL/L and ~0.001 at 6mL/L. The hysteresis parameters, E, H1 and H3 do not need fine tuning.

There was only one major outlier, sample #589 from event #108. There was very high variability at the beginning of this bottle stop, so this is quite likely a matter of the Niskin being filled by a shed wake and never equilibrating to ambient values. So the value likely reflects the bottle contents; no flag is justified.

Plots of Titrated DO and CTD DO against CTD salinity were examined. The only outlier was that from sample #589.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor.

There was a wide range of CHL values from ~0 to 22.6ug/L. As usual, fluorescence generally reads much higher than CHL for low CHL, in this case by as much as 9 times. Incomplete flushing of Niskin bottles likely explains much of the scatter in the fits for higher CHL values, as the sample may reflect conditions a short distance below the CTD depth. Since fluorescence usually has a fairly sharp peak a sample that comes from only a short distance below the CTD observation can have a much higher or lower value.

For CHL>1.5ug/L the ratio of FL/CHL is about 1. This is a better correspondence than we usually see.

A fit forced through the origin has a slope of 0.84.

For full details of the comparison see file 2015-09-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

There are so many bottle stops in these data and shed wakes due to noisy descent and ascent rates that it is hard to judge what setting brings the SBE:DO trace into best alignment with temperature. Settings between +2.5s and +4.5s all look reasonable. For 3 other recent cruises using this equipment an advance of +2.5s looked best; however, +3s looks better for these data.

ALIGNCTD was run on all casts using +3s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. During other recent cruises using the same equipment, the default setting of (α = 0.0245, β=9.5) looked slightly better for both conductivity channels than others tested. One cast was checked for this cruise and the default setting does improve the data.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. A few results from an earlier cruise (shaded entries) are shown for comparison, including 2 cruises that did not go very deep.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2015-01-0040 | 1000 | -0.0013 | -0.0001 | -0.0001 | “ |
| “ | 3000 | -0.0007 | -0.0001 | -0.0011 | “ |
| 2015-01-0093 | 1000 | -0.0008 | -0.0002 | -0.0010 | “ |
| “ | 3000 | -0.0006 | -0.0002 | -0.0014 | “ |
| 2015-03-0017 | 600 | +0.0001 | +0.0000 | +0.0004 | Moderate, steady |
| 2015-17-0058 | 325 | +0.0008 | +0.0001 | -0.0001 | High, Steady |
| 2015-09-0018 | 1000 | -0.0009 | -0.0001 | -0.0006 | High, V.Noisy |
| “ | 1900 | -0.0007 | -0.0002 | -0.0016 | “ |
| 2015-09-0041 | 1000 | -0.0015 | -0.0002 | -0.0010 | High, X Noisy |
| “ | 1900 | -0.0009 | -0.0002 | -0.0016 | “ |
| “ | 3000 | -0.0009 | -0.0002 | -0.0021 | “ |
| 2015-09-0078 | 1000 | -0.0011 | -0.0001 | -0.0007 | High, Moderate |
| “ | 1900 | -0.0007 | -0.0002 | -0.0014 | “ |
| “ | 3000 | -0.0005 | -0.0002 | -0.0018 | “ |
| “ | 4000 | -0.0005 | -0.0002 | -0.0015 | “ |
| 2015-09-0087 | 1000 | -0.0009 | -0.0001 | -0.0008 | High, Moderate |
| “ | 1900 | -0.0007 | -0.0002 | -0.0014 | “ |
| “ | 3000 | -0.0007 | -0.0002 | -0.0014 | “ |
| “ | 4000 | -0.0006 | -0.0002 | -0.0015 | “ |
| 2015-09-0104 | 1000 | -0.0007 | -0.0001 | +0.0006 | High, Moderate |
| “ | 1900 | -0.0012 | -0.0002 | -0.0008 | “ |

The differences are small and show no significant change through the cruise; they are a little lower later in the cruise probably due to a steadier descent rate. There is also no significant drift over several cruises. In the deeper waters of the current cruise the salinity differences show pressure dependence that disappears at about 3000m, at which point low gradients make any misalignment effects insignificant.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book and two station names were fixed as noted in the sampling notes.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.9db which is reasonable for the Tully, though in offshore waters we might expect it to be a little higher. The lowest pressure recorded was +0.07db at the beginning of a file; there were in-water salinity and temperature values but the pumps were off.

A header check was run. There were many negative values that were investigated to see if they are single-point spikes or if there may be a calibration problem:

- Fluorescence values were very high during the upcast of event #18 at about 70m. There were about 200scans full of spikes to high values but not off-scale. This was during a stop for a bottle, but does not occur during the 10s-window of data included in the bottle file.

- Oxygen:Dissolved:SBE was negative for about 100 scans during the upcast of event #89. There was no stop during that period so no bottle data are affected.

- Both conductivity channels had spikes just below 300m during the downcast of event #48 and the pumps were turned off at that point. This cast is not to be processed further. #49 replaces it.

No other problems were noted.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. Plots were made of altimetry near the bottom for most casts and the headers look appropriate for all files except for casts #38 and 45 in the CLN files and #20, 22, 49 and 79 in the SAMAVG files. Those altimeter headers were removed.

Water depths differed from those in the log in many cases. When the difference was >3m, the entry was changed in the CLN and SAMAVG files. Changes were made to events #:10, 13, 15, 17, 19, 65, 77, 84, 85, 86, 87.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 4 casts to choose settings for alignment of the 2 conductivity channels. For the primary conductivity the best results were found with -0.9 records. For the secondary the best results were with a setting of +1.0 records, but the secondary data look very noisy with all settings.

SHIFT was run on all casts using -0.9 records for the primary conductivity and +1.0 records for the secondary conductivity.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning concerned the fragment event #90 from the upcast after a computer crash.

##### Other Comparisons

Previous experience with these sensors –

Conductivity, pressure and dissolved oxygen sensors have been used on 6 previous cruises since they were last recalibrated in late 2013 or early 2014. There was no sampling from the 1st and poor comparisons for the 3rd, 4th and 5th. For the 2nd cruise both T/C sensor pairs produced salinity within 0.001 of bottles. Problems were found in the secondary pump after the 4th cruise. During 2015-20 there were further problems with the secondary pump or some other element of the secondary system that rendered much of the secondary data unusable. The primary salinity was found to be close to bottles, on average, but there was a lot of scatter in the comparison and the analyst noted poor precision in the bottle duplicates. The secondary salinity was found to be low by about 0.0015.

The pressure sensor was found to have drifted lower by about 1.25db during the first 3 of these cruises so the offset was increased by that amount after that.

The dissolved oxygen sensor corrections from the previous cruises varied from about 1.5% to 3% with no obvious drift pattern.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Salinity values were within the climatology except for some values below the minimum around 400m for casts at P12 and P16 and some high values around 100-140m at P20 and P21. Temperatures were above the historic ranges at many casts between P4 and P12 around 80m. There was a lot of variability among the many casts at P4 but the temperature was above the maximum for some of the many P4 casts on the outward journey and for the one sampled on the inward journey. Temperatures were below the minimum at P21 around 250m. It is unusual to have excursions from the climatology out as far as P16, but this fits the observations of the North Pacific “blob” as reported widely in 2014/2015.

Repeat Casts – There are many repeat casts but most include a deep and a shallow cast, so those are not good for checking repeatability. At P26 there were 2 very deep casts just over 24 hours apart. At about 4000m the temperatures vary by about 0.002C° and salinity by 0.0002 along lines of constant sigma-T. At this level we expect very little change and the CTD does well on this test of repeatability.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

Since the secondary temperature and salinity channels had some serious problems during 2015-20, the primary channels were selected for archiving, and hence, editing. The primary channels were also closest to the bottles.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. There were some excursions in salinity not seen in the temperature channel. These are mostly associated with deceleration of the CTD through temperature gradients. This appears to be due to misalignment of T and S since the excursions are frequently in opposite directions in the two salinity channels, perhaps due to variations in the flow rate or swinging of the CTD. In such areas salinity points were sometimes interpolated but many such points were removed, but temperature data were left in place.

All files required some editing.

##### Further Recalibration

No adjustment appears necessary for the pressure sensor.

The primary salinity was found to be high by about 0.0005 based on the bottle comparison and by 0.0001 based on the 22 bottles fired at 2000m during a single cast. If the scatter in the comparison is due to evaporation or desorption of samples, then the CTD salinity would likely be reading higher than it appears, though many samples were analyzed quickly enough to make a significant effect unlikely. No recalibration is justified.

File 2015-09-recal2.ccf was created to correct the Oxygen:Dissolved:SBE channel using:

CTD DO Corrected = CTD DO * 1.0246 + 0.0452

This correction was first applied to the SAM and MRGCLN2 files. COMPARE was rerun and the results confirm that the recalibrations were applied properly. The average of differences in the DO fit when outliers were removed was +0.0002mL/L and the standard deviation was 0.018mL/L.

(See file 2015-09-DO-comp2.xlsx for details.)

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. There is a lot of scatter but most of it is in the range 3.5 to 5.5mL/L.

When outliers were removed, the CTD DO was higher than the bottles by an average of ~0.013mL/L and standard deviation of 0.051mL/L. But if you only look at DO in the range 0-2mL/L and 6-7mL/L the downcast DO is very close to upcast bottles. Most outliers occur in areas where dissolved oxygen varies most rapidly, so even a slight inefficiency in Niskin flushing would lead to lower DO values in the bottles. That could explain why the CTD looks higher than bottles, though another explanation is that the CTD DO sensor is reading high because it cannot responded quickly enough. No further recalibration is justified. See 2015-09-dox-comp3.xlsx for details.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some unstable features in 3 casts but for cast #121 which was near shore the features could be real and for casts #10 and #38 the features are in areas with many corrupted records and only removing large sections of data will remove the unstable features. No further editing was applied.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

The PAR channel was removed from casts: #20-40, 46-67, 76-89, 93-115 and 121.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and unedited except that

some records were removed in editing temperature and salinity.

A graphical editor was used to remove records from obviously bad sections of

temperature and salinity data, and some salinity data were smoothed to

remove small spikes. Where there was corruption of salinity data that

was not seen in temperature data, only salinity points were removed.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

Downcast Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 150db

±0.2 mL/L from 150db to 400db

±0.08 mL/L from 400db to 600db

±0.04 mL/L below 600db

For details on the processing see document: 2015-09_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

A cross-reference list was produced.

The sensor history was updated.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted.

Values were high (110%-130%) at most near-shore sites (P2, P3, P4, LB01 and LB02) but only 90% at P1. In the offshore (P4 to P26 and LB04 and LB06) saturations were between 103% and 106%.

These values look reasonable though possibly slightly high as we typically see about 103% offshore. The casts don’t go right to the surface but temperature and DO look quite well mixed near the surface.

##### Final Bottle Files

The MRGCOR2 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

The PAR channel was removed from casts: #20-40, 46-67, 76-89, 93-115 and 121.

A second SBE DO channel was added for both the CTD DO and titrated DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

Data were exported from the CHE files to file 2015-09-bottles-final.xlsx. The entries were compared with the rosette log sheets to ensure no samples had been missed and the following discrepancies were found:

- Sample #56, Cast #10 - There is no salinity sample but it was shown on sampling log. There is no such sample is in the spreadsheet. It is assumed that it was never taken.

- Sample #138, Cast #18 – Duplicates were taken but were not averaged in the spreadsheet. The values were averaged and flag 6 added to the CHE file and the analyst was notified to update the spreadsheet.

- Sample #280, cast #38 - There is no salinity sample but it was shown on the sampling log. There is no such sample in the spreadsheet. It is assumed that it was never taken.

- Sample #333, cast #49 – Was entered in spreadsheet as event #48. Fixed file.

Standards check and a header check were run on all files and no errors were found.

A cross-reference listing was produced for the CHE files.

##### Thermosalinograph Data

There were loop nutrients, extracted chlorophyll and salinity samples taken, some while stopped and some while underway.

a.) Checking calibrations

The 7 configuration files identical. One file was renamed as 2015-09-tsg.xmlcon. No errors were found in the calibration parameters. There were errors in the names of the first 4 files with names starting 2015-20* instead of 2015-09*. The names were corrected.

b.) Conversion of Files

The 7 files were converted to CNV files using configuration file 2015-09-tsg.con.

They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots were produced.

- File #1 – flow starts ~0.95, drops to ~0.6 then goes up to ~1.2. The salinity is very noisy when the flow rate is low.

- File #2 – short record – flow rate ~1.2

- File #3 -  starts with flow ~1.2, then drops to ~1 and stays there

- File #4 and 5 – flow rate ~1

- File #6 – flow rate ~1 with drop to 0 at end

- File #7 – flow ~1.05 throughout

There are occasional spikes in salinity but they are few. File #1 is the only one with a significant spiking problem for about 9 hours when the flow rate was low.

It was noted at sea that the fluorescence values rose throughout the cruise and did not compare at all well with the CTD fluorescence near the surface. During file #7 it goes off-scale and stays there. Comparison with the CTD and loop samples may enable a decision on the value of data from the first file, but it is expected that the fluorescence channel will have to be removed from most of the data.

The track plot looks fine. The plot was added to the end of this report.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db and exported to a spreadsheet which was saved as 2015-09-ctd-tsg-comp.xls. There were 64 casts which overlapped with TSG files.

The TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, lab temperature, salinity and fluorescence and the files were reduced to the times of CTD files and loop samples.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude and longitude were all≤0.0010° and only one difference was >0.00023°; the medians in both were 0.0000°. This shows both the times and positions are reliable for both systems. For the one larger difference the longitude was varying more than usual at the time given for the CTD start which may indicate the ship had not quite stopped or was drifting.

The loop samples were combined in file 2015-09-loop-TSG-rosette-comp.xlsx. There were 5 cases where loop samples were taken during a rosette cast. The rosette salinity and CHL bottle data were combined with the loop sampling and median (over 5 records) TSG readings were added. The only nutrient sampling was from underway loop samples so there are no comparisons possible.

The two spreadsheets will be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Loop & Rosette Samples and TSG and CTD data

- T1 vs T2 The intake temperature sensor worked throughout the cruise. The median difference between intake temperature and lab temperature while the ship was stopped was 0.178C° with a much wider range than usual from -0.625C° to 1.171C°. This variability is mostly towards the end of the cruise when the ship ran near-shore lines through rapidly waters with rapidly-varying temperatures. When those data are excluded the median difference is similar and the value looks typical for the time of year when intake water temperatures are moderately high.

- TSG vs CTD The spreadsheet comparing CTD and TSG files was then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG.

1. Intake Temperature The intake temperature is higher than the CTD temperature by a median of 0.0055C° (standard deviation 0.289C°). When the near-shore data near the end of the cruise are excluded the median difference is 0.0037C° (standard deviation 0.006C°).

2. LAB TEMP The lab temperature was higher than the CTD temperature by a median value of 0.191 C° or 0.186C° if outliers are excluded.

3. SALINITY The TSG salinity data are lower than the CTD salinity by a median value of 0.030 (standard deviation 0.019). On a plot of differences versus standard deviation in the TSG Salinity suggests the differences are very noisy but appear to be converging on anywhere from 0.01 to 0.03 when standard deviations are low. As the standard deviation increases the differences become larger and almost always towards the direction of TSG salinity looking lower. A possible explanation is that where salinity is changing rapidly there is mixing and maybe more bubbles. Between casts #18 and #103 (stations P6 to P26) the differences do not vary a lot and the median difference is 0.026, with a standard deviation of ~0.01. A subset of those casts with the least variation (casts #18-59) has a median difference of 0.023. Shoreward of P6 the TSG salinity is usually much lower than the CTD.

4. FLUORESCENCE

As noted at sea the TSG fluorescence signal is odd. It started with values a little lower than those from the CTD fluorometer, but by the 4th cast it was higher and continued to rise relative to CTD fluorescence until near the end when it had reached the top of the scale. It is possible that the first few values are ok, but not likely. (See 2015-09-ctd-tsg-comp.xls.)

- Loop Bottle - TSG Comparisons

For the 4 loops from Juan de Fuca Strait the TSG fluorescence ranges from 30% to 90% of the loop CHL values with the highest ratio associated with the lowest CHL values, which is typical for this type of fluorometer. For all other samples the TSG fluorescence is much higher than the loop CHL.

The TSG salinity is lower than loop salinity by about 0.032 to 0.034, with little difference between loops taken during stops and those taken while the ship was underway.

While the TSG flow rate was occasionally very low, it was between 1 and 1.2 for the comparisons in this study.

(See 2015-09--loop-TSG-rosette-comp.xlsx.)

- Loop vs Surface Rosette (salinity and chlorophyll)

The Loop salinity is lower than the samples from the rosette by from 0.002 to 0.005 and lower than the CTD salinity during the near-surface bottle stop by from 0.001 to 0.008.

The loop chlorophyll is mostly very close to the rosette chlorophyll samples with the ratio of loops/rosette varying from 0.83 to 1.14. Compared to the CTD fluorescence the loop is mostly lower; in 3 cases the loop is about 65% of the CTD fluorescence and in only 1 of 5 cases is it higher. This type of fluorometer tends to read higher than CHL when CHL is low, and none of these loop CHL samples were >0.5ug/L.

- Calibration History

The temperature and conductivity sensors were recalibrated in December 2013 and have been used for 6 other cruises since then, of which 5 have been processed.

During 2014-21 the TSG salinity was found to be lower than loop samples and CTD salinity by ~0.03 but the difference varied with flow rate which was highly variable.  No recalibration was applied due to the variability in the comparisons and the fact that such a large drift in calibration on its first use seemed unlikely. During 2014-18 the salinity was found to be low by 0.014 but the TSG was so noisy that this was not trusted. During 2014-19 the TSG salinity was found to be low by ~0.02, and the TSG temperature was found to be higher than the CTD temperature by ~0.005Cº. 2014-22 results were not trusted. 2015-01 salinity data were recalibrated by adding 0.02. Intake temperature was found to be higher than CTD temperature by from 0.005 Cº to 0.007 Cº. TSG fluorescence was about 130% of the CTD fluorescence, but comparisons were limited by the absence of any CHL>1.5ug/L.

Conclusions

1. The TSG clock worked well.

2. The TSG flow rate was mostly ok, but for parts of file #1 it was low and the data were noisy during that section, so this will be examined in CTDEDIT and possibly the data should be removed. During file #6 there was a section at the end with zero flow where data should be removed.

3. The temperature in the loop increases by about 0.19Cº.

4. The TSG intake temperature was higher than the CTD temperature by ~0.004Cº to 0.006Cº, which is similar to the observations of 2014-19 and 2015-01. While this could be a real difference, it is more likely that the TSG water is coming from a little higher in the water column. No recalibration is justified.

5. The TSG Salinity is lower than the CTD salinity by between 0.01 and 0.04, with the median during a period of   fairly steady values (casts #18 to 103) is 0.026 and lower than loop samples by ~0.03. Previous cruises have used a correction of +0.02. Recalibration will be applied by adding 0.03. The 3 loop salinity samples taken during CTD casts were within 0.001 of the CTD salinity and 1 sample was very close. The low TSG values are likely due to bubbles as well as calibration drift and, as with temperature, the loop may be drawing water from a little shallower than 4m. Though bubbles were not a huge problem for this cruise there are likely small ones present and the results near-shore are likely due to there being more bubbles in surface waters in those areas.

6. The TSG fluorescence is bad for most of the time and we have no way to being sure it is not bad for all the data, so the channel should be removed.

7. The loop and the rosette CHL and SAL samples compared as well as we can expect.

8. The overall quality of the salinity data was high with little of the spiking that has plagued this equipment in the past. This enables a closer study of the data and the conclusion that the accuracy of salinity measurements is likely limited by the presence of bubbles, which may vary greatly with time and space.

f.) Editing

The ATC files were copied to *.EDT.

Each file was opened in CTDEDIT and the following editing was applied:

File #1 – Salinity data were removed in sections where the flow rate was ~0.6 because there was corruption not seen in the temperature data; likely this was caused by increased presence of bubbles.

File #2 – Short file – no editing applied.

File #3 – Light editing of isolated salinity spikes; there are sections of noisy salinity but not clear which records are good or bad, so not edited.

File #4 – Light editing of isolated salinity spikes.

File #5 – Light editing of isolated salinity spikes.

File #6 – Light editing of isolated salinity spikes and 3 final records removed because the flow rate dropped to 0.

File #7 – Light editing of isolated salinity spikes.

g.) Recalibration

File 2015-09-tsg-recal.ccf was prepared to add 0.03 to the salinity channel and was applied to all *.EDT files.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Difference, Conductivity:Primary, Fluorescence:URU:Seapoint, Flag and Position:New channels.

The flow channel was not removed since it may be useful for further studies of optimal flow rate and there were changes through the cruise.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats.

A note was added to indicate the pond filter was bypassed for files 3 through 7 in an unsuccessful attempt to improve the fluorescence data.

Those files were saved as TOB files.

Header Check was run. Some files had entries for station names that are descriptions of the starting position of the files. Since we have positions and maps, these are not useful, so they were removed.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

##### Loop File

The CHE files were put through program DERIVE to obtain sigma-t.

Data from those files were exported to file 2015-09-che-surface.csv. Data from below 8.9db were removed.  Columns were rearranged to fit a model 6-line header. A sampling method column was added with ROS entered for the CHE data.

Loop data were then added to the file and lined up appropriately. USW was added to the sampling method column for these data.

The data were sorted on event number and then time.

The file break column was filled with value 1 so all data will be in a single file when converted.

The file was then saved as 2015-09-surface-6linehdr.csv.

A comment file was prepared which was essentially the same as the one used in preparing CHE files, with the addition of some details about the TSG. A note was added that the pond filter had been bypassed after 2100 on June 15.

CONVERT was run to produce an IOS Header file. The flags and comments were entered in the headers in the conversion process.

CLEAN was run to get start and stop times and positions.

Header Edit was used to add comments.

The final file was renamed as 2015-09-surface.loop. A track plot looks reasonable and a plot of temperature and salinity versus longitude looks reasonable.

# Particulars

PAR ON: 5, 6, 8, 9, 17, 23, 34, 42, 45, 51, 76, 77, 84, 89, 91, 94, 100, 111, 112, 113, 114, 116, 118, 119,120,121.

Out of Order firing: 8, 9, 12, 15, 17, 18, 23, 32, 38, 42, 51, 55, 76, 82, 90, 94.

Loops: 13 underway and 6 while on station – for Loops P22, P23, P26 and P7 loop was taken at same time as 5m rosette fired.

TSG: Problems with fluorescence. Efforts to clean pond filter proved almost impossible. Bypassing filter did not help. Perhaps problem is with the sensor.

General: Many stops on way up for casts 2000m and deeper due to spooling problems.

5/6/8. Wrong cruise number on rosette log and labels. Should be 2015-09 Not 2015-20. FIXED

17. Slightly off station due to weather.

19. Stopped at 100 and400m on way down due to spooling issue on winch. Also stops on upcast.

23. Niskin 18 closed instead of Niskin 14 for surface samples.

33. Wrong station name in header – should be P13 NOT P12. FIXED

41. Stop at 530m during downcast.

48. Computer crash on downcast, so CTD brought back to surface and restarted as 49. 48 looks bad, probably not worth processing. NOT PROCESSED

49. Wrong station name in header should be P19 not P20.

89. Computer crashed at ~100db of upcast. Lowered to 150m and started cast 90.  Rename bottle file 89.

116. Two Niskins closed at surface but only Niskin 1 is needed in CHE file. Niskin 2 not sampled.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 31Jan2013 | Factory |  |  |
| Conductivity | 1763 | 1Jan2014 | Factory |  |  |
| Secondary Temp. | 5013 | 27Feb2013 | Factory |  |  |
| Secondary Cond. | 3394 | 3Jan2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 1438 | 3Jan2014 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2011 |  |  |  |
| pH | 0691 | 29Dec2010 | Factory |  |  |
| SeaPoint Fluor. | 3640 | n/a |  |  |  |
| Pressure Sensor | 506 | 30Dec2013 | Factory |  |  |
| Altimeter | 62354 | n/a | Factory |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3363       Cruise ID#:	2015-09

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3363 | 28Dec13 | Factory |  |  |
| Conductivity | 3363 | 28Dec13 | Factory |  |  |
| SeaPoint Fluorometer | 2356 |  |  |  |  |
| Temperature:Secondary | ? | ? |  |  |  |
| Flow meter | ? | n/a |  |  |  |