## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 27 Mar 2027 | ### Updated channel names & formats in TOB & Loop files. GG & SH |
| ### 24 July 2024 | ### Added Cesium data to 9 casts. SH |
| ### 18 January 2023 | ### Corrected some flags & comments; removed DIC/ALK from cast 31, pH from cast 115. GG |
| ### 1 December 2021 | ### Corrected Salinity:Bottle Precision lost during the addition of HPLC, re-did the addition of DIC to casts 16, 22, 31, 90 115, SH. |
| ### 19 July 2021 | ### Removed previous DIC & TA additions, added updated DIC, TA and pH data. SH. |
| ### 15 April 021 | ### Added DIC and TA data to 5 files. S.H. |
| ### 11 Dec 2020 | ### RE-ran addition of HPLC Data for minor corrections. S.H. |
| 16 July 2020 | Added HPLC data & corrected truncated flag channel name for DMSP_D |
| 27 Nov 2019 | Added DMSP-D&DMSP-T – values all padded due to bad blanks G.G. |
| 19 Nov. 2018 | Corrected 2015-10-0075.CHE. See note at end of report. |
| 26Jan2016 | Updated CHL sample 889 flag and comment. |

## PROCESSING NOTES

Cruise: 2015-10

Agency: OSD

Location: North-East Pacific

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: 18 August 2015 – 3 September 2015

Processed by: Germaine Gatien

Date of Processing: 27 November 2015 – 18 December 2015

Number of original HEX files: 65	Number of CTD files: 64 (cast #47 only upcast data)

Number of bottle files: 	58		Number of bottle casts processed: 57

Number of original TSG files:  5	Number of processed TSG files:	 4

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0443) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#997) on the secondary pump, a SeaPoint Fluorometer (#3640) on the primary pump, a Biospherical QSP-400 PAR sensor (#4565), and an altimeter (#62354).

A thermosalinograph (Seacat 21 S/N 3363) was mounted with a Wet Labs WETstar fluorometer (S/N ws-3s 953p), remote temperature sensor and a flow meter.

Seasave version V7 22.4 was used for acquisition.

The deck unit was a Seabird model 11, serial number 0619.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

The oxygen kit was Scripps kit #2.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log, rosette log sheets and analysis logs were in good order. Sampling notes from the chief scientist were very useful.

The secondary salinity was found to be low by about 0.0018 with a standard deviation of 0.0024, based on the bottle comparison and by 0.0014 based on the 22 bottles fired at 2000m during a single cast when 1 outlier was removed. The Autosal accuracy is considered ±0.002 and there may be a small effect from incomplete flushing of Niskin bottles which would make the CTD look a little low. No recalibration was applied. The primary salinity was low by about 0.002; the primary signal was noisier than the secondary.

As usual CTD fluorescence is mostly higher than extracted chlorophyll samples at low CHL values, but close to CHL for 0.8<CHL<4ug/L. For CHL>10ug/L the fluorescence reads between 46% and 66% of the CHL (with the exception of one outlier). The large differences at high CHL may be partly due to incomplete flushing of Niskin bottles in high gradient areas. There were no extracted chlorophyll samples in the range 4ug/L to 13.5ug/L.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 200db

±0.2 mL/L from 200db to 500db

±0.04 mL/L below 500db

Dissolved oxygen and salinity data were removed from cast #74 between about 175m and 210m because of bad values that must be related to an interruption of flow in the secondary plumbing. The primary salinity was better through that section, but much worse elsewhere.

There were problems with the thermosalinograph. There was a 3.5-day period during which positions were not logged and a 4.5-hour period with no acquisition, and there were a lot of small spikes in the salinity (~0.2 to 0.4psu). Position data were obtained from other underway files and added to the TSG files by matching times. The larger salinity spikes were cleaned using a graphical editor, but they may well be indicative of a more general problem of bubbles reducing salinity. The TSG salinity was found to read lower than both CTD and loop salinity samples with median differences of about 0.17psu. This is a much larger difference than usual and is unlikely to be due to calibration drift. Salinity was recalibrated by adding 0.17psu. Fluorescence is given in volts because the comparison with CTD fluorescence and extracted chlorophyll samples from the loop did not lead to a satisfactory means of calibration.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

There were sampling notes provided by the chief scientist.

Nutrients, extracted chlorophyll, dissolved oxygen, DMS and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked. Only the pressure sensor had been used since the last factory recalibration.

The XMLCON files did not change through the cruise.

The calibration constants were checked for all instruments and the only correction needed was to fix a spelling error. The corrected file was saved as 2015-10-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2015-10-ctd.xmlcon to create CNV files.

A few casts were examined and all expected channels are present.

The temperature and conductivity channels are close at depth. Near the surface there is a lot of variability especially during upcasts. There were spikes in the primary sensors.

The descent rate was often extremely noisy with complete reversals of the CTD, especially later in the cruise.

The altimetry appears to have worked well but most casts do not get close enough to the bottom to have a signal.

Fluorescence, transmissivity and PAR look ok.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2015-10-ctd.xmlcon.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. As usual there were many near the surface, but these may be real variations. For a few casts there were some deeper outliers, so those files were opened in CTDEDIT:

Event 21. Salinity was cleaned (both channels).

Event 57. Salinity was cleaned (primary only).

Event 98. Salinity was cleaned (both channels).

The output files were copied to *.bot.

A preliminary header check and a cross-reference check were run and no problems were found. CTD fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

Sort was used to ensure that the ADDSAMP file was in sample number order.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

The SAM files were bin-averaged on bottle # and called SAMAVG.

At this point it was discovered that some of the salinity and nutrient samples came from CTD casts run for another institution, so those were removed.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2015-10-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2015-10chl*.xls. The file included comments and flags and a precision study. The loop samples were removed to spreadsheet, 2015-10-loop-data.xlsx. A simplified version of the spreadsheet was prepared and saved as 2015-10chl.csv. The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2015-10oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and was then saved as 2015-10oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2015-10SAL.xls. The analysis was done between 3 and 19 days of collection. The files were simplified and saved as 2015-10sal.csv. The loop samples were removed to the loop spreadsheet, 2015-10-loop-data.xlsx. File 2015-10sal.csv was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2015-10_nutrients*.xlsx. This includes a precision study. The file was simplified, reordered on sample numbers and saved as 2015-10-nuts.csv. The file was converted to individual NUT files.

DMS

DMS data were obtained in file DMS summary (2015-10).xls. Values given as < were changed to 0 and the comments that will go into the header will explain that 0 means below detectable level. (There was a separate report on analysis techniques and problems, 2015-10 DMS report.doc.) The file was then saved as 2015-10DMS.csv and converted to individual DMS files.

The SAL, CHL, OXY, NUT and DMS files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. Errors found or things to be adjusted were:

- Oxygen: All DO casts. Flag 6 added to duplicates.

- Cast #31 – removed Niskin #1 from SAMAVG and MRG file.

- Cast #33, sample #198 – flag removed – new rosette log sheet explains this

- Cast #47 – sample #344 flag changed from 2 to 9.

- Cast #88 – only sample was for water, no need for chem file, so removed.

- Cast #102 – samples 876 to 878 were entered with pad values but should be 0. OK in analyst S/S.

- Cast #111 – Sample 943, salinity was entered as pad and flag 9. Planned but no sample taken.

After these corrections the dissolved oxygen file was reconverted to *.oxy, the merge process repeated and CLEAN rerun.

##### Compare

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The comparison has scatter that looks like data are in 2 different groups. To see if this was due to hysteresis, the data above 1500m were excluded from the comparison to see if they stood out as different from those below that level, and they did. The DO sensor settings have never been checked since it was last serviced. Usually changes to factor E in the dissolved oxygen calibration parameters improves such results. E=0.036 is the default value. Tests were run using a variety of E settings and E=0.0375 brought the deep DO into line with that above 1500m in COMPARE.

A return was made to the beginning of building the bottle files by reconverting all data using this new setting. COMPARE was then rerun.

There was only one significant outlier, with most of the data falling into a reasonably tight group.

When outliers were removed based on residuals the fit was:

CTD DO Corrected = CTD DO * 1.0151 + 0.0426

To test for hysteresis points below 1500m were also excluded and the fit was then:

CTD DO Corrected = CTD DO * 1.0154 + 0.0465

This is very close to the full fit despite there being few points in the deep fit, so the value used for factor E looks good.

There were no significant outliers.

For more details see document 2015-10-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. There were no outliers.

Salinity

Compare was run with pressure as reference channel.

Two major outliers were identified:

- Event #22, sample #53: There is a note on the rosette log that Niskin bottles 4 and 5 were fired out of order. The original plan was that Niskin bottles were to be fired out of order, but they were actually fired in order. Sample #53 was supposed to come from Niskin #6 at 5m, but the value of sample #53 makes it clear that it came from 30m, so from Niskin #4. So it should have had sample #51. This is a case of mislabelling caused by confusion in how bottles were fired.

- Event #85, sample #693: One of the duplicates has a value that is clearly from near the surface. Reject that value and use only 34.5807 with flag 2.

- Event #70, sample #557. There were duplicates and they are in very poor agreement, 34.3677 & 34.4377. The higher value is close to the CTD salinity and the other is much lower. This does not look like the sample came from another bottle. The value was changed to 34.4394 and flag 2 entered.

After changing those samples, COMPARE was rerun.

When data above 150db are excluded along with a few bottles with standard deviation in the CTD salinity >0.0008, the primary CTD salinity is lower than bottles by an average of 0.0020 (standard deviation 0.0015) and the secondary lower by 0.0018 (standard deviation 0.0024). The primary fit is flatter, but only slightly so. However, when the outliers are examined, there are larger ones for the secondary. These are above 200m but may be evidence of less reliable CTD data. On the other hand the average of standard deviations in the bottle comparison is roughly the same for both sensor pairs.

A study of the 22 bottles fired at 2000m during cast 390 shows similar results though the primary salinity is lower than bottles by more than in the cruise as a whole. The secondary results were similar from both though again the difference is not large.

|  | Study of bottles fired at 2000m Cast #90 |  |  |  |
| --- | --- | --- | --- | --- |
|  | Sal 0 |  | Sal1 |  |
|  | average | stdev | average | Stdev |
| all | -0.0025 | 0.0015 | -0.0017 | 0.0015 |
| excl. 1 | -0.0023 | 0.0012 | -0.0014 | 0.0011 |

Examination of the differences versus time suggests a different picture with the secondary having a flatter fit. However, for the secondary there is a lot of variability from one cast to another, one high and then another low, so that the flat fit is likely accidental.

For full details for the COMPARE run see file 2015-10-sal-comp1.xls.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor.

CHL values ranged from ~0 to 20.6ug/L, but there were no values between 4 and 13.5ug/L. As usual, at very low CHL values the CTD fluorescence looks higher than chlorophyll but the two are reasonably close for 0.8<CHL<4ug/L. For CHL>10ug/L the fluorescence reads between 46% and 66% of the CHL (with the exception of one outlier). The large differences at high CHL may be partly due to flushing problems in high gradient areas.

For full details of the comparison see file 2015-10-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on a few casts and results with a setting of +2.5 to +3.5 all looked reasonable, but +3.5 was best overall.

ALIGNCTD was run on all casts using +3.5s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. During other recent cruises, the default setting of (α = 0.0245, β=9.5) looked better for both conductivity channels than others tested. One cast was checked for this cruise and the default setting does improve the data.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2015-10-0028 | 500 | -0.0020 | -0.00040 | -0.0013 | High, Noisy |
|  | 1000 | -0.0011 | -0.00010 | -0.0001 | “ |
| 2015-10-0039 | 500 | -0.0012 | -0.00034 | -0.0024 | High, Moderate |
|  | 1000 | -0.0010 | -0.00015 | -0.0005 | “ |
|  | 1900 | -0.0011 | -0.00007 | +0.0003 | “ |
| 2015-10-0070 | 500 | -0.0010 | -0.00016 | -0.0008 | High, Very Noisy |
|  | 1000 | -0.0010 | -0.00006 | +0.0003 | “ |
|  | 1900 | -0.0013 | -0.00002 | +0.0009 | “ |
|  | 2900 | -0.0013 | -0.00001 | +0.0013 | “ |
|  | 3500 | -0.0015 | 0 | +0.0015 | ‘ |
| 2015-10-0099 | 500 | -0.0011 | -0.00016 | -0.0009 | High, Very Noisy |
|  | 1000 | -0.0010 | -0.00006 | +0.0004 | “ |
|  | 1900 | -0.0013 | -0.00003 | +0.0010 | “ |
|  | 2900 | -0.0014 | -0.00001 | +0.0014 | “ |
|  | 3500 | -0.0015 | 0 | +0.0015 | “ |

The differences are small and show no significant change through the cruise; they are slightly higher at 500m for the shallow casts, probably due to higher local gradients so slight misalignment becomes more significant.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. One station name was given as P15ish in the file and log book; this was changed to PAR as it was not very close to P15. It was noted that the times in the log book are often later than in the file headers. This is likely due to some people recording the time when the CTD is turned on and others when acquisition starts. There can be 10 to 20 minutes difference between those actions.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 3.0db which is reasonable for the Tully. The lowest pressure recorded was +0.79db at the beginning file #71; the associated conductivity and salinity have “in-water” values and the pumps were on. The transmissivity is very high, but not as high as expected in air. So the pressure looks about right.

A header check was run. The range of primary temperature, conductivity and salinity values suggests at least some spikes in primary channels. No other problems were noted.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. After checking maximum pressures against water depth, plots were made to check the altimetry for casts where the entries look unlikely. The entries were incorrect for casts #23, 24, 26, 31, 32, 40, 43, 47, 62, 92 & 102 in the CLN files and #42, 43, 47 & 62 in the SAMAVG files. Those altimeter headers were removed.

Water depths differed from those in the log book in many cases. When the difference was >3m, the entry was changed in the CLN and SAMAVG files to match the log. For cast #55 neither the log nor the header looked reasonable, but an estimate was made using the altimetry and maximum pressure.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 3 casts to choose settings for alignment of the 2 conductivity channels. For the primary conductivity there was little improvement with any setting tried, but the best was with +0.2 records. For the secondary the effect was much more obvious, with the best results being with a setting of -0.7 records, but -0.5 also looks good.

SHIFT was run on all casts using +0.2 records for the primary conductivity.

SHIFT was run on all casts using -0.7 records for the secondary conductivity.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only casts with warnings were cast #47 which had upcast data only and #55 with warnings that pertain to the bottom of the cast or upcast section, so no action was needed.

##### Other Comparisons

Previous experience with these sensors –

There was no history available for the temperature, conductivity and dissolved oxygen sensors. They have not been used on any other cruise that has been processed since they were last calibrated at the factory.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Salinity values were within the climatology except for some values above the maximum around 100 to 140m at P20 and P13 and near the minimum between 275m and 400m at P9, P12 and P15. The salinity excursions are not systematic, so not suggestive of a calibration problem. Temperature was frequently above the maximum between 25m and 100m at P5 to P12 and between 25m and 40m at P26. Temperature was higher than or at the maximum at about 275m at P3 and P4. That was not seen at P4 when it was occupied at the end of the cruise. There have been many observations of unusually high temperatures in the region at those depths so the observations do not suggest a calibration problem.

Repeat Casts – There are many repeat casts but most include a deep and a shallow cast, so those are not good for checking repeatability. At P26 there were 2 casts to at least 1000m about 7 hours apart. At about 1000m the temperatures vary by about 0.0015C° and salinity by 0.0002 along lines of constant sigma-T. At this level we expect very little change and the CTD does well on this test of repeatability.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The comparison with bottles shows little difference between the two salinity channels with the secondary being very slightly closer to bottles. An examination of a few casts shows a lot more noise in the primary T-S plots, so the secondary was selected for editing and eventual archiving.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. Some bad salinity points were removed. For cast #74 all salinity points between 184db and 210db were removed; it appeared that something went wrong in the flow to the conductivity sensor. (The primary data was somewhat better, but spikes elsewhere in the profile made the primary a poor alternative.)

All files required some editing and most required heavy editing due to corruption by shed wakes.

After editing was complete, an error was discovered in the dissolved oxygen concentration derivation. It was correct in the bottle files but incorrect in files that had been processed through the editing phase. So DERIVE was run a second time and the output files were put through Conversion to IOS HEADER format, CLEAN and DELETE with distinctive extensions, IOSDO, CLNDO, DELDO. MERGE was then used to select all channels except the oxygen concentration from the edited files and dissolved oxygen concentration from the new files with pressure used as the merge channel.

##### Further Recalibration

No adjustment appears necessary for the pressure sensor.

The secondary salinity was found to be low by about 0.0018 with a standard deviation of 0.0024, based on the bottle comparison and by 0.0014 based on the 22 bottles fired at 2000m during a single cast when 1 outlier was removed. The Autosal accuracy is considered ±0.002 and there may be a small effect from incomplete flushing of Niskin bottles which would make the CTD look a little low. Although sample analysis was prompt, there could also be very small effects from evaporation or desorption of samples which would also lead to the CTD looking slightly low. No recalibration is justified.

File 2015-10-recal2.ccf was created to correct the Oxygen:Dissolved:SBE channel using:

CTD DO Corrected = CTD DO * 1.0151 + 0.0426

This correction was first applied to the SAM and MRGCLN2 files. COMPARE was rerun and the results confirm that the recalibrations were applied properly. The average of differences in the DO fit when outliers were removed was -0.0003mL/L and the standard deviation was 0.019mL/L.

(See file 2015-10-DO-comp2.xlsx for details.)

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. The first run turned up a major problem in the downcast CTD DO caused by an error in the derivation of DO concentration. This was fixed as described in section 15.

There is a lot of scatter but most of it is from depths where the vertical DO gradient was high. When outliers were removed, the CTD DO was higher than the bottles by an average of ~0.005mL/L and standard deviation of 0.081mL/L. There are two sources of error which unfortunately have the same effect on the result so we can’t distinguish which is the larger effect:

- Even a slight inefficiency in Niskin flushing would lead to lower DO values in the bottles above the DO minimum where the highest vertical gradients occur. Below the DO minimum the error would have the opposite sign but the effect would be smaller due to lower gradients.

- The CTD values are also likely to lag slightly due to slow response which would make them read high above the DO minimum and low below the minimum. The errors would be most significant above the DO minimum because vertical gradients are higher there.

In COMPARE the average difference below 1400db is -0.006mL/L with a standard deviation of 0.021. Above 1200m the averaged is +0.009mL/L with a standard deviation of 0.09.

No further recalibration is justified. See 2015-10-dox-comp3.xlsx for details.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. A problem was found in cast #74 with bad dissolved oxygen data between about 175m and 210m. The salinity data from this section were already removed and the DO sensor was on the same pump. This was clearly a problem with flow to the sensors. A text editor was used to replace the DO data with pad values for the records in that pressure range that already have padded salinity data. No other problems were found.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

The PAR channel was removed from casts: #20-40, 46-67, 76-89, 93-115 and 121.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and unedited except that

some records were removed in editing temperature and salinity.

A graphical editor was used to remove records from obviously bad sections of

temperature and salinity data, and some salinity data were smoothed to

remove small spikes. Where there was corruption of salinity data that

was not seen in temperature data, only salinity points were removed.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

Downcast Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 200db

±0.2 mL/L from 200db to 500db

±0.04 mL/L below 500db

For details on the processing see document: 2015-10_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

A cross-reference list was produced.

The sensor history was updated.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted.

Values were between 74% and 86% in Haro Strait and near shore from stations LB1 to LB4. The values rise for the LC line until peaking at P2 and P3 in the 110%-130% range. From P4 west the saturation is mostly between 100% and 110% with just a few lower values and 102% being most frequent. These values are close to the usual 103% commonly seen offshore, suggesting that the DO calibration is reasonably good.

##### Final Bottle Files

The MRGCOR2 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

The PAR channel was removed from casts: #20-40, 46-67, 76-89, 93-115 and 121.

A second SBE DO channel was added for both the CTD DO and titrated DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

A second set of data were prepared using the steps above except that Oxygen:Voltage:SBE was not removed. Those data were exported to a spreadsheet file for the use of Glenn Cooper in his study of a Rinko DO sensor that was also in use. Later the Voltage 6 and Voltage 7 data were converted and added to this file as they are required for the calculation of dissolved oxygen from the Rinko sensor.

The track plot looks ok.

Data were exported from the CHE files to file 2015-10-bottles-final.xlsx. The entries were compared with the rosette log sheets to ensure no samples had been missed and a few discrepancies were found and corrected including a salinity sample that was flagged 0 but should have been 9, and several DMS values that had been entered with pad values but should have been 0

Standards check and a header check were run on all files. The only problem found was that there were some very low draw temperature values in casts #70 and 79. The analyst provided new values, so the derivation of DO in mass values was redone and Reorder and Head Edit were rerun.

Plots of each file were examined to ensure no further problems were found.

A cross-reference listing was produced for the CHE files.

##### Thermosalinograph Data

There were loop nutrients, extracted chlorophyll and salinity samples taken, some while stopped and some while underway.

a.) Checking calibrations

The 5 configuration files identical. One file was renamed as 2015-10-tsg.xmlcon. No errors were found in the calibration parameters.

b.) Conversion of Files

The 5 files were converted to CNV files using configuration file 2015-10-tsg.con, but there were no data in the 5th HEX file, so only 4 CNV files were produced.

Those CNV files were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

An initial track plot shows a spatial gap between files #3 and #4, due to the positions failing to update at for the latter part of file #3. Why this happened is explained in the following extract from the post-cruise report:

The TSG NMEA got disconnected and we lost the signal for four days. Although the position data wasn’t being recorded, at least we were still getting temperature, salinity and fluorescence values. When rebooting the TSG computer to try getting the NMEA data back – before realising that it was simply disconnected – some settings got changed and the TSG stopped communicating altogether. We then noticed that the ground cable had gotten disconnected. That was part of the problem. After reconnecting the ground cable to a pipe going to the manifold, we had to reset all the settings. It turns out that the instructions written in the TSG logbook are wrong. Seaterm had to be used to reconnect to the TSG.

If the time was downloading properly, we should be able to use other time-series files to fill in the positions. An SCS file with time and positions was found that includes the missing period. That file contains readings every 10s, which the TSG is recording every 30s. The SCS file was opened in EXCEL and thinned to every 3rd record, using a mod 3 calculation. The CNV file was opened in EXCEL and the SCS positions were added to the file, as needed. That file was saved as 2015-10-0003.csv and was then converted to file 2015-10-0003.IOS2. File 2015-10-0003.IOS was put through REMOVE to remove latitude and longitude with output named as *.IOS1. Merge was used to combine the IOS1 file with the latitude and longitude from the IOS2 file to create a new IOS file. This was put through CLEAN and ADD TIME CHANNEL. This provided a file that appears to be good, with reasonable track and time-series plots. There is still a 4.5 hour break between files #3 and 4, but the 3.5days with “stuck positions” have been restored.

A quick check shows the ship stopping for a CTD cast at the right time and the temperature readings are in reasonable agreement with the CTD cast. Further checks will be done later.

Time-series plots were produced. The flow rate is ~1 and steady through the cruise except for a brief drop to 0 in the first file and 0 values for a few records at the beginning and end of file #4. Fluorescence spiked at the end of file #4, presumably due to flow stopping. There are small spikes in salinity, ~0.4, throughout the record. The heating in the loop varies with intake temperature as expected.

The track plot looks fine. The plot was added to the end of this report.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db and exported to a spreadsheet which was saved as 2015-10-ctd-tsg-comp.xls. There were 61 casts which overlapped with TSG files.

The TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, lab temperature, salinity and fluorescence and the files were reduced to the times of CTD files and loop samples.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude and longitude were all≤0.0008° and only 2 differences were >0.0002°; the medians in both were 0.0000°. This shows both the times and positions are reliable for both systems. The “patched” section for TSG file #3 looks fine.

The loop samples were combined in file 2015-10-loop-TSG-rosette-comp.xlsx. There were 3 cases where loop samples were taken during a rosette cast. The rosette salinity and CHL bottle data were combined with the loop sampling and median (over 5 records) TSG readings were added. The only nutrient sampling was from underway loop samples so there are no comparisons possible.

The two spreadsheets will be used in step (d) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from Loop & Rosette Samples and TSG and CTD data

- T1 vs T2 The intake temperature sensor worked throughout the cruise. The median difference between intake temperature and lab temperature in the 4 files and the median intake temperatures were calculated. As usual, the heating in the loop varies inversely with intake temperature.

| File # | Median (Tlab -Tintake) | Median Intake Temperature |
| --- | --- | --- |
| 1 | 0.174 | 13.5 |
| 2 | 0.127 | 18.0 |
| 3 | 0.159 | 15.3 |
| 4 | 0.132 | 17.3 |

- TSG vs CTD The spreadsheet comparing CTD and TSG files was then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG.

1. Intake Temperature The intake temperature is higher than the CTD temperature by a median of 0.0044C° (standard deviation 0.13C°) or 0.0027C° if only the 20 casts with the smallest standard deviation in the TSG temperature are included.

2. LAB TEMP The lab temperature was higher than the CTD temperature by a median value of 0.161 C° or 0.131C° if only the 20 casts with the smallest standard deviation in the TSG intake temperature are included.

3. SALINITY The TSG salinity data are lower than the CTD salinity by a median value of 0.179 (standard deviation 0.055). If only the 20 casts with the lowest standard deviation in the TSG salinity are included the median TSG salinity is low by 0.161. On a plot of differences versus standard deviation in the TSG Salinity the differences are very noisy but appear to be converging on ~ 0.01 as the standard deviations decrease. Presumably where the standard deviation is high, there is a strong vertical gradient; if the TSG is sampling from higher in the water column than the CTD it would read low.

4. FLUORESCENCE

The TSG fluorescence data are uncalibrated and expressed in volts. The ratio of the TSG fluorescence to that from the CTD varies greatly. The lowest ratios come early in the cruise when fluorescence was highest based on loop samples or CTD data. The highest ratio is from a section with very low CTD fluorescence. The median TSG fluorescence is about 15% of the TSG fluorescence or 25% if only data from the 20 casts with the lowest standard deviation in the TSG fluorescence are included. (See 2015-10-ctd-tsg-comp.xls.)

- Loop Bottle - TSG Comparisons

For the 4 loops from Juan de Fuca Strait the TSG fluorescence ranges from 30% to 75% of the loop CHL values with the highest ratio associated with the lowest CHL values, which is typical for this type of fluorometer. For the 2 samples with CHL>1ug/L, the TSG fluorescence value equals only 5% and 9% of the CHL. More high CHL samples would be needed to make a rough estimate of how to calibrate the TSG fluorometer.

The TSG salinity is lower than loop salinity by about 0.032 to 0.034, with little difference between loops taken during stops and those taken while the ship was underway.

While the TSG flow rate was occasionally very low, it was between 1 and 1.2 for the comparisons in this study.

(See 2015-10-loop-TSG-rosette-comp.xlsx.)

- Loop vs Surface Rosette (salinity and chlorophyll)

There were only 3 rosette casts that overlapped with loop samples, and only 1 of those had a salinity sample. The loop salinity was higher than the rosette sample by 0.005. The extracted CHL samples were very close to the rosette samples, one low by ~5%, one high by ~6% and one equal, but all values were very low so the differences are not significant.

- Calibration History

The temperature and conductivity sensors were recalibrated in December 2013 and have been used for 8 other cruises since then. It was mounted on the Vector for 1 of those. The fluorometer has not been used before with this equipment and no calibration information is available.

During 2014-21 the TSG salinity was found to be lower than loop samples and CTD salinity by ~0.03 but the difference varied with flow rate which was highly variable.  No recalibration was applied due to the variability in the comparisons and the fact that such a large drift in calibration on its first use seemed unlikely. During 2014-18 the salinity was found to be low by 0.014 but the TSG was so noisy that this was not trusted. During 2014-19 the TSG salinity was found to be low by ~0.02, and the TSG temperature was found to be higher than the CTD temperature by ~0.005Cº. 2014-22 results were not trusted. 2015-01 salinity data were recalibrated by adding 0.02. Intake temperature was found to be higher than CTD temperature by from 0.005 Cº to 0.007 Cº. TSG fluorescence was about 130% of the CTD fluorescence, but comparisons were limited by the absence of any CHL>1.5ug/L. For 2015-09 the intake temperature was higher than CTD temperatures by 0.004 to 0.006, and salinity was lower than loop samples and CTD salinity by about 0.03. There was no useful comparison from 2015-46.

Conclusions

1. The TSG clock worked well.

2. The TSG flow rate was good, except for a few brief drop-outs that will be addressed in editing the files.

3. The temperature in the loop increases by roughly 0.13 to 0.17Cº depending on intake temperature.

4. The TSG intake temperature was higher than the CTD temperature by ~0.005C which is similar to other cruises since the last factory calibration. While this could be a real difference, it is more likely that the TSG water is coming from a little higher in the water column. No recalibration is justified.

5. The TSG Salinity is lower than the CTD salinity by between 0.1 and 0.7 and lower than the loop samples by a median value of 0.17. While the differences are at the low end when standard deviations are low, the loop samples argue for a larger correction than +0.01. This is a very large change from previous uses of this equipment.

6. The comparison of the TSG fluorescence with loop CHL samples includes mostly samples with low CHL when fluorometers tend to read high relative to CHL. The 2 high CHL cases suggest a correction factor of 11 or 20 while the other samples suggest a correction factor of about 2. The comparison with CTD fluorescence suggests a factor of 16 and an offset of 1, but that would clearly overcorrect for most of the record when chlorophyll values are low. No recalibration will be applied, but the raw data will be archived.

7. The loop and the rosette CHL and SAL samples compared well.

8. There were many small spikes towards low salinity, ~0.2 to 0.4psu. This may be caused by bubbles and might account for low overall TSG salinity. When there are fewer bubbles that would lower the standard deviation in salinity.

9. The salinity will be recalibrated by adding 0.17 to the salinity. This may overcorrect for areas with few spikes, but for most of the record it will under-correct. This error is likely not due to calibration drift.

f.) Editing

The ATC files were copied to *.EDT.

Each file was opened in CTDEDIT and the following editing was applied:

File #1 – Salinity cleaned. Some points removed when flow was off.

File #2 – Salinity cleaned.

File #3 – Salinity cleaned.

File #4 – Salinity cleaned. Some points removed when flow was off.

g.) Recalibration

File 2015-10-tsg-recal.ccf was prepared to add 0.17 to the salinity channel and was applied to all *.EDT files.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Difference, Conductivity:Primary, Flag and Position:New channels.

The flow channel was not removed since it may be useful for further studies of optimal flow rate and there were changes through the cruise.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats.

Those files were saved as TOB files.

The standards check found some problems in file #3 which had been created in a different way due to the missing positions. Adjustments were made until the standard check found no further problems.

Header Check was run. Some files had entries for station names that are descriptions of the starting position of the files. These are not useful, so they were removed.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

##### Loop File

The CHE files were put through program DERIVE to obtain sigma-t.

Data from those files were exported to file 2015-10-che-surface.csv. Data from below 8db were removed.  Columns were rearranged to fit a model 6-line header. A sampling method column was added with ROS entered for the CHE data.

Loop data were then added to the file and lined up appropriately. USW was added to the sampling method column for these data.

The data were sorted on event number and then time.

The file break column was filled with value 1 so all data will be in a single file when converted.

The file was then saved as 2015-10-surface-6linehdr.csv. Cases where there were comments were checked to ensure that all comments got included. In one case there were comments for 2 different variables and only one was included, so the comment was adjusted.

A comment file was prepared which was essentially the same as the one used in preparing CHE files, with the addition of some details about the TSG.

CONVERT was run to produce an IOS Header file. The flags and comments were entered in the headers in the conversion process.

CLEAN was run to get start and stop times and positions.

Header Edit was used to add comments.

The final file was renamed as 2015-10-surface.loop. A track plot looks reasonable and a plot of temperature and salinity versus longitude looks reasonable.

# Particulars

PAR ON: 2, 4, 7, 8, 9, 10, 11, 13, 14, 15, 16, 18, 21, 22, 23, 24, 26, 28, 31, 32, 40, 50, 53, 60, 61, 71, 82, 88, 101, 102.

Out of Order firing: 16, 18, 30, 31, 34, 39, 45, 50, 53, 55, 61, 68, 72, 75, 82, 98, 102, 103, 111, 115.

1. Supposed to be test cast but problem with wire. Not run.

2. File misnamed as 2015-01-0002.

4. Niskin 6 not closed, not needed.

22. Niskins were meant to be fired out of order but were not, so Niskin 4 was closed at 30 instead of 5m and #5 was closed at 5 instead of 30m.

31. Spigot on Niskin 1 didn’t look right as deployed so Niskin 21 fired at 300 as well and Niskin 21 was sampled.

35. Accidently came up to 640db then back down.

40. PAR cast – brought PAR out of water at start and end of cast.

47. Archiving only on for upcast.

47. Niskin 19 skipped, so Niskins #19 to 22 were fired at the wrong depths, 50, 25, 10, 5m respectively.

55. CTD hit bottom. No IOS sampling.

88. One bottle fired for water but no CHE file needed – no sample #.

102. Niskins 15 and 16 not sampled, 6 and 7 sampled instead.

103. Niskins 1 and 2 closed at 1000m but not sampled – removed from CHE file. Keep Niskins 5 to 23.

115. Station P4 – written P25 on rosette log.

Correction – 19 November 2018 – Though not noted in the rosette log, Niskin #4 did not fire during event #75.

This caused confusion in preparing the ADDSAMP file and some chemistry data were assigned to incorrect bottles.

A salinity sample was given a sample number associated with Niskin #4, but no samples were found missing, so it is unknown where it really came from. The ADDSAMP file and file 2010-10-0075.CHE were corrected

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 17Dec2014 | Factory |  |  |
| Conductivity | 2280 | 19Dec2014 | Factory |  |  |
| Secondary Temp. | 2663 | 15Jan2015 | Factory |  |  |
| Secondary Cond. | 2754 | 19Dec2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 17Jan2015 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| SeaPoint Fluor. | 3640 | n/a |  |  |  |
| Pressure Sensor | 0443 | 17Dec2014 | Factory |  |  |
| Altimeter | 62354 | n/a | Factory |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3363       Cruise ID#:	2015-10

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3363 | 28Dec13 | Factory |  |  |
| Conductivity | 3363 | 28Dec13 | Factory |  |  |
| WetLabs Fluorometer | Ws3s-953p |  |  |  |  |
| Temperature:Secondary | ? | ? |  |  |  |
| Flow meter | ? | n/a |  |  |  |