## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27 March 2025 | Updated channel names & formats in TOB files. G.G. |
| 15 March 2021 | Added Carbon data to files, 1, 4, 15, 28, 46, 58, 78 and 80. S.H. |
| 16 July 2020 | Added HPLC data G.G. |

## PROCESSING NOTES

Cruise: 2015-17

Agency: Ocean Sciences Division

Location: Strait of Georgia and Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Strait

Party Chief: Chandler P.

Platform: Vector

Date: April 1, 2015 – April 8, 2015

Processed by: Germaine Gatien

Date of Processing: 4 May 2015 – 31 August 2015

Number of original HEX files: 81

Number of CTD files: 81	Number of bottle casts:	25

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0506) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#1438), a SeaPoint Fluorometer (#3642), a Biospherical QSP-400 PAR sensor (#4565), a surface PAR (#16504), a pH sensor (#0851) and an altimeter (#1252).

The data logging computer was Vector CTD Laptop.

The data acquisition program was Seasave.

The CTD deck unit was an SBE model 11+, serial number 0619.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log, rosette log sheets and analysis logs were in good order.

There were a few errors in the headers entered. The file names all had 2015-07 instead of 2015-17. “Stn” was entered where the word “Station” should be used, so in conversion to IOS HEADERS the station names were not picked up. A spreadsheet was used to merge this information with the other location headers.

There were no rosette sheets for 2 casts that had only a single Niskin bottle fired. There was a note of the sample numbers in the Daily Science Log Book, but no indication of what sampling was done. That prevents a check to make sure no data are missing from the bottle files. There were many cases in which the event number was entered incorrectly so that samples were not picked up in the building of the bottle files. With no rosette log sheets as a source of information, such missed samples might not get found.

The configuration file used at sea had the pH and PAR sensors on the wrong external voltages. When the parameters were reversed the data looked fine. This reversal caused considerable confusion at sea as the pH values looked bad. It is a good idea to send files from an early cast to IOS for review to ensure all expected data are being recorded properly and profiles look reasonable; while this project is a short one, there is still a chance to catch trouble early. There is usually (but not always) someone available to do such a review.

The surface PAR sensor produced no useful data for this cruise. The sensor is believed to have malfunctioned.

The secondary pump was found in need of repair after this cruise and there is evidence that it was not performing well during the previous cruise, so the primary sensors were selected for archiving. This could also affect the dissolved oxygen data which was on the same pump, possibly making the response time a little longer than usual.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.13 mL/L from 0 to 30db

±0.1 mL/L from 30db to 100db

±0.08 mL/L from 100db to 250db

±0.03 mL/L below 250db

There was a lot of scatter in the comparison of CTD salinity with bottles. Poor flushing of Niskin bottles is suspected in some cases. The correction applied is an estimate based on a few bottles from areas of low vertical salinity gradient and the results of cruise 2015-01 when the same sensors were used.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data were available at the time of processing. Calibration is required for each cast to get absolute values although

general trends within a cast are likely real.

An SBE-21 Thermosalinograph recently installed on the Vector was put into use on this cruise. After some initial difficulties 2 useful files were acquired. No loop sampling is available at this time, so comparison with the CTD is the only check on calibration. The temperature in the lab is about 0.1C° higher than that measured by the CTD which is likely to be due to heating in the loop. However, it is also possible that water in the loop comes from a little higher than the 2m data used in the comparison, which would also lead to values looking a little high. There were variations in the pump pressure which would also affect heating in the loop. The salinity comparison was extremely noisy but median values show the salinity being reasonably close to the CTD values with a range from being low by ~0.009 to high by ~0.0004, depending on how many outliers were excluded. Due to the many uncertainties no proxy for intake temperature was derived and only Temperature:Lab will be archived. The setting for the Position:New channel should be adjusted so that it does not read 0 while the ship is moving.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX. All files had the wrong cruise number in the file names. Those were corrected.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets for dissolved oxygen, extracted chlorophyll and salinity.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked.

The XMLCON files did not change through the cruise.

The calibration constants were checked for all instruments and some corrections were needed:

- The chief scientist advised that the pH and PAR entries might have been reversed. A test conversion showed that was indeed the case, so they were changed in the con file.

- The date of calibration was added to the PAR sensor entry.

- The date formats for the transmissometer and pH sensor were changed.

- Based on the results of the last 3 cruises that used this CTD (#506), the pressure offset was updated to value -0.25.

The corrected file was saved as 2015-17-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2015-17-ctd.xmlcon to create CNV files.

A few casts were examined.  As noted at sea there is only noise in the SPAR channel. All other expected channels are present. The upcast and downcast temperature and conductivity channels are reasonably close, with the upcasts generally noisier than downcasts. The primary and secondary are also reasonably close though there significant differences occasionally. There are spikes in the primary conductivity. Dissolved oxygen, fluorescence, PAR, pH, altimetry and transmissivity profiles all look normal. As usual for this region the descent rate is high and steady in inland areas and very noisy with many complete reversals of direction for some casts near the mouth of Juan de Fuca Strait.

A problem was found in the secondary pump after this cruise. Evidence of a problem with either pump was sought and the only thing noted is the spikes in the primary conductivity that are not seen in temperature and a few spikes in the secondary conductivity as well. There is no obvious correspondence to features in either fluorescence or dissolved oxygen traces, but the former is so spiky and the DO so smooth that we might not see an effect if there were one. On a previous cruise there was evidence in the secondary channels of slow response at the surface and after a stop, as shed wakes moved past the sensors.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2015-17-ctd.xmlcon.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and none were found.

MERGE: CSV Files to Headers was used to add station names as was done for the full files.

A preliminary header check was done and no problems were found. Fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine.

Those files were bin averaged on bottle number.

For event #1 the 9 bottles fired as a test and having no sample numbers were removed from the SAMAVG file.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2015-17-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2015-17chl*.xls. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2015-17chl.csv which was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2015-17oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2015-17oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2015-17SAL.xls. The analysis was done within 14-19 days of collection. The files were simplified and saved as 2015-17sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2015-17nuts.xls. This includes a precision study.

Then the file was simplified, reordered on sample numbers and saved as 2015-17-nuts.csv. The file was ordered on sample numbers and converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

There was one problem in this process: sample #303 was said to be from event #98 in the CHL spreadsheet, but the log book indicates it is from cast #97 and there is no rosette log sheet for that sample.  The station name is the same for events #97 and 98, but the latter was a net cast. The event number was changed to #97.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. There were a few problems including the rosette sheet indicating that sampling occurred during event #81 but it was really 80. The CHL samples were missing from 4 events because they had been misnamed. There were often CTD and NET casts at the same station, and the CHL are from the ROS files. They were corrected.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There is a lot of scatter in the comparison, but if cases are excluded if the standard deviation in the CTD salinity is >0.001, most outliers are removed.

There are 2 major outliers remaining as follows:

- Event #1- This bottle was fired at the bottom of the cast in Saanich Inlet. This was the only outlier for which the bottle value was lower than the CTD salinity and this makes sense when a bottle is fired at the bottom of the cast; if the flushing of the bottle is incomplete it would contain water of lower salinity than expected. The descent rate in Saanich Inlet was very steady which likely reduces flushing efficiency. Sampling is not recommended at the bottom of a cast and it is assumed this was a mistake as it was not done during the rest of the cruise.

- Event #4 – This bottle has a high standard deviation in the CTD salinity and there is a drop of about 0.03 during the bottle stop. Poor flushing is likely given little vertical motion during the stop. The CTD salinity varies by ~0.01 during the stop even after any shed wakes would have disappeared which is typical of Haro Strait. Combining all these factors there is no justification in flagging the bottle.

There are minor outliers that are mostly explained by high standard deviation in the CTD salinity. Those that are not easily explained in that way are:

- Event #42 – Low standard deviation in the CTD salinity; likely incomplete flushing.

- Event #76 – Low standard deviation in the CTD salinity; large shed wake with salinity close to bottle value seen as CTD stopped. Likely incomplete flushing.

To assess the calibration of the CTD salinity we usually look for samples taken in low salinity gradients since this reduces the error associated with poor flushing and casts with noisier descent rates are thought to reduce the flushing problem. We also try to avoid areas of high variability by excluding cases of high standard deviation in the CTD salinity during the bottle stop. There were few samples from this cruise that fit these requirements. Comparisons were run using different methods for selecting data:

- When the large outliers were excluded as well as cases where the standard deviation in the CTD salinity is >0.001, the primary salinity is found to be lower than bottles by an average of 0.0006 and the secondary low by 0.0001.

- We expect some problems with incomplete flushing of Niskin bottles in inland waters especially when the descent rate is quiet. For the 2 bottle casts near the mouth of Juan de Fuca Strait the descent/ascent rates are very noisy, the primary CTD read higher than bottles by 0.0018 and 0.0009 and the secondary read higher by 0.0023 and 0.0018.

- Using bottles below 200m (excluding one outlier) the primary is high by an average of 0.0007 and the secondary high by 0.0012.

- For 5 casts below 300db the primary is high by an average of 0.0008 and the secondary high by 0.0010.

- The 2 casts below 300db with the lowest vertical salinity gradient indicate that the primary is high by 0.0025 and 0.0015 the secondary high by 0.0027 and 0.0018.

The most reliable comparisons show that both sensors are reading high and the differences between the sensors are similar to the results found in section 9.

For full details for the COMPARE run see file 2015-17-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There is a lot of scatter in the fit. The only major outlier comes from the low DO waters in Saanich Inlet where the DO sensor is always challenged due to a very high gradient; either incomplete flushing or just the distance between the Niskin and CTD could be factors. If other outliers are excluded based on residuals, there is a reasonable fit:

CTD DO Corrected = CTD DO * 1.0282 +0.0553

Fits of the data from near the mouth of the Juan de Fuca Strait (where flushing is likely to be fairly good but vertical gradients are fairly high) produces quite a different fit:

CTD DO Corrected = CTD DO * 1.0406 +0.0247

A fit of points in the central Strait of Georgia (where there are frequent DO reversals and flushing is likely to be incomplete due to steady CTD motion) is different again:

CTD DO Corrected = CTD DO * 1.0146 +0.0993

This comparison may also be affected by the problems noted in the secondary pump, but for the most part the DO sensor would have equilibrated by firing time. So the issue is probably primarily due to flushing variations which are not a reflection of calibration and not relevant to downcast data.

For more detail see document 2015-17-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. The only outlier of note was the one already mentioned from Event #1.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor. There was a large range of CHL values.

A plot of FL/CHL versus CHL does not have the usual shape. Usually we expect the ratio to be high for very low CHL (as high as 2X to 10X). For these data it is rarely above 1. We then expect the ratio to gradually drop as CHL rises. While that is generally the case with these data, the drop goes lower than usual (down to ~0.4) for CHL between 1 and 8ug/L and then rises about 0.7 for the highest CHL. Geography may explain the differences.

A plot of FL versus CHL forced through the origin has a slope of 0.64. Most cases of CHL between about 5 and 10ug/L fall below the trendline; these casts are found in Juan de Fuca Strait and the central Strait of Georgia where CHL values are mostly fairly high. Haro Strait samples are close to the trendline. The northern Strait of Georgia is where the ratio is mostly above the trendline. The fluorescence is often lower on the upcast than for the downcast even after a bottle stop.

For full details of the comparison see file 2015-17-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Based on 2015-01 and 2015-03 which preceded this cruise and had the same equipment, an initial estimate was made that the DO leads the temperature traces by about 2.5s.  ALIGNCTD was run using a setting of +2.5s to advance the DO signal. The results look good. One cast was tested using 2s and 3s to see if either looked better and they didn’t. ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. Tests on previous cruises using these data showed the default setting of (α = 0.0245, β=9.5) did the best job and it does improve the data for both conductivity channels for these data.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. While none are very deep, it does allow us to check for sudden changes.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2014-50-0019 | 330 | ~0 N | -0.00008N | -0.0006 N | Mod, V.Noisy |
| 2014-50-0037 | 325 | ~0 VN | -0.00004 | -0.0004 | Mod, Steady |
| 2015-01-0040 | 325 | -0.0027 | -0.00002 | +0.002 XN | High, Very Noisy |
| “ | 1000 | -0.0013 | -0.00013 | -0.0001 | “ |
| “ | 3000 | -0.0007 | -0.00013 | -0.0011 | “ |
| 2015-01-0093 | 325 | -0.0008 | -0.00011 | -0.0005 | High, Very Noisy |
| “ | 1000 | -0.0008 | -0.00015 | -0.0010 | “ |
| “ | 3000 | -0.0006 | -0.00016 | -0.0014 | “ |
| 2015-03-0017 | 600 | +0.0001 | +0.00003 | +0.0004 | Moderate, steady |
| 2015-03-0027 | 450 | +0.0004 | +0.00005 | +0.0003 | “ |
| 2015-03-0075 | 350 | +0.0004 | +0.00004 | +0.0003 | “ |
| 2015-03-0087 | 450 | +0.0003 | +0.00004 | +0.0002 | “ |
| 2015-17-0058 | 325 | +0.0008 | +0.00005 | -0.0001 | High, Steady |
| 2015-17-0068 | 300 | +0.0006 | +0.00004 | +0.0001 | High, F.Noisy |
| 2015-17-0083 | 300 | +0.0010 | +0.00006 | -0.0002 | High, Steady |

The differences are all small.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

There are no station names in the headers because of a format error in the SeaBird headers, so a spreadsheet was prepared with files names and station names as found in the log book; this was used with routine “MERGE: CSV file to Headers”.

##### Checking Headers

A cross-reference list was checked against the log book and no problems were found.

A header check was run. No problems were detected.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.0db which is reasonable for the Vector, though perhaps slightly low. The lowest pressure found was about 0.5db and the associated salinity was ~15, so this looks about right.

The altimeter and water depth readings from the headers of the MRH and SAMAVG files were exported to a spreadsheet. Plots were made of altimetry near the bottom for most casts and the headers look appropriate for both sets of files.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 4 casts using a variety of shifts and the best results were found with -0.5 records for the primary and -0.6 records for the secondary.

SHIFT was run on all casts using -0.5 records for the primary conductivity and -0.6 records for the secondary conductivity.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. Tests were run using values between +30 records and +70 records and +50 looks like the best choice overall.

SHIFT was run on the pH:SBE channel using a setting of +50 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

Conductivity, pressure and dissolved oxygen sensors were all recalibrated in late 2013 or early 2014. They were used for 1 cast during 2014-19 and all of 2014-50, 2015-01 and 2015-17. Both T/C sensor pairs produced salinity within 0.001 of bottles for 2014-50, while from 2015-01 it appeared that the primary salinity was high by ~0.003 and the secondary high by ~0.0015. For 2015-03 the primary salinity was found to be low by 0.0036 and the secondary low by 0.0029 but there was evidence of poor flushing and fairly high vertical salinity gradients.

The pressure sensor was found to have drifted lower by about 1.25db during 2014-50, 2015-01 and 2015-03.

For 2014-50 the dissolved oxygen sensor was corrected using a linear fit of slope 1.0281 since there was too little sampling of waters with low DO values to estimate an offset. For 2015-01 the fit used had a slope of 1.0187 and offset of +0.056mL/L. For 2015-03 the results were similar to 2015-01 with a slope of 1.0147 and offset of +0.0647mL/L. The lower slope may have been due to incomplete flushing since the sensor drift leads to DO values reading low, but if the samples were from lower in the water column they will generally be reading low too. The secondary pump problem may have affected the fit as well, but it does not appear to have been a large effect.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. The salinity was low at depth near the mouth of Juan de Fuca Strait and near the surface in many of the casts, especially in the central Strait of Georgia. Low near-surface salinity has been observed often in 2014/2015. The temperature is mostly within the climatology but slightly high around 80-100db at some casts in the Northern Strait of Georgia where the vertical gradients are quite low.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available but the secondary pump was found to need repair after 2015-17.

##### DETAILED EDITING

There is little to distinguish between the primary and secondary T and S channels. However, the CTD technician’s report that there was a problem with the secondary pump after this cruise and evidence that the problem was present for the previous cruise, led to the choice of the primary channels for editing and eventual archiving. The previous cruise had some very sharp gradients that made the problem more obvious.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. All files required some editing.

##### Initial Recalibration

The most reliable comparisons show that both sensors are reading high. The history of these sensors is not terribly helpful since 2 other cruises were in waters even more prone to flushing problems and the best comparison, from Line P cruise 2015-01, was noisier than usual. When 22 bottles were fired at 2000m the primary was found to be high by an average 0.0029 and the secondary by 0.0016. It is noted that the primary was higher than the secondary in Line P but the opposite is true for this cruise. It will be seen in section 9 that there was a change in sign in the salinity differences between 2015-01 and 2015-03. This may be related to the problem found in the secondary pump after 2015-17. The differences between the primary salinity and the bottles fired in low gradient waters are reasonably close to the 2015-01 comparison.

It is most likely that the primary salinity is reading high, with the 2 most reliable cases showing it high by 0.0015 and 0.0025. For 2015-01 an estimate was made that the primary salinity was high by 0.003.  For this cruise a correction of -0.002 will be applied.

Pressure does not need recalibration.

For dissolved oxygen the large scatter in the fits with quite different results in different parts of the cruise is probably due to poor flushing of Niskin bottles and possibly problems with the secondary pump though that seems less likely to have affected the CTD while stopped. Using the results of cruise 2015-01 looks to be the wisest choice though the comparisons for that cruise were also rather noisy.

CALIBRATE was run using file 2015-17-recal1.ccf to subtract 0.002 from the Salinity:T0:C0 channel and to correct the Oxygen:Dissolved:SBE channel in the SAM and MRGCLN2 files using:

CTD DO Corrected = CTD DO * 1.0187 + 0.0569

COMPARE was rerun for salinity and dissolved oxygen using the recalibrated values. The results confirm that the recalibrations were applied properly. (See files 2015-17-sal-comp2.xlsx and 2015-17-DO-comp2.xlsx for details.)

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed, the CTD DO was lower than the bottles by an average of ~0.004mL/L and standard deviation of 0.04mL/L; the results vary with depth with the CTD reading slightly low below 250m. No further recalibration will be applied. See 2015-17-dox-comp3.xlsx for details.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

The CTD files from rosette casts were clipped to 50m; sigma-T was derived and the data were exported to a single file, 2015-17-SOG.csv.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some unstable features near the surface and for a few casts in vertically well-mixed areas. The PAR:Reference channel contains no useful data for any casts. The pH values are extremely low for event #26 (station 64) with values <5.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, PAR:Reference, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and unedited

except that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.13 mL/L from 0 to 30db

±0.1 mL/L from 30db to 100db

±0.08 mL/L from 100db to 250db

±0.03 mL/L below 250db

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see the report: 2015-17-proc.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values were low ranging from 67% to 138%. The highest value was in Saanich Inlet and the lowest was at a near-shore cast between Nanaimo and Baynes Sound. Casts in Haro Strait and eastern Juan de Fuca Strait had low saturation rates, as usual, due to strong vertical mixing. Some casts in the middle of Juan de Fuca Strait had values that were slightly high. Most values were between 90% and 110%. There is no indication of a problem with the dissolved oxygen calibration.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, PAR:Reference, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check and a header check were run on all files and no errors were found.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

##### Thermosalinograph Data

There were no loop samples, flow meter or intake thermistor.

The only method to check calibration is to compare with the CTD casts. There are 3 issues that complicate this.

- Time channel – Is there a time offset or is time drifting?

- Depth of sampling – The intake is about 2m below the surface, but CTD data is not always available from that shallow and the pump may suck water in from a little higher than the intake level.

- Pump pressure – This is variable, so the time in the loop may be vary and affect heating.

Only the 1st of these issues can be established easily. More experience and loop sampling are needed to begin to address the other questions.

a.) Checking calibrations

There were many small test files.

All casts had non-standard names.

Most of the configuration files were identical with the only differences in the others being the failure to enter the TSG serial number in the headers.

One configuration file was renamed as 2015-17-tsg.xmlcon. No errors were found in the calibration parameters.

b.) Conversion of Files

The 9 files with proper configuration files were converted to CNV files using configuration file 2015-17-tsg.con. The names were changed to standard format. They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots were produced. Those from files 7, 8 and 9 look good. Files 1 and 2 contain no temperature or salinity data, though the positions look fine. Files 4 and 5 contain only a few records each. Files 3 and 6 might have some useful data but they are short.

The track plots mostly look fine, though there are large gaps in the early portion and about half of file #3 has positions that got “stuck”.

c.)  Checking Time Channel

This issue is critical before doing any other comparisons.

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.3db of 2db and exported to a spreadsheet which was saved as 2015-17-ctd-tsg-comp.xls. There were 47 casts which overlapped with TSG files.

The 5 TSG files with data that overlapped with CTD data were opened in EXCEL, median and standard deviations (over 5 records) were calculated for lab temperature and salinity and the files were reduced to the times of CTD files.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences generally started small and grew with time in files #3, 6 and 7. For file #8 the comparison of positions was generally ok, though the last cast had slightly high differences. However, the explanation probably lies in the site of the cast which was at station 21 in a strong current; there was considerable variation in the CTD position around the time given in the header of the CTD file – enough to explain the differences easily. Similarly for file #9 the final cast shows somewhat higher differences in positions but examination of the file suggests that the ship had not quite stopped when acquisition started.

A note in the log says “TSG working” as file #8 began and it does look as though files 8 and 9 are reliable and the others should not be archived. A track plot of files #8 and #9 was added to end of this report.

d.) Comparison of T and S from TSG and CTD data

- TSG vs CTD The spreadsheet comparing CTD and TSG files was then examined to find the differences between the salinity and temperature channels for the CTD and the TSG. The differences are higher for the last 5 casts which were in narrow channels or close to shore.

1. LAB TEMP The lab temperature was higher than the CTD temperature by a median value of 0.1C° even if the last 5 casts are excluded. The standard deviation is 0.113C° using all data and 0.025C° excluding 5.

2. SALINITY The TSG salinity data are lower than the CTD salinity by a median value of 0.009 (standard deviation 0.289) using all casts and 0.008 (standard deviation 0.0185) if the last 5 casts are excluded. When 9 outliers are excluded the TSG salinity is very close to the CTD salinity at 2m (median difference 0.0003). However, the standard deviations are very high in the salinity.

(See 2015-17-ctd-tsg-comp.xls.)

- Calibration History

The TSG temperature and conductivity were recalibrated in February 2013 and there is no record of them having been used since then.

This is the first use of the TSG system recently installed on the Vector.

Conclusions

1. After some initial problems, the TSG clock worked well.

2. The TSG lab temperature was higher than the CTD temperature by ~0.1Cº which is could be due to heating in the loop and/or water coming from slightly higher than 2m.

3. The difference between the TSG Salinity and that from the CTD is highly variable. Depending on how outliers are identified the TSG values are low by a median of ~0 to ~0.009. If it is low, this could be due to the water coming from a little higher than 2m.

4. Matching the TSG and CTD data in time is difficult for this cruise due to the fact that the CTD acquisition appears to have started while the ship was moving for some casts. Examining that in detail is time-consuming, and given other limitations in the comparison it is likely not worth the time it would take. Another complicating factor in the comparison is that the pump pressure is known to vary from time to time and could account for some of the variability noted.

5. Overall the quality of the data was good with little spiking.

6. The Postion:New channel contains many 0 values when the ship is clearly underway. This is not seen in Tully TSG data but whether that is because of higher ship speeds or a different setting on the TSG is unknown. If the latter is the case, then it should be adjusted for the Vector.

7. There are too many variables at work to enable reliable recalibration of temperature and salinity, so only Temperature:Lab will be archived, rather than the usual method of making an estimate of the intake temperature by recalibrating the lab temperature.

f.) Editing

No editing is necessary.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Conductivity:Primary, Flag and Position:New channels. The latter can sometimes be of some interest, but it appears the difference was set so that it frequently appeared the ship was stopped when it was underway.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats. The Temperature:Primary was renamed as Temperature:Lab to make it clear that it is not the intake temperature.

Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

# Particulars

General: 	SPAR not working

pH and PAR reversed channels in configuration file used at sea.

1. 18 bottles fired to check but only 9 sampled

1. SPAR constant value – replaced cable and still didn’t work

7-10. Chains on

16. Chains off

36. Check pressure term in calibration equation

45. TSG working – open outlet valve in engine room

57. Hit bottom, mud on rosette

59. Swapped spigot on bottle #16

83. Didn’t turn off pump, had to power down

87. Current off Campbell River

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 31Jan2013 | Factory |  |  |
| Conductivity | 1763 | 1Jan2014 | Factory |  |  |
| Secondary Temp. | 5013 | 27Feb2013 | Factory |  |  |
| Secondary Cond. | 3394 | 3Jan2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 1438 | 3Jan2014 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2014 |  |  |  |
| pH | 0851 | 4Sept2010 | Factory |  |  |
| SeaPoint Fluor. | 3642 | n/a |  |  |  |
| Pressure Sensor | 506 | 30Dec2013 | Factory |  |  |
| Altimeter | 1252 | n/a |  |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3363       Cruise ID#:	2015-17

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 28Dec13 | Factory |  |  |
| Conductivity | 2487 | 28Dec13 | Factory |  |  |