## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| May 29, 2025 | Cleaned Flag values. SH |
| 27 Mar 2025 | Updated channel names & formats in TOB files. |
| 17 July 2020 | HPLC data added to CHE files. G.G. |

## PROCESSING NOTES

Cruise: 2015-18

Agency: Ocean Sciences Division

Location: Strait of Georgia and Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Strait Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: June 14, 2015 – June 29, 2015

Processed by: Germaine Gatien

Date of Processing: 3 July 2015 – 14 October 2015

Number of original HEX files: 81

Number of CTD files: 81	Number of bottle casts:	25

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0506) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#1438), a SeaPoint Fluorometer (#3640), a Biospherical QSP-400 PAR sensor (#4565), a pH sensor (#0692) and an altimeter (#62354).

The data logging computer was Vector CTD Laptop.

The data acquisition program was Seasave.

The CTD deck unit was an SBE model 11+, serial number 0619.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log book was generally in good order, but there was information missing on some sampling log sheets. Since the analysts depend on the sampling logs, problems should be noted there - for example, if there are no samples because there was no water in the Niskin bottle. The analyst will then enter a pad value for that sample rather than waste time trying to find the missing sample. In one case there appears to have been a deliberate change to planned chlorophyll sampling with no note made on the sampling log. A correction was made to at least one sheet by erasing the event # and replacing it. Since labels may have been prepared using the earlier entry, it is better to put a line through the error and enter the correction so we can understand discrepancies.

In the raw files headers, “Stn” was entered where the word “Station” should be used, so in conversion to IOS HEADERS the station names were not picked up. A spreadsheet was used to merge this information with the other location headers.

There were many problems in the first half of the cruise. There were 12 data files with the wrong names; these always occurred after a non-CTD event when the hex file was not given the same name as entered in the log book. The combination of that error and not having station names entered correctly in the files led to considerable confusion and could have led to the conclusion that some data were lost.

There were several bottle misfires, many reports of leaking bottles and one case of sampling from the wrong bottle.

Cast #1 lacked NMEA data so latitude and longitude were entered based on the log record.

This cruise provided an opportunity to test whether choosing a single recalibration of channel Oxygen:Dissolved:SBE for the whole cruise is appropriate in areas where variations in the comparison of bottles and CTD DO appear be geographic or temporal in nature. Selecting casts in rougher waters produced a result similar to that found when the whole data set was used, while that for quieter conditions differed significantly. Most outliers in the full comparison came from quiet areas where poor flushing of Niskin bottles is likely. So the geographic variations likely do not reflect calibration change, but rather variations in the efficiency of flushing. A single recalibration algorithm is appropriate because the variations reflect sampling issues, not the CTD DO sensor performance.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±1.0 mL/L from 0 to 25db

±0.2 mL/L from 25db to 100db

±0.1 mL/L from 1000db to 200db

±0.06 mL/L below 200db

The salinity comparison was very noisy with almost all differences showing the CTD salinity to be lower than bottles. Poor flushing of bottles for some casts and evaporation due to a long wait for sample analysis would both lead to higher bottle values. Picking out a few casts for which flushing was likely reasonably good, the primary and secondary CTD salinity channels were found to be lower than bottles by about 0.003. If there was a small amount of evaporation or desorption, say ~0.002, then the CTD salinity is likely within 0.002 of bottles. No recalibration was applied.

The CTD fluorometer values were higher than the extracted CHL for CHL<1ug/L, but overall fluorescence was about 65% of the extracted chlorophyll. The noise in the comparison may be partly due to flushing issues, but since some samples are from above the CHL maximum and others below, the effect will not be systematic.

A thermosalinograph was in use. There was no flow meter or intake thermistor limiting our ability to judge performance. There are large differences between the TSG temperature and salinity from those seen in overlapping CTD data at 2m. This appears to be due, primarily, to a mismatch in depth between the two sets of observations which causes very large differences in the presence of the high near-surface gradients seen at many sites. Notes in the log indicate that the flow was off for a section early in the record and there are other sections where data values barely change or have 0 values, so there must have been other interruptions in the flow. Temperature and salinity data were replaced with pad values where there is evidence of flow stopping. No recalibration was applied.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data were available at the time of processing. Calibration is required for each cast to get absolute values although

general trends within a cast are likely real.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX. All files had the wrong cruise number in the file names. Those were corrected.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked.

A problem was found and repaired in the secondary pump before this cruise.

The XMLCON files did not change through the cruise except that for the first cast there was no NMEA download. The calibration constants were checked for all instruments and the only correction made was to the date format for the transmissometer. Based on the results of the last 4 cruises that used this CTD (#506), the pressure offset was updated by using -0.25.

The corrected file was saved as 2015-18-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2015-18-ctd.xmlcon to create CNV files.

A few casts were examined.  All expected channels are present. The upcast and downcast temperature and conductivity channels are reasonably close, with the upcasts generally noisier than downcasts. The primary and secondary are also reasonably close though there are occasionally significant differences. There are spikes in the primary conductivity. Dissolved oxygen, fluorescence, PAR, pH, altimetry and transmissivity profiles all look normal. As usual for this region the descent rate is high and steady in inland areas and very noisy with some complete reversals of direction for some casts near the mouth of Juan de Fuca Strait.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2015-18-ctd.xmlcon.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

The position information was added to file 2015-18-0001.BOT because NMEA had not been acquired.

Temperature and salinity were plotted for all BOT files to check for outliers and none were found.

File names 11, 28, 29, 34 were changed to 12, 29, 30, 36 to match log event numbers; see §11 for details.

A preliminary header check was done and no problems were found. Fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records. There were 10 sample numbers skipped. All bottles were fired during event #1 but only the first 10 Niskin bottles were sampled.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine. Those files were bin averaged on bottle number.

For event #1 Niskin bottles 11-24 have no sample numbers and were removed from the SAMAVG file.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2015-18-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2015-18chl*.xls. The file included comments, flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2015-18chl.csv which was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2015-18oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2015-18oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2015-18SAL.xls. The analysis was done within 57-62 days of collection. The files were simplified and saved as 2015-18sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2015-18nuts.xls.

Then the file was simplified, reordered on sample numbers and saved as 2015-18-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. There were a few discrepancies:

- Event #4 – The oxygen sample was identified as from event #5. This was fixed.

- Event #16 – The nutrient data for samples 62 and 63 are missing. The DO s/s shows that #63 was flagged 9 with the comment that the Niskin bottle was empty, so flag 9 was added to the nutrient data as well. For sample 62 there is no explanation. Tentatively flagged 1 with a comment that the sample is missing.

- Event #51 – The oxygen sample was identified as from event #50. This was fixed.

- Event #77 – The chlorophyll samples shown on the rosette sheet were samples 223, 224 and 226 but the samples in the spreadsheet are 223, 225 and 226. The labels are in agreement with the spreadsheet, so no change was made.

- Events #93 and 98 – Only phyto sampling but CHE file will be produced.

After the fixes the merge process was rerun and no further problems were found.

CLEAN was run on the MRG files to add 0 flags to empty flag channels and to update header limits. (MRGCLN2)

##### Compare

Salinity

Compare was run with pressure as reference channel.

There is a lot of scatter in the comparison. There are 2 major outliers remaining as follows:

- Event #1- This bottle was fired at the bottom of the cast in Saanich Inlet. All samples look like they are from near the surface. The sample was replaced with a pad value and flag 5.

- Event #8 – Sample #26 appears to be a mis-sample. Other samples from this bottle (nutrients and DO) fit their profiles very well and so this does not look like a misfire or leaky bottle. The sample looks as if came from a little below 100m, lower than Niskin #6. Niskin #1 misfired at about the right depth to make sense for this sample. So we likely have a combination of mis-sample and misfire. The sample was replaced with a pad value and flag 5.

The scatter in the comparison would appear to have 2 sources:

- There was a long delay in analyzing the salinity samples so that either evaporation or desorption may be significant; evaporation is likely to be fairly random and we might expect an effect of ~0.002. For desorption the effect is thought to be ~0.0024 in 2 months. So these 2 effects do not seem enough to explain differences >0.01 of which there were 5.

- Niskin bottles do not flush perfectly and this is particularly true when the CTD moves very steadily through the upcast. If the vertical salinity gradient is high, the combined effect will lead to CTD salinity looking significantly lower than bottle samples. The 5 casts with the largest differences between the CTD and bottles were examined. Events #6 and 10 had fairly steady descent rate and events #51, 77 and 78 were very steady. The local vertical salinity gradients were fairly high for all but one. The differences were greatest for the shallowest of those 5 samples and least for the deepest.

Unfortunately, there are no useful samples from the cast at the mouth of Juan de Fuca Strait where flushing is likely quite efficient. Casts #12 and 16 in the western part of the Strait are also less likely to have flushing problems and the vertical salinity gradient at the level of the bottles was low. For those the differences were 0.0033 and 0.0034. This may indicate that the CTD is reading low by about 0.003, but if there is also an effect from evaporation/desorption, say ~0.002, then the CTD salinity is likely well within 0.002 of the bottles. For the secondary sensors the differences for those two bottles were both 0.0026. The difference between salinity channels is ~0.0006 with the primary lower than the secondary. This differs from the 2 previous cruises when the secondary was lower than the primary by about 0.0015, but those cruises included some very deep comparisons and there is pressure-dependence in differences.

The differences between the sensors are similar to the results found in §9.

For full details for the COMPARE run see file 2015-18-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There is a lot of scatter in the fit. The most severe outliers were:

- Sample # 8 Event #1 at 10m – Fairly high DO gradient and steady ascent rate, so probably explained by poor flushing as CTD is higher than bottle by 0.2mL/L. No flag justified.

- Sample # 9 Event #1 at 5m – CTD high by 0.9mL/L, very high gradient, high std dev in CTD DO, likely poor flushing or noisy CTD to blame. No flag justified.

- Sample # 25 Event #8 at 252m – CTD lower by 1.0mL/L. Other samples are also out of line. Bottle value should be replaced with pad value and note about misfire.

- Sample # 35 Event #8 at 20m – CTD DO low by 0.4mL/L, so not far off. The std dev in CTD DO is fairly high. No flag justified.

- Sample # 76 Event #19 at 5m – CTD DO low by 0.7mL/L. Already flagged due to leaky vent. Std Dev in CTD DO high and plot shows very large range of CTD DO values during stop. Lot of motion before stop so DO sensor may have had a problem responding, so this may be poor sensor response and/or a problem with the sample. Leave flag as is.

- Sample # 184 Event #62 at 20m – Complex T and DO around 20m – could be slow response of sensor. No flag justified.

- Sample # 224 Event #77 at 10m – Complicated profile with many reversals. Steady ascent rate. CTD DO high by +0.45. Could be flushing problem or slow response of sensor. No flag justified.

- Sample # 248 Event #78 at 5m – CTD DO low by 0.6mL/L, DO reversal. Could be flushing problem or slow response. No flag justified.

When outliers are excluded based on residuals, the fit is:

CTD DO Corrected = CTD DO * 1.0224 +0.026   (R2 = 0.68)

When the data were broken down into subsets based on region, the fits were:

Western Juan de Fuca:        CTD DO Corrected = CTD DO * 1.0211 +0.0344   (R2 = 0.80)

Southern Strait of Georgia: CTD DO Corrected = CTD DO * 1.0236 +0.0164   (R2 = 0.71)

Northern Strait of Georgia: CTD DO Corrected = CTD DO * 1.0191 +0.0406   (R2 = 0.29)

This cruise provides a good mix of conditions, so we can study whether it is reasonable to use a single fit for a whole cruise. The fit in the northern Strait of Georgia is quite different from that in the southern part of the Strait and the one for the whole data set? Should it be recalibrated using a separate fit? One explanation for such variations is that the Niskin bottles do not flush completely so that the Niskin contains water from deeper than that measured by the CTD (though this error is offset slightly by the CTD being below the Niskin bottle). If the local vertical DO gradient is low, this is not significant, but in the presence of a large DO gradient it will be. It seems likely that the flushing will be better when the Niskin bottles is bobbing around, and the descent/ascent rate of the CTD gives a notion of how much ship motion there is.

Descent rates were examined and the casts were broken into 2 groups, those having generally steady descent rates and those with noisier descent rates, with the same number of casts in each. Plots were made that would suggest both groups have about the same number of casts with very large DO gradients near the surface and there are deep casts in both groups, though more in the “noisy” group. The fits for the two groups were;

Steady descent rate: CTD DO Corrected = CTD DO * 1.0183 +0.0366   (R2 = 0.47)

Noisy descent rate:  CTD DO Corrected = CTD DO * 1.0234 +0.0260   (R2 = 0.78)

These results suggest that flushing of Niskin bottles is a factor that must be considered in recalibrating dissolved oxygen. The low R2 values for the steady cases is likely due to the scatter caused by the varying effects of poor flushing depending on local gradients. The fact that the “noisy” result is close to the overall result is also likely because a majority of the outliers removed came from the casts with steady descent rate.

Using the overall fit or the one from the “noisy” casts looks best. As will be discussed later, those results are also close to those from 2 offshore cruises that preceded this one and used the same sensor. A single recalibration algorithm is appropriate because the variations are due to sampling issues, not the CTD DO sensor performance.

For more detail see document 2015-18-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor. There was a large range of CHL values.

A plot of FL/CHL versus CHL has the usual shape. The ratio is not as high at the low CHL end, but that may be because there were no extremely low CHL values. At the high end the ratio settles to about 0.7 with the exception of the largest CHL value when it was about 0.3.

A plot of FL versus CHL forced through the origin has a slope of 0.64. When 2 outliers are excluded the slope is 0.68.

For full details of the comparison see file 2015-18-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Based on 2015-20 results ALIGNCTD was run using a setting of +3s to advance the DO signal. The results look good. Two casts were tested using 2.5s and 3.5s to see if either looked better and they didn’t. ALIGNCTD was run on all casts using +3s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. Tests on previous cruises using these sensors showed the default setting of (α = 0.0245, β=9.5) did the best job and it does improve the data for both conductivity channels for these data.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. While none are very deep, it does allow us to check for sudden changes. Differences from earlier cruises using the same equipment are also shown with dark shading.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2015-01-0040 | 1000 | -0.0013 | -0.0001 | -0.0001 | High, V Noisy |
| “ | 3000 | -0.0007 | -0.0001 | -0.0011 | “ |
| 2015-01-0093 | 1000 | -0.0008 | -0.0002 | -0.0010 | High, V Noisy |
| “ | 3000 | -0.0006 | -0.0002 | -0.0014 | “ |
| 2015-03-0017 | 600 | +0.0001 | +0.0000 | +0.0004 | Moderate, steady |
| 2015-17-0058 | 325 | +0.0008 | +0.0001 | -0.0001 | High, Steady |
| 2015-20-0042 | 600 | -0.0018 | -0.0001 | +0.0006 XN | High, X Noisy |
|  | 1000 | -0.0010 | -0.0001 | +0.0004 N | “ |
|  | 1900 | -0.0008 | -0.0002 | -0.0010 | “ |
| 2015-20-0078 | 600 | -0.0015 | -0.0001 | +0.0009 | High, Noisy |
|  | 1000 | -0.0011 | -0.0001 | -0.0003 | “ |
|  | 1900 | -0.0010 | -0.0002 | -0.0007 | “ |
| 2015-20-0081 | 1000 | -0.0011 | -0.0001 | -0.0003 | High, Noisy |
|  | 1900 | -0.0009 | -0.0001 | -0.0010 | “ |
| 2015-20-0111 | 1000 | -0.0007 | 0 | +0.0010 | High, Noisy |
|  | 1900 | -0.0008 | -0.0001 | -0.0007 | “ |
|  | 2400 | -0.0007 | -0.0002 | -0.0012 | “ |
| 2015-09-0018 | 1000 | -0.0009 | -0.0001 | -0.0006 | High, V.Noisy |
| “ | 1900 | -0.0007 | -0.0002 | -0.0016 | “ |
| 2015-09-0041 | 1000 | -0.0015 | -0.0002 | -0.0010 | High, X Noisy |
| “ | 1900 | -0.0009 | -0.0002 | -0.0016 | “ |
| “ | 3000 | -0.0009 | -0.0002 | -0.0021 | “ |
| 2015-09-0104 | 1000 | -0.0007 | -0.0001 | +0.0006 | High, Moderate |
| “ | 1900 | -0.0012 | -0.0002 | -0.0008 | “ |
| 2015-18-0008 | 230 | -0.0017 | -0.0010 | +0.0008 | High, Noisy |
| 2015-18-0056 | 345 | -0.0024 | -0.0002 | +0.0002 | F.High, F.Steady |
| 2015-18-0062 | 335 | -0.0019 | -0.0001 | +0.0002 | High, F.Noisy |
| 2015-18-0085 | 350 | -0.0025 | -0.0003 | -0.0005 | High, F.Steady |

The differences are all reasonably small. The temperature differences are a little higher and the salinity differences lower than for previous cruises, but the differences in depth are significant. Salinity differences tend to increase with pressure and certainly did for cruises 2015-20 and 2015-09. Temperature differences tend to decrease with pressure for those 2 cruises as well.

Differences were examined in shallower water for 2015-09 and the temperature differences were similar at ~0.002 at 350m, but the conductivity and salinity differences were higher at ~0.0002 and 0.0015. These differences could be due to higher gradients at that depth in the earlier cruise or differences in sensor alignment or pump performance, or there may be a real change in calibration. The salinity comparison does confirm that the two channels were closer than has been noted in other recent cruises.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

The NMEA positions were not acquired for event #1, so after conversion values were entered in the header based on the log entry.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

There are no station names in the headers because of a format error in the SeaBird headers, so a spreadsheet was prepared with files names and station names as found in the log book; this was used with routine “MERGE: CSV file to Headers”.

##### Checking Headers

A cross-reference list was checked against the log book. There were many discrepancies found:

- File #9 has the position of event #10 – station 103. Change file name to #10.

- File #10 has the position of event #11 – station 76. Change file name to #11.

- File #11 has the position of event #12 – station 75. Change file name to #12.

- File #12 has the position of event #13 – station 74. Change file name to #13.

- File #28 has the position of event #29. Change file name to #29.

- File #29 has the position of event #30. Change file name to #30.

- File #30 has the position of event #31. Change file name to #31.

- File #31 has the position of event #32. Change file name to #32.

- File #32 has the position of event #33. Change file name to #33.

- File #33 has the position of event #34. Change file name to #34.

- File #34 has the position of event #36. Change file name to #36.

- File #35 has the position of event #37. Change file name to #37.

Corrections were made to the files after the running of DERIVE. Convert to IOS Headers, CLEAN and MERGE:CSV file to Headers routines were rerun. A new cross-reference list was produced and checked and no further problems found.

A header check was run. There are clearly some spikes in temperature and salinity for some casts; the most severe case was checked and was found to have occurred at the surface, so it should not be a problem.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.4db which is reasonable for the Vector.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and as usual, there is a lot of variability with the up and downcast traces sometimes closer than temperature and sometimes further apart. This is likely due to varying vertical gradients. Overall the choice made earlier looks appropriate, so no further alignment will be applied.

Conductivity

Tests were run on 4 casts using a variety of shifts and the best results varied from cast to cast and feature to feature, but overall -0.9 records for the primary and +1.1 records for the secondary looked best.

SHIFT was run on all casts using -0.9 records for the primary conductivity and +1.1 records for the secondary conductivity.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. Tests were run using values between +30 records and +70 record and overall +50 looks like the best choice.

SHIFT was run on the pH:SBE channel using a setting of +50 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

The altimeter and water depth readings from the headers of the DEL and SAMAVG files were exported to spreadsheets. A few checks showed that there were many cases of spikes at the bottom which can easily be misinterpreted by the algorithm. So plots were made of altimetry near the bottom for all casts. The altimetry entry generally looks about right; in one case where there was no header entry, the CTD was near the bottom but it is too hard to pick out a value. The water depths differed in many cases but usually not by more than 2m; in most cases the header looked better but for 6 casts the log entry makes more sense, so changes were made. The same changes were made to 2 casts that were also bottle casts.

##### Other Comparisons

Previous experience with these sensors –

Conductivity, pressure and dissolved oxygen sensors were all recalibrated in late 2013 or early 2014. They were used for 1 cast during 2014-19 and all of 2014-50, 2015-01, 2015-03, 2015-17, 2015-20 and 2015-09.

- Both T/C sensor pairs produced salinity within 0.001 of bottles for 2014-50, while from 2015-01 it appeared that the primary salinity was high by ~0.003 and the secondary high by ~0.0015. For 2015-03 the primary salinity was found to be low by 0.0036 and the secondary low by 0.0029 but there was evidence of poor flushing and fairly high vertical salinity gradients. The comparison for 2015-17 was noisy and both sensors were found to be high by about 0.002. For 2015-20 and 2015-09 also had noisy comparisons with the primary very close to bottles and secondary salinity was low by -0.0015. Repairs were made to the secondary pump after 2015-17.

- The pressure sensor was found to have drifted lower by about 1.25db during 2014-50 and that offset has been used since then.

- For 2014-50 the dissolved oxygen sensor was corrected using a linear fit of slope 1.0281 since there was too little sampling of waters with low DO values to estimate an offset. For 2015-01 the fit used had a slope of 1.0187 and offset of +0.056mL/L. For 2015-03 the slope was 1.0147 and the offset +0.0647mL/L. The lower slope may be due to incomplete flushing since the sensor drift leads to DO values reading low, but if the samples are from lower in the water column they will generally be reading low too. The secondary pump problem may have affected the fit as well, but it does not appear to have been a large effect. For 2015-17 the result of 2015-01 was used because of huge scatter. For the offshore cruises 2015-20 and 2015-09 the slope/offset values applied were 1.0235/0.0248 and 1.0246/0.0452.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. The salinity was low at depth at the southern station at the mouth of Juan de Fuca Strait. Salinity was frequently high near the bottom of casts in the southern Strait of Georgia. Temperature was at or a little above the maximum at the surface around Haro Strait and at mid-depths in the Strait of Georgia. It was well above the maximum around 50m at a few casts in the northern part of the Strait of Georgia. There are many reports of high temperatures for 2014 and 2015 and the salinity outliers are not systematic, so these results do not look symptomatic of calibration problems.

Repeat Casts – There was 1 repeat cast but the site was too shallow to assess repeatability.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

There is little to distinguish between the primary and secondary T and S channels. The primary channels have been used most often since the sensors were last calibrated, but that seemed to be due to a non-sensor problem, the pump performance having been questioned at one point. For this cruise there is no evidence of that sort of problem and overall, the T-S plots look slightly more stable. The secondary salinity was slightly closer to the bottles even allowing for a little evaporation of samples. So, the secondary temperature and salinity were selected for editing.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. All files required some editing.

##### Initial Recalibration

Pressure does not appear to need recalibration.

It is harder to judge salinity calibration, but based on the comparison and history, it is likely that the sensors are within to ±0.002. No recalibration will be applied.

For SBE DO the 2 fits that look most reliable after outliers are excluded based on residuals are:

CTD DO Corrected = CTD DO * 1.0224 +0.0260 – using all casts

CTD DO Corrected = CTD DO * 1.0234 +0.0260 – using casts with “noisy” ascent rates.

The results are very close, so it doesn’t make much difference which we choose. The “noisy” choice had a slightly higher R2 value and a slope closer to those found for the 2 previous cruises using this sensor, so choosing that one makes sense.

CALIBRATE was run using file 2015-18-recal1.ccf to correct the Oxygen:Dissolved:SBE channel in the SAM and MRGCLN2 files using:

CTD DO Corrected = CTD DO * 1.0234 + 0.0260

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly. The average of differences in the fit once outliers were removed was +0.0038mL/L, but the standard deviation is 0.0188mL/L. This fit included all casts.

See file 2015-18-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed, the CTD DO was higher than the bottles by an average of ~0.02mL/L and standard deviation of 0.02mL/L; the results vary with depth with the CTD reading low right at the surface, high between 10m and 100m and close to bottles below 100m. This pattern is likely due to incomplete flushing in the presence of significant DO gradients above 100m and frequent reversals of DO gradient near the surface; slow response time of the CTD sensor could partly account for this pattern as well. No further recalibration of DO is justified.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

The CTD files from rosette casts were clipped to 50m; sigma-T was derived and the data were exported to a single file, 2015-18-SOG.csv.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some small unstable features, but that is normal for this region in near-shore and well-mixed casts. No problems were found.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, pH and PAR data are nominal and unedited

except that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±1.0 mL/L from 0 to 25db

±0.2 mL/L from 25db to 100db

±0.1 mL/L from 1000db to 200db

±0.06 mL/L below 200db

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see the report: 2015-18-proc.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values were highly variable ranging from 70% to 165%. The lowest values were found in Haro Strait and southern Gulf Islands region and at the mouth of Juan de Fuca Strait. The highest values were found in Saanich Inlet, and 1 site in the southern Strait of Georgia and 1 in the southern Gulf Islands and are all associated with high near-surface vertical gradients in DO and the highest fluorescence values of the cruise (at 5 to 10m depth). Most values were between 90% and 130%. Given the wide range in DO gradients and surface saturation, these observations are not helpful in assessing the DO sensor calibration.

##### Final Bottle Files

As was done for the full files, routine MERGE:CSV Files to Headers was used to add station names to the MRGCLN2 files. (MRGMRH)

CALIBRATE was run on those files and then the MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, PAR:Reference, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check and a header check were run on all files and no errors were found.

The track plot looks ok.

Data from the CHE files were exported to a spreadsheet and compared with rosette sheets; this showed that there was a 9 flag missing for one sample. That was added to all files where it should have been entered. The spreadsheet was exported again and saved as 2015-18-bottles_final.csv. No further errors were found.

Plots of each file were examined to ensure no problems had crept in. One dissolved oxygen value looked odd, but the SBE DO data from the full file show the same shape in both downcast and upcast.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

##### Thermosalinograph Data

There were no loop samples, flow meter or intake thermistor.

a.) Checking calibrations

There was 1 TSG file. It had a non-standard name.

The configuration file was renamed as 2015-18-tsg.xmlcon. No errors were found in the calibration parameters.

b.) Conversion of Files

The HEX file was converted to a SeaBird CNV file and the name changed to standard format.

The CNV file was converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots were produced. The temperature and salinity look either bad or very suspicious between about 0600 and 2200 on June 25th. There is a note in the log book saying that the pump had been shut off sometime after 0500 and turned back on at about 1700. The plot shows it was off again between 2000 and 2200.

The track plot looks fine. The plot was added to the end of this report.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 2db and 4db and exported to spreadsheets which were saved as 2015-18-ctd-tsg-comp2.xlsx and 2015-18-ctd-tsg-comp4.xlsx. There were 77 casts which overlapped with TSG files at 4m, and 32 at 2db. These include cases where there was no flow or low flow rate.

The TSG file was opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, lab temperature, salinity and fluorescence and the files were reduced to the times of CTD files and loop samples.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in longitude were all ≤0.0008°; there was only 1 difference in latitude >0.0008°. The average differences were 0.0001° for both. A check was made of the one case where latitude differed by 0.0013°; times and date matched, but there was more motion during the cast than usual based on the positions from the beginning to the end of the CTD cast. This shows both the times and positions are reliable for both systems.

d.) Comparison of T and S from TSG and CTD data

- TSG vs CTD An initial comparison of data from 4m showed the TSG temperature being higher than that from the CTD while the salinity was lower. This suggests the CTD data should come from higher in the water column. This is expected since the intake is at about 2m, but there are more data at 4m so comparisons can be interesting if there are well-mixed surface waters. However, the near-surface gradients are mostly high. The comparison at 2m still showed TSG values looking out of line. There are few records above 2m.  Because some of the early casts probably had a problem with the pump they were excluded from the comparison. The TSG temperature was higher than the CTD at 4m by a median of 0.52C° and higher than the CTD at 2m by 0.28C°. The TSG salinity is lower than the CTD by 0.28 and 0.40 at 2m and 4m, respectively. This suggests that the differences are due, primarily, to water in the loop coming from quite high in the water column, at least while the ship is stopped. Unfortunately, the data above 2m are few and they are noisy. Checking a few casts, no CTD data were found with values close to the TSG temperature and salinity.

(See 2015-18-ctd-tsg-comp2.xlsx and 2015-18-ctd-tsg-comp4.xlsx.)

- Calibration History

The TSG temperature and conductivity were recalibrated in February 2013 and have been used for only 1 other cruise since then. During 2015-17 in the same region the temperature was found to be high by 0.01 and the salinity was close to the CTD, on average.

Conclusions

1. The TSG clock worked well.

2. The TSG lab temperatures were much higher than the CTD temperatures at any depth for which we have data. We usually expect heating in the loop to create some difference, but not as much as observed. Moreover, the water was so warm that the effect should be slight as it was likely close to the ambient ship temperature. Similar observations were found during a recent Tully cruise off the west coast of Vancouver Island, when water actually appears to have cooled off in the loop.

3. The TSG Salinity is much lower than the CTD salinity at any depth sampled by the CTD. While the presence of bubbles can account for low TSG salinity, there is little evidence in the record of a significant problem.

4. The comparison taken together with the history of the sensors appears to support that the problem is not likely due to calibration drift. No recalibration is justified.

5. Because of the likely mismatch in depths of in the comparison, we can’t use the differences to judge whether the pump was working. The very largest differences do appear to occur during a period when the log notes the pump was off, but there are other large differences from other times when there is no hint of a pump problem. Those are also places where the near-surface gradients were large.

6. Other than some obvious problems with the pump, the overall quality of data was high with little of the spiking that might indicate problems with bubbles.

7. No reasonable estimate can be made of the intake temperature because we have no CTD data from high enough in the water column, and the estimate of loop heating from 2015-17 is inappropriate because near-surface temperatures were much higher for this cruise, so that loop heating is likely to be less, or there could even be loop cooling.

f.) Editing

The ATC file was opened in CTDEDIT and a few single-point spikes in salinity were cleaned and sections of temperature and salinity data were removed where there was evidence of little or no flow. One section on June 28th showed little variation, but an examination of the log showed the ship moved little during that time.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Conductivity:Primary, Flag and Position:New channels.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats.

Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

# Particulars

4. File saved as #3 – fixed in raw files.

12. Noticed pump to TSG off at 1700 June 26. Probably turned off after 0500.

17. Bottle 11 empty.

18. Rosette bumped A-frame on recovery – need to use fog lines, no damage.

38. Trip mechanism replaced because latch #11 was sticky – tested okay.

44. Altimeter not working – just says 68.0

45. Altimeter working

72. No oxygen sampling

73. Bottom varies 350-330m

81. No oxygen sampling.

88. Station name wrong – should be 14 (corrected in processing)

97. Communications between computer and CTD interrupted near surface on upcast. Power off to stop pump to bring on board.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 31Jan2013 | Factory |  |  |
| Conductivity | 1763 | 1Jan2014 | Factory |  |  |
| Secondary Temp. | 5013 | 27Feb2013 | Factory |  |  |
| Secondary Cond. | 3394 | 3Jan2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 1438 | 3Jan2014 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2014 |  |  |  |
| pH | 0851 | 4Sept2010 | Factory |  |  |
| SeaPoint Fluor. | 3640 | n/a |  |  |  |
| Pressure Sensor | 506 | 30Dec2013 | Factory |  |  |
| Altimeter | 1252 | n/a |  |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3363       Cruise ID#:	2015-18

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 28Dec13 | Factory |  |  |
| Conductivity | 2487 | 28Dec13 | Factory |  |  |