## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| May 29, 2025 | Cleaned flag values. SH |
| 27 Mar 2025 | Updated channel names & formats in TOB files. G.G. |
| 17 July 2020 | Added HPLC data to CHE files. G.G. |

## PROCESSING NOTES

Cruise: 2015-19

Agency: Ocean Sciences Division

Location: Strait of Georgia and Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Strait Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: September 29, 2015 – October 4, 2015

Processed by: Germaine Gatien

Date of Processing: 27 January 2016 – 4 February 2016

Number of original HEX files: 82

Number of CTD files: 81	Number of bottle files:	23

Number of TSG files: 1

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0443) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#0997) on the secondary pump, a SeaPoint Fluorometer (#3640) on the primary pump, a Biospherical QSP-400 PAR sensor (#4565), a pH sensor (#0692) and an altimeter (#1253).

The data logging computer was Vector CTD Laptop.

The data acquisition program was Seasave.

The CTD deck unit was an SBE model 11+.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log book and sampling/rosette log sheets were generally in good order but there were discrepancies between the labels for extracted chlorophyll sampling and the rosette sampling records.

The times in the file headers are usually earlier than those in the log book. When this is a difference of a few minutes, there is no big problem. But it is often more than 5 minutes and judging by the positions and water depth the ship does not appear to have fully stopped when the file header information is acquired. Drift during casts can account for some differences, but it is noteworthy that larges differences in positions tend to occur when times differ most. For casts #19 and 20, the differences in time were 2 and 8 minutes, respectively. The depths in the log differed from that in the file by 0m and 7m, respectively.  On one occasion the file header information appears to have been acquired during the net cast that preceded the CTD cast, a matter of 30 minutes difference. Fortunately, there was little drift during that cast. This is a relatively new problem and has been noted on other ships and it renders header information less reliable. It is especially troublesome when assessing TSG data because data are selected for the comparison by reading TSG values at the time of the CTD casts. For this cruise the matching of positions was poorer than usual, so the reliability of the comparison with CTD temperature and salinity is reduced.

In the raw files headers, “Stn” was entered where the word “Station” should be used, so in conversion to IOS HEADERS the station names were not picked up. A spreadsheet was used to merge this information with the other location headers.

The comparison of CTD and bottle salinity has a lot of scatter. 7 of the 23 bottles had high standard deviations in the CTD salinity. Another bottle came from the bottom of a cast where comparisons are usually poor. When those outliers were removed from the comparison the primary salinity was found to be low by an average of 0.0081 and the secondary was low by 0.0076. However, the standard deviations in the fits were still high at ~0.006. There was a delay of almost 2 months in analysis of salinity samples which can lead to evaporation and adsorption which could easily increase salinity by 0.005; salt crystals on a few samples may indicate that evaporation would be even higher than usual. Evaporation also explains the scatter that remains even after the 8 outliers were removed. There may also be a small increase in bottle values due to incomplete flushing of Niskin bottles. So the CTD salinity is likely more accurate than it appears. The best evidence from the history of the sensors suggests salinity is within 0.002. No recalibration is justified.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.8 mL/L from 0 to 30db

±0.4 mL/L from 30db to 50db

±0.25 mL/L from 50db to 125db

±0.05 mL/L below 125db

The CTD fluorometer values were higher than the extracted CHL for CHL<1ug/L, but overall fluorescence was about 55% of the extracted chlorophyll.

The chlorophyll analyst noted that the labels for samples from events #1 to #18 indicate that the samples came from 20m, 5m and the surface. The rosette log indicates they are from 20m, 10m and surface. After cast #18 the labels and rosette log are in agreement. Comparison with fluorescence from the one highly stratified of these casts suggests that the samples came from 20m, 5m and the surface. Further, it is highly likely that samplers read the labels just before sampling and followed that instruction. The label entries were used in preparing the bottle files.

There were many problems with the Thermosalinograph during this cruise. The flow was obviously stopped during some parts of the record and in other parts the salinity traces look odd and compare poorly with CTD data. No cause could be determined since there was no loop sampling, intake temperature or flow rate data available, and as noted above, the comparison with CTD data is of lower quality than usual. Flow rate variations and/or fresh water getting into the loop are possible reasons for poor data, or the TSG itself may have malfunctioned. Temperature and salinity data that were clearly bad were removed, but all temperature and salinity data should be treated as nominal. No recalibration was applied; based on the results of an earlier cruise calibration drift is thought to be small. Times and positions are considered reliable.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data were available at the time of processing. Calibration is required for each cast to get absolute values although

general trends within a cast are likely real.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. There was an error in event # for station 59 – it should be 4, not 2.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked.

The XMLCON files did not change through the cruise. The calibration constants were checked for all instruments and the only corrections made were to the date format for the transmissometer and factor E in the dissolved oxygen parameter (based on 2015-10 tests).

The corrected file was saved as 2015-19-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2015-19-ctd.xmlcon to create CNV files.

A few casts were examined.  All expected channels are present. The primary and secondary temperature and conductivity channels are reasonably close during downcasts, but as usual the upcasts differ more due to noise in both channels. Dissolved oxygen, fluorescence, PAR, pH, altimetry and transmissivity profiles all look normal. As usual for this region the descent rate is high and steady in inland areas and very noisy with some complete reversals of direction for casts near the mouth of Juan de Fuca Strait.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2015-19-ctd.xmlcon.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and problems were seen in casts #18 and 21. CTDEDIT was used to clean a few secondary salinity points at 30m for event #18. The output file was copied to the BOT file. The noise in the primary salinity of cast #21 is more complicated and it is not clear which data are good and which not, so no editing was applied.

A preliminary header check was done and no problems were found. Fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records. There were notes in the log book about cast #88: the pump stopped so the cast was rerun. It is clear from the rosette sheet that the sampling was done from the rerun (#89) and the 2 bottles fired during #88 were not sampled, or at least, not analyzed. So cast #88 was removed from the ADDSAMP file.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine. Those files were bin averaged on bottle number.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2015-19-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2015-19chl*.xls. The file included comments, flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2015-19chl.csv, event numbers were added, and the file was then converted to individual CHL files.

The analyst noted that the rosette spreadsheet indicated that sampling was done at 5m for casts 1-18, but the labels indicate it was at 10m.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2015-19oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2015-19oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2015-19SAL.xls. The analysis was done within 52-56 days of collection. The files were simplified and saved as 2015-19sal.csv. That file was then converted to individual SAL files. The sample from station 59 had the wrong event number – should be #4.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2015-19nuts.xls.

Then the file was simplified, reordered on sample numbers and saved as 2015-19-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. No problems were found.

CLEAN was run on the MRG files to add 0 flags to empty flag channels and to update header limits. (MRGCLN2)

##### Compare

Salinity

Compare was run with pressure as reference channel.

There is a lot of scatter in the comparison. There are 7 cases where differences between bottles and CTD were greater than 0.02. Of these one had a high standard deviation in the CTD data and was flagged due to salt crystals around the bottle neck and was unstable during analysis. Two others had a higher standard deviation than we would usually allow into the comparison and one had an extremely high standard deviation. Another was a bottle from the bottom which usually do not compare well. One bottle is likely mislabelled as the sampling log showed it coming from 250m but the label shows it from 300m. That leaves just one outlier without an obvious explanation. That was from event #9 and both CTD channels were lower than that sample by about 0.02. The delay in analysis could easily explain that outlier.

The bottle that was mislabelled was fixed and COMPARE was rerun. When outliers were removed for standard deviation in the CTD salinity >0.0015 and one other outlier that came from a bottle fired at the bottom, the average differences show the CTD salinity being lower than the bottles by 0.0081 for the primary and 0.0076 for the secondary. The standard deviations in the fits were 0.0055 and 0.0056.

It is likely that the salinity from bottles is too high due to evaporation and adsorption which could easily account for a difference of 0.005.

For full details for the COMPARE run see file 2015-19-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There was a lot of scatter in the fit but when the bottles from Saanich Inlet are excluded the fit becomes much tighter. Saanich Inlet is a poor place to do oxygen calibration due to extreme gradients. We expect the CTD to read lower than bottles especially where DO values are at the high end, so cases where the CTD read higher or very close to bottles were examined. CTD DO profiles showed that all such cases came from areas where there were local reversals in DO gradients or very high gradients. Even slight inefficiency in Niskin flushing could easily account for these outliers. When those outliers were removed fits were found using 3 methods. There are no bottle samples with DO<2mL/L, so getting a fit with a reliable offset is a problem.

- With offset forced to = 0 the fit was

CTD DO Corrected = CTD DO * 1.0304   (R2 = 0.32)

- With the offset free

CTD DO Corrected = CTD DO * 1.0217 +0.032   (R2 = 0.38)

- With the offset forced to equal to -0.0309 which was the offset found during 2015-21 which used this equipment and included many low DO samples, the fit was:

CTD DO Corrected = CTD DO * 1.0220 +0.0309   (R2 = 0.38)

And when a few more values were removed based on residuals the fit was

CTD DO Corrected = CTD DO * 1.0216 +0.0309   (R2 = 0.53)

The fit from 2015-21 was

CTD DO Corrected = CTD DO * 1.0225 +0.0309   (R2 = 0.90)

The 2015-19 fit is remarkably close to that when we force the offset to match. There were a lot more data available from 2015-21.

For more detail see document 2015-19-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor. There was a fairly large range of CHL values.

A plot of FL/CHL versus CHL has the usual shape. The ratio at the low end is somewhat lower than usual, but that may be due to the absence of really low CHL values. At the high end the ratio settles to about 0.5 which is typical of this type of sensor.

A plot of FL versus CHL forced through the origin has a slope of 0.56, but the scatter is very large.

For full details of the comparison see file 2015-19-fl-chl-comp1.xlsx.

Data for the first 6 bottle files were examined to see if there is confirmation about whether samples came from 5m or 10m. Only cast #1 has a sufficiently stratified fluorescence profile to provide reasonable evidence. The fluorescence at 10m and 20m is much lower than at 2m and 5m. The chlorophyll sample in question looks reasonably close to the fluorescence at 2 and 5m, but much higher than that from 10m. This supports the labels being correct. Given that samplers generally read labels and then sample accordingly, this confirms that the best choice is to believe the labels.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Based on 2015-21 results ALIGNCTD was run using a setting of +3.5s to advance the DO signal. The results were difficult to judge because the temperature and DO profiles are so complex. Some casts were tested using 3s and 4s to see if either looked better, but overall +3.5s looked best.

ALIGNCTD was run on all casts using +3.5s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. Tests on previous cruises using these sensors showed the default setting of (α = 0.0245, β=9.5) did the best job and it does improve the data for both conductivity channels for these data.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. While none are very deep, it does allow us to check for sudden changes. Differences from earlier cruises using the same equipment are also shown with dark shading.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2015-10-0039 | 500 | -0.0012 | -0.00034 | -0.0024 | High, Moderate |
|  | 1000 | -0.0010 | -0.00015 | -0.0005 | “ |
|  | 1900 | -0.0011 | -0.00007 | +0.0003 | “ |
| 2015-10-0070 | 500 | -0.0010 | -0.00016 | -0.0008 | High, Very Noisy |
|  | 1000 | -0.0010 | -0.00006 | +0.0003 | “ |
|  | 1900 | -0.0013 | -0.00002 | +0.0009 | “ |
| 2015-10-0099 | 500 | -0.0011 | -0.00016 | -0.0009 | High, Very Noisy |
|  | 1000 | -0.0010 | -0.00006 | +0.0004 | “ |
|  | 1900 | -0.0013 | -0.00003 | +0.0010 | “ |
| 2015-21-0034 | 1000 | -0.0013 | -0.00014 | -0.0002 | High, Mod |
|  | 1900 | -0.0014 | -0.00006 | +0.0008 | Med, Noisy |
| 2015-21-0068 | 1000 | -0.0013 | -0.00008 | +0.0002 | High, VNoisy |
|  | 1900 | -0.0012 | -0.00004 | +0.0009 | “ |
| 2015-21-0099 | 1000 | -0.0014 | -0.00008 | +0.0004 | High, XNoisy |
|  | 1900 | -0.0013 | -0.00002 | +0.0012 | “ |
| 2015-19-0053 | 350 | -0.0013 | -0.00004 | +0.0007 | High, Steady |
| 2015-19-0075 | 350 | -0.0011 | -0.00003 | +0.0008 | High, Moderate |
| 2015-19-0093 | 350 | -0.0010 | -0.00003 | +0.0007 | High, V.Steady |
|  | 400 | -0.0012 | -0.00002 | +0.0008 | “ |
| 2015-19-0097 | 350 | -0.0004 | -0.00003 | +0.0008 | High, V.Steady |
|  | 400 | -0.0003XN | -0.00006VN | +0.0015VN | “ |

The temperature differences are a little high but they are also very noisy. Conductivity and salinity differences are small.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

There are no station names in the headers because of a format error in the SeaBird headers, so a spreadsheet was prepared with files names and station names as found in the log book with the exception of that shown for event #76 which looks wrong. The station for event #76 was given as #8 in the file header and that makes more sense. There is another event that has station #9 at a position that does make sense. The spreadsheet (2015-19-merge-hdr.csv) was used with routine “MERGE: CSV file to Headers” to add station names to the headers.

##### Checking Headers

A cross-reference list was checked against the log book. The station names were all correct. There are sometimes differences in time of more than 5 minutes and in those cases the positions usually differ more than expected. The positions match well when the times match well. The log entries were likely made just before deployment or acquisition started while header times come from when the CTD system is turned on. In one case the header time corresponds to the net cast that was run before the CTD cast.

A header check was run. No problems were found. No off-scale fluorescence values were found.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.6db which is reasonable for the Vector. During cast #75 the rosette came out of the water briefly after the initial soak at 10m. The CTD itself rose to about 0.6db for about 10swhile and the pumps were on. When the pressure was approximately <0.7 the salinity was <5, but below 0.7db salinity was ~28 in both channels. There were some variations as the rosette was probably tipped so that the two sensors were at different levels, but in general the pressure is consistent with a CTD that was very close to the surface.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and as usual, there is a lot of variability with the up and downcast traces sometimes closer than temperature and sometimes further apart. This is likely due to varying vertical gradients. Overall the choice made earlier looks appropriate, so no further alignment will be applied.

Conductivity

Tests were run on 4 casts using a variety of shifts and the best results varied from cast to cast and feature to feature, but overall -0.2 records for the primary and -0.8 records for the secondary looked best.

SHIFT was run twice on all casts using those settings.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. Tests were run using values between +20 records and +70 records. The best choice varies from cast to cast, but overall +50 looks like the best choice.

SHIFT was run on the pH:SBE channel using a setting of +50 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

The altimeter and water depth readings from the headers of the DEL and SAMAVG files were exported to spreadsheets. A few checks showed that there were many cases of spikes at the bottom which can easily be misinterpreted by the algorithm. So plots were made of altimetry near the bottom for all casts.

There was one cast with an entry based on spikes where the CTD clearly did not get to the bottom so the entry was removed. For 4 casts the algorithm failed due to spikes, but the CTD clearly was near the bottom. In those cases an estimate based on a plot was entered in the headers together with a note of explanation.

As noted earlier there are many cases where the log and header entries differ because the header entry is from earlier before the cast began. It appears that sometimes the ship had not fully stopped. There is further evidence of this in the depth entries which sometimes differ significantly. No attempt was made to change times and positions as this would be extremely time-consuming, but in 2 cases (events 48 & 91) depths are very different so corrections were made since the maximum pressure and altimetry show that the header entries don’t make sense.

##### Other Comparisons

Previous experience with these sensors –

The sensors were used during 2015-10 and 2015-21 which immediately preceded this cruise. The primary and secondary salinity were both found to be low by~ 0.002 for the earlier cruise. They were low by 0.0045 and 0.0032, respectively, for 2015-21, but delayed salinity analysis reduced the reliability of those results. No problems were found with the pressure for either cruise. Dissolved oxygen fits were reasonably similar given very different pressure and DO ranges included in them.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. The salinity was low at depth at the southern station at the mouth of Juan de Fuca Strait. It was also a little low for parts of casts on the east side of the Strait of Georgia including Saanich Inlet and Baynes Sound where this climatology is not considered appropriate. Temperatures were frequently above the climatology maxima. In the northern part of the Strait of Georgia there were often sections of high temperature between 60 and 80db associated with an incursion of warmer, saltier water at gamma ~23.5. In Haro Strait the temperature was a little high around 200db. There have been many reports of temperatures above the climatology in 2014 and 2015, so these are not considered evidence of calibration problems.

Repeat Casts – There was only 1 repeat cast and the pumps were off on the first one.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

There is little to distinguish between the primary and secondary T and S channels. The primary channels have slightly less spiking so those were selected.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. All but 1 file required some editing.

##### Initial Recalibration

Pressure does not appear to need recalibration.

It is harder to judge salinity calibration, but based on the comparison and history, it is likely that the sensors are within ±0.002. No recalibration will be applied.

The recalibration used for cruise 2015-21 looks close to that found for the current cruise and was based on a little more data including many bottles with DO<2mL/L values. The parameters from 2015-21 were selected for recalibration of the 2015-19 DO data. CALIBRATE was run using file 2015-19-recal1.ccf to correct the Oxygen:Dissolved:SBE channel in the SAM and MRGCLN2 files using:

CTD DO Corrected = CTD DO * 1.0225 +0.0309

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly. The average of differences in the fit once outliers were removed was +0.0038mL/L, but the standard deviation is 0.028mL/L.

See file 2015-19-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed, the CTD DO was higher than the bottles by an average of ~0.05mL/L and standard deviation of 0.08mL/L; the results vary with depth with the CTD reading low right at the surface, high between 10m and 125m and close to bottles below 125m. This pattern is likely due to incomplete flushing in the presence of significant DO gradients above 100m and frequent reversals of DO gradient near the surface; slow response time of the CTD sensor could partly account for this pattern as well. No further recalibration of DO is justified.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

The CTD files from rosette casts were clipped to 50m; sigma-T was derived and the data were exported to a single file, 2015-19-SOG.csv.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some small unstable features, but that is normal for this region in near-shore and well-mixed casts. The pH values are very low for cast #68.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, pH and PAR data are nominal and unedited

except that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.8 mL/L from 0 to 30db

±0.4 mL/L from 30db to 50db

±0.25 mL/L from 50db to 125db

±0.05 mL/L below 125db

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see the report: 2015-19-proc.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values were highly variable ranging from 65% to 125%. The lowest values were found in Haro Strait, most of Juan de Fuca Strait, the southern Gulf Islands region and at one well-mixed cast at station 16. The highest values were found in Saanich Inlet, and some sites in the central and northern Strait of Georgia. Most values in the Strait of Georgia were between 90% and 125%. Given the wide range in DO gradients and surface saturation, these observations are not helpful in assessing the DO sensor calibration.

##### Final Bottle Files

As was done for the full files, routine MERGE:CSV Files to Headers was used to add station names to the MRGCLN2 files. (MRGMRH)

CALIBRATE was run on those files and then the MRGCOR1 files were put through SORT to order on increasing pressure.

Two casts were not processed further:

- #5 – This bottle was fired for phyto sampling, but no rosette file is needed. The sample was meant to come from a net cast at the previous station and so CTD data at this site is not useful.

- #88 – The pumps were off so the cast was rerun.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check and a header check were run on all files and no errors were found.

The track plot looks ok.

Data from the CHE files were exported to a spreadsheet and compared with rosette sheets and no errors were found.

Plots of each file were examined to ensure no problems had crept in. One dissolved oxygen value looked odd, but the SBE DO data from the full file show the same shape in both downcast and upcast.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

##### Thermosalinograph Data

There were no loop samples, flow meter or intake thermistor.

a.) Checking calibrations

There was 1 TSG file. It had a non-standard name.

The configuration file was renamed as 2015-19-tsg.xmlcon. No errors were found in the calibration parameters.

b.) Conversion of Files

The HEX file was converted to a SeaBird CNV file and the name changed to standard format.

The CNV file was converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots were produced. The temperature and salinity look bad for much of the record. For the first few hours there was clearly no flow to the TSG. Once flow started the data look reasonable until about 01:00 on October 1 when the temperature starts to rise and salinity looks flat; flow must have stopped.  At about 11:00 the flow starts again and the data look reasonable until October 2 when salinity looks very odd, though temperature might be ok. Salinity continues to look bad until about 0930 on October 3. Another bad patch that looks like flow stopped was between CTD casts #98 and 101 since temperature gradually rises and salinity shows a smooth and slight rise. The end of the file looks ok.

The track plot looks fine. The plot was added to the end of this report.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 3db and exported to a spreadsheet which was saved as 2015-19-ctd3-tsg-comp.xlsx. There were 71 casts which overlapped with TSG files and had data in the 3m bin.

The TSG file was opened in EXCEL, median and standard deviations (over 5 records) were calculated for lab temperature and salinity and the files were reduced to the times of CTD files.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The median differences in latitude and longitude are ~0.0001°. The differences in latitude were all ≤0.0008° but there were 4 differences in longitude >0.001 with the most extreme outlier differing by -0.0025°. This is puzzling as it suggests a problem with the longitude download, but that seems unlikely. The explanation may lie in the discussion of section 11 about differences in times between CTD headers and log book entries. The largest outlier from the longitude differences came from event #67, so the differences between log and file headers for that CTD cast were checked. The latitudes differed by about 0.002° and the longitudes by 0.006°. So the CTD cast is not in the same position as the TSG and the scatter in the differences is due to a mismatch in time, not problems with the NMEA download.

d.) Comparison of T and S from TSG and CTD data

- TSG vs CTD Lines were removed from the spreadsheet for those casts where time-series plots make it obvious that flow had stopped. But the only casts that look convincingly good are casts #4 to #16. For those the TSG lab temperature is higher than the CTD temperature by a median of 0.146C° (standard deviation 0.135C°) and the salinity is lower than the CTD by a median value of 0.300 (standard deviation 0.089). For the casts when the flow was obviously stopped the temperature is much higher than the CTD temperature and the difference is rising with time as we would expect due to heating in the lab. The salinity difference looks random, which again makes sense as TSG salinity would change little while CTD values would vary from cast to cast.

The more problematic part of the record overlaps with CTD casts #28 to #71. The TSG salinity values are much lower than the CTD (with a few exceptions) and the values generally get lower with time. The temperature trace does not suggest the flow had stopped, though there is a lot of noise in the differences. Very low salinity suggests fresh water is getting into the loop which might have a complex effect on the temperature difference. Bubbles in the water might account for low salinity but shouldn’t affect temperature. The low salinity values are seen when the ship is underway and while stopped, so this is not something related to the CTD casts, like shed wakes or other ship effects peculiar to a stop. A malfunction in the TSG itself may account for bad salinity while not effecting temperature much.

For casts #72 to #90 the traces look normal. The TSG salinity is lower than the CTD salinity but by only a median value of 0.033. The standard deviation is high at 0.18. The temperature differences are very noisy and lower than the CTD by a median value of 0.126C°, so a slightly lower difference than for the early casts that were considered good. A higher flow rate might lead to the temperature heating less in the loop and the intake temperature is slightly higher than for the earliest casts, which also leads to less heating. It seems unlikely that a higher flow rate would lead to fewer bubbles and hence higher salinity. However, it is possible that the higher rate causes the intake to draw water from a little lower in the water column leading to higher TSG salinity and lower TSG temperature.

(For more details see 2015-19-ctd3-tsg-comp.xlsx.)

- Calibration History

The TSG temperature and conductivity were recalibrated in February 2013 and have been used for 2 other cruises since then. During 2015-17 in the same region the temperature was found to be high by 0.01 and the salinity was close to the CTD, on average. During 2015-18 the TSG lab temperatures were much higher than the CTD temperatures at any depth for which we had data. The intake water had a temperature close to air temperatures so it was not expected that there would be a lot of heating in the loop. Similarly, the TSG salinity was much lower than the CTD salinity and there was no evidence of bubbles. There could have been some problem with fresh water getting into the system or else the water was being drawn from very close to the surface. A TSG malfunction might be the explanation.

Conclusions

1. The TSG clock is believed to have worked well, but problems with time in the CTD file headers made this study less reliable than usual.

2. There were large sections of data which show clear evidence that the flow had stopped. In other sections the salinity data look bad but the cause must be different – either fresh water getting into the loop or a change of flow rate so that the water entering the loop comes from a different depth than for other parts of the cruise. In parts the salinity is clearly bad and the temperature is not clearly bad; this could be a problem with the TSG itself.

3. There is no reliable evidence from this cruise on which to base a recalibration. The best evidence about salinity is from cruise 2015-17 when the TSG salinity was found to be close to the CTD salinity. For temperature the heating in the loop depends on flow rate and intake temperature. During 2015-17 the TSG temperature was found to be high by 0.01 but the intake temperatures were lower so we would expect less heating during this cruise. It is likely that there is little error due to calibration drift. No recalibration will be applied these data.

f.) Editing

The ATC file was opened in CTDEDIT and data were removed that were clearly bad. In areas where the salinity was very bad but the temperature seemed ok, only salinity data were removed.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Conductivity:Primary, Flag and Position:New channels.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats.

Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

# Particulars

22. Changed the bottom 0-ring on bottle 11

57. Bottle 16 was not tripped

75. The rosette came out of the water briefly at the start.

83. Stayed at bottom for 3 minutes for camera footage.

88. Pump off at depth so brought up and redid cast.

89. Niskin #2 and #4 – valve open no water.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 17Dec2014 | Factory |  |  |
| Conductivity | 2280 | 19Dec2014 | Factory |  |  |
| Secondary Temp. | 2663 | 15Jan2015 | Factory |  |  |
| Secondary Cond. | 2754 | 19Dec2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 17Jan2015 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| Surface PAR | 16504 | 16Mar2014 |  |  |  |
| pH | 0851 | 4Sept2010 | Factory |  |  |
| SeaPoint Fluor. | 3640 | n/a |  |  |  |
| Pressure Sensor | 0443 | 17Dec2014 | Factory |  |  |
| Altimeter | 1252 | n/a |  |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3363       Cruise ID#:	2015-19

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 28Dec13 | Factory |  |  |
| Conductivity | 2487 | 28Dec13 | Factory |  |  |