## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27 Mar 2025 | Updated channel names & formats in TOB files. G.G. |
| 21 July 2020 | Added HPLC data. G.G. |

## PROCESSING NOTES

Cruise: 2015-20

Agency: OSD

Location: West Coast Vancouver Island

Project: La Perouse

Party Chief: Yelland D.

Platform: John P. Tully

Date: 26 May 2015 – 6 June 2015

Processed by: Germaine Gatien

Date of Processing: 5 August 2015 –  2 October 2015

Number of original HEX files: 113		Number of CTD files: 114 (2 files for cast #8)

Number of bottle files: 64			Number of bottle casts processed: 60

Number of original TSG files:  	7		Number of processed TSG files:	 5

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0506) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#1438) on the primary pump, a SeaPoint Fluorometer (#3640) on the secondary pump, a Biospherical QSP-400 PAR sensor (#4565), and an altimeter (#62354).

A thermosalinograph (Seacat 21 S/N 3363) was mounted with a SeaPoint fluorometer (#2356), remote temperature sensor and a flow meter.

The data logging computer was #2. Seasave version V7 22.4 was used for acquisition.

The deck unit was a Seabird model 11, serial number 0508.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

The oxygen kit was IOS kit #2.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log, rosette log sheets and analysis logs were in good order.

There was a serious but intermittent problem with the channels on the secondary pump that suggest that the flow stopped even though the pumps appeared to be turned on. The primary temperature and salinity were selected for archiving (except for cast #89 for which the primary channels looked bad). The fluorometer was pumped on the secondary system and has many spikes at depth, but most disappeared after a cable change. Transmissivity was also spiky at depth but the spikes are not aligned with those in fluorescence.

The fluorometer cable was installed incorrectly for events #85-115, so the Fluorescence:URU:Seapoint channel was removed from those files.

The comparison between salinity bottles and CTD salinity had more scatter than expected and the precision was poorer than usual. While there was a wait of up to 6 weeks for analysis the results do not really suggest a problem with sample evaporation, and since plastic inserts were used in the bottles we don’t really expect that to be a big problem. There were only 7 duplicates. There was some evidence of intermittent problems with the salinometer during the analysis of 2015-09, but later tests of the salinometer detected no problems. While there was more scatter in the 2015-20 comparison that of 2015-09, the average of differences between the CTD salinity channels and bottle salinity were very close. Due to the low precision in the bottle salinity data those values are only given to 3 decimal places. There is no evidence of a problem with the CTD salinity, so those data are given with the usual 4 decimal places.

Problems were found in the following casts:

- Cast #8 – The pumps were turned on at about 50m, so file 2015-20-0008.ctd is missing the top 50m. In order to have some information about the top 50m, a separate file, 2015-20-9008.ctd was prepared using upcast data, but the quality of that data is considered lower than usual because upcast data are corrupted by moving though the wake.

- Cast #14 – The pumps were turned on at 27m. There were too many bottles fired during the upcast to get a reasonable profile, so an upcast file was not prepared.

- Cast #40 – The pumps were turned off between 312.4 and 580.2db. Pumped channels were replaced with pad values from 312.4db to 584.8db using a text editor. Pressure and transmissivity were left unchanged.

- Cast #67 – The pumps were not turned on until the CTD was at about 6.8db.

- Cast #89 – Secondary sensors were selected for this cast due to the primary T/S data looking bad.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±1.0 mL/L   from 0 to 20db

±0.5 mL/L   from 20 to 100db

±0.3 mL/L   from 100db to 250db

±0.08 mL/L from 250db to 400db

±0.04 mL/L below 400db

The Thermosalinograph fluorometer performed poorly with values that looked low compared to the CTD fluorometer and extracted CHL samples from both the loop and near-surface rosette bottles. The values were particularly low when extracted chlorophyll was high. The fluorescence channel was removed. The temperature and salinity traces look fine and the flow rate was very steady. Salinity was found to read lower than that from the CTD by between 0.02 and 0.10 depending on how outliers were identified. The larger differences may be due to the loop drawing water from higher than the intake depth, a significant factor when near-surface waters are not well-mixed as was often the case during this cruise. Bubbles in the loop water may also lead to lower TSG salinity. Salinity was recalibrated by adding 0.025 based on the history of the sensor and examination of parts of the cruise with well-mixed surface waters. This correction may well correct for the effect of small bubbles as well as calibration drift.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Nutrients, extracted chlorophyll, dissolved oxygen, NH4 and salinity data were obtained in QF spreadsheet format from the analysts.

There were many notes including the following that are especially important for processing:

- Event 1- Test firings, no Niskin sampling – no CHE file needed.

- Event 2 - Niskin 1 sampled, all others closed just to test.

- Event 14 - Archiving started at 24metres.

- Events 87-115 - Fluorometer cable installed incorrectly – remove FL from casts 87-115.

- Event 102 - Niskin 1 closed but no bottle file needed.

- Event 163. Started archiving before down to 10 and back – need to remove data from initial drop.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked.

The XMLCON files did not change through the cruise. There was a change in fluorometer cable, but it was done incorrectly so that there are no fluorescence data for events #85 to 115.

The calibration constants were checked for all instruments and the only correction needed was to fix a spelling error. The corrected file was saved as 2015-20-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2015-20-ctd.xmlcon to create CNV files.

A few casts were examined and all expected channels are present. The altimetry, dissolved oxygen, PAR and transmissivity look normal. However, the pumped channels have frequent odd features, mostly in the secondary channels, but occasionally in the primary.

The temperature and conductivity channels are sometimes very close, but often far apart even during downcasts. For some casts it is not completely clear which is bad. There is a lot of noise in the secondary particularly on the upcast near the surface. As usual, upcast traces are generally farther apart than downcasts. But for casts #18 the opposite is true. Usually it is clear which system is at fault, but it is not obvious here. The shape of the secondary traces seems odder and there is a lot of noise in that system. However, if flow were impeded in the secondary system we might expect the secondary temperature to look too high on the downcast and too low on the upcast. For event #18 it is too low on the downcast. For cast is #98 it is too low towards the end of the upcast.

For event #117 the secondary temperature is clearly out of line in both downcast and upcast. The two temperature channels are very close until the CTD reaches ~300m when the secondary temperature shifts upwards by ~0.1C°. Below that the trace is very smooth as temperature falls very gradually. On the upcast the temperature reads too low and by the end of the cast it is too low by ~2.0C°. There is a deep temperature reversal detected by the primary sensor on both downcast and upcast; the secondary sensor shows just a slight suggestion of that feature. The fluorescence looks normal on the downcast. At the point where the secondary temperature begins to look bad, there would normally be no fluorescence signal, so no problem shows up other than a tiny spike. But the upcast fluorescence values are all very low, so something odd was going on there. We might expect higher fluorescence if the flow rate was low, but this looks more like there was no flow.

Cast #34 has an obvious problem on the downcast with the two temperature sensors far apart until about 150m. At about 100m they become very close. The fluorescence is much higher on the upcast than the downcast which is unusual. The trace certainly looks odd with values ~1ug/L around 20m, dropping to ~0.6 from 40 to 70m and then going up to ~1 between 100 and 150m with a sharp drop at 150m. On the upcast the traces has the usual shape and peaks at about 4ug/L at 10m and is <0.5ug/L below 10m. However, the extracted CHL samples look more like the down trace than the up.

Cast #89 has bad primary data from 25m to the bottom and the DO trace is bad too, but the upcast is probably ok for both

There are many spikes in the fluorescence at depths where they are not expected. This was noticed at sea and the cable was replaced, incorrectly for a few casts, then correctly. The change did help.

A quick look at a few casts towards the end of the cruise shows that there continue to be cases where there are significant differences during downcasts where there are high gradients, though no very large differences were noted. An initial look at the cruise that followed showed similar small differences where gradients are high. A few cases were found where there were large differences between the two sensors continued through a bottle stop, but both channels had very noisy data.

The descent rate was kept high, on average, with most casts being moderately noisy and a few extremely noisy with complete reversals of the CTD.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2015-20-ctd.xmlcon.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. CTDEDIT was used to remove a few outliers in the primary salinity in casts #63, 89 and 92 and the secondary salinity in casts #89 and 118. There was quite a lot of noise near the surface of cast #157 but this may be real variability. The output files were copied to *.BOT.

Some ROS files will not be processed because there is no need for a CHE file. This includes event #1 which was just a test, events #102 & 114 had no samples drawn and the only sample from event #42 was for an NH4 blank with no sample number.

A preliminary header check was run; fluorescence values were very high for cast #61, but examination of the full cast shows that it did not go off-scale during the downcast.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Those files were then bin-averaged and called SAMAVG. The lines for bottles #2 to #24 were removed for event #2 as they were just fired to test; there was no sampling.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2015-20-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2015-20chl*.xls. The file included comments and flags and a precision study. The loop samples were removed to a separate spreadsheet, 2015-20-loop-data.csv. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2015-20chl.csv. The csv file was then converted to individual CHL files.  Note that loop sampling was complex with some samples having been through the filter and others not.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2015-20oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2015-20oxy.csv. The spreadsheet was reordered on sample number. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2015-20SAL.xls. The analysis was done within 14-57 days of collection. The files were simplified and saved as 2015-20sal.csv. The loop samples were removed to a separate spreadsheet, 2015-20-tsg-loop-ros-comp.xlsx. File 2015-20sal.csv was then converted to individual SAL files. The precision was lower than usual. The analyst recommended using format F8.3 rather than the usual F9.4.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2015-20nuts.xls. This includes a precision study. The loop samples were removed to spreadsheet 2015-20-tsg-loop-ros-comp.xlsx. The QF file was simplified, reordered on sample numbers, saved as 2015-20-nuts.csv and converted to individual NUT files.

NH4

NH4 data were obtained in file 2015-20_NH4.xls which included a report on precision. The file was simplified and saved as 2015-20NH4.csv and converted to individual NH4 files.

The SAL, CHL, OXY, NUT and NH4 files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files using the Bottle_Number as merge channel.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. Errors found or things to be adjusted were:

- Event 2 – Niskins 2-24 were removed from the SAMAVG file – bottles closed but not sampled

- Event 22 – Removed Niskin 1 from SAMAVG file - not sampled.

- Event 68 – DO was missing because the event # was given as 67 – it should be 68.

- Event 121 – Oxygen duplicate sample 315 not 311 as indicated on rosette sheet. File ok.

- Event 167 – Niskin #1 removed from SAMAVG file – not sampled.

- Event 199 – Niskin #2 removed from SAMAVG file – not sampled.

- Event 216 – Niskins #2-24 were removed from SAMAVG file – bottles closed but not sampled.

The MERGE process was repeated and the data exported until all corrections were made.

##### Compare

Salinity

Compare was run with pressure as reference channel.

The comparison has an unusually large scatter. While analysis took up to 6 weeks, the use of plastic liners in the bottles would be expected to prevent significant evaporation of samples. The analyst noted poor precision based on duplicates.

Excluding bottles from above 450m, the primary CTD salinity was lower than bottles by an average of +0.0001 but the standard deviation was 0.0027; the secondary CTD salinity was low by 0.0015 with a standard deviation of 0.0026. When fits were done with less data by excluding cases with standard deviations >0.001 in the CTD data, or differences between bottles and CTD >0.006, the results are similar. There are insufficient casts with sampling to detect temporal drift.

The 4 most extreme outliers (differences >0.007) were all associated with high standard deviation in the CTD salinity. No quality flags are justified based on this comparison.

The analyst noted that sample bottle #276 from event #115 had a thin translucent residue covering the inside of the bottle.  No explanation has been found for this film which proved to have bacteria all over it when examined under a microscope and could not be dissolved in seawater. The salinity from that bottle was lower than the CTD primary salinity by 0.0033 and lower than the secondary by 0.0013, so it cannot be described as an outlier in light of the scatter found in this comparison.

Precision was fairly low with Sp=0.0044 but there were only 7 duplicates. This might be due to evaporation or desorption raising salinity values in the bottles, a process that is somewhat random from one bottle to another, or there may have been a problem with analysis. During 2015-09 which followed this cruise, analysis was done faster, there were more duplicates and the precision was better than for 2015-20, though not great. There was some suggestion of intermittent problems with the salinometer. The results of the 2 comparisons were very similar – except that 2015-09 had less scatter. So this suggests that evaporation was not a major problem. Tests run later on the salinometer using substandard samples turned up no problems.

For full details for the COMPARE run see file 2015-20-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There are some clear outliers, but most of the data fall into a reasonably tight group. There was a wide range of values with many <2mL/L. When outliers were removed based on residuals the fit was:

CTD DO Corrected = CTD DO * 1.0235 + 0.0248

The 6 most significant outliers are all from areas of high temperature and dissolved oxygen gradients and have fairly high standard deviations in the CTD DO. There is no justification for flagging these samples.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further notable outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor.

There was a large range of CHL values. As usual, fluorescence reads higher than CHL for very low CHL and there is a very large scatter for values <5ug/L. Above 5ug/L the ratio FL/CHL is between 0.45 and 1.9. For full details of the comparison see file 2015-20-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature channels in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on 5 casts to see what setting brings the SBE:DO trace into best alignment with temperature. Settings between +2.5s and +3.5s look reasonable. For 3 other recent cruises using the same DO sensor and mostly the same other sensors, an advance of +2.5s looked best but for the cruise that followed this one and had exactly the same equipment, a setting of +3s looked best.

ALIGNCTD was run on all casts using +3s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. During other recent cruises using the same equipment, the default setting of (α = 0.0245, β=9.5) looked slightly better for both conductivity channels than others tested. One cast was checked for this cruise and the default setting does improve the data.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. A few results from an earlier cruise and one that followed (shaded entries) are shown for comparison, including 1 cruise that did not go very deep. The deepest casts are highlighted in yellow.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2015-01-0040 | 1000 | -0.0013 | -0.0001 | -0.0001 | High, V Noisy |
| “ | 3000 | -0.0007 | -0.0001 | -0.0011 | “ |
| 2015-01-0093 | 1000 | -0.0008 | -0.0002 | -0.0010 | High, V Noisy |
| “ | 3000 | -0.0006 | -0.0002 | -0.0014 | “ |
| 2015-03-0017 | 600 | +0.0001 | +0.0000 | +0.0004 | Moderate, steady |
| 2015-17-0058 | 325 | +0.0008 | +0.0001 | -0.0001 | High, Steady |
| 2015-20-0042 | 600 | -0.0018 | -0.0001 | +0.0006 XN | High, X Noisy |
|  | 1000 | -0.0010 | -0.0001 | +0.0004 N | “ |
|  | 1900 | -0.0008 | -0.0002 | -0.0010 | “ |
| 2015-20-0078 | 600 | -0.0015 | -0.0001 | +0.0009 | High, Noisy |
|  | 1000 | -0.0011 | -0.0001 | -0.0003 | “ |
|  | 1900 | -0.0010 | -0.0002 | -0.0007 | “ |
| 2015-20-0081 | 1000 | -0.0011 | -0.0001 | -0.0003 | High, Noisy |
|  | 1900 | -0.0009 | -0.0001 | -0.0010 | “ |
| 2015-20-0111 | 1000 | -0.0007 | 0 | +0.0010 | High, Noisy |
|  | 1900 | -0.0008 | -0.0001 | -0.0007 | “ |
|  | 2400 | -0.0007 | -0.0002 | -0.0012 | “ |
| 2015-09-0018 | 1000 | -0.0009 | -0.0001 | -0.0006 | High, V.Noisy |
| “ | 1900 | -0.0007 | -0.0002 | -0.0016 | “ |
| 2015-09-0041 | 1000 | -0.0015 | -0.0002 | -0.0010 | High, X Noisy |
| “ | 1900 | -0.0009 | -0.0002 | -0.0016 | “ |
| “ | 3000 | -0.0009 | -0.0002 | -0.0021 | “ |
| 2015-09-0104 | 1000 | -0.0007 | -0.0001 | +0.0006 | High, Moderate |
| “ | 1900 | -0.0012 | -0.0002 | -0.0008 | “ |

The differences are small and show no drift through the cruise. There is also no significant drift over several cruises.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book and no problems were found.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.7db which is reasonable for the Tully. At the end of event #203 the CTD was held very close to the surface with pumps turned on. The primary salinity was about 20 while the secondary was about 30 for pressures very close to 0 (-0.01db to +0.10db). For cast #40 with pumps on there are some spikes to very low conductivity with pressures between -0.2db and +0.05db. For cast #7 the conductivity goes very low just before the pump is turned off at 0.1db. These observations suggest that the pressure readings are reasonably accurate, with in 0.2db. No pressure adjustment will be applied.

There were 2 casts for which acquisition started late – event #8 started at 50m and #14 at 27m. The latter was noted in the log but not the former. Both downcast and upcast files were processed for cast #8 with file names 2015-20-0008.CTD (downcast) and 2015-20-9008.CTD (upcast) since there was only one bottle stop. For #14 there were many stops so the data was too corrupted to be useful.

A header check was run. There were negative values in many channels that appear to be due to spikes near the surface.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet.

Water depths in the headers differed significantly (by >5m) from those in the log book for 8 casts. For events #14, 42, 45, 89, 120, 197 and 199 the log reading looks correct and so the headers in the CLN files were adjusted. For 4 of those casts there were bottle files, so the headers for those were adjusted in the SAMAVG files. For event #130 both readings look unlikely given the altimetry reading, but possibly there was drift between the start of the cast and when the CTD reached bottom. The header entry looks more reasonable than the log, so no change was made.

Finally, all bottle files with sampling only at the surface were checked to ensure there are no altimetry headers entered – sometimes there is a spike at the surface that would give the impression the bottle was from near the bottom. That happened for events #170 and 210; the altimetry header was removed from those two SAMAVG files.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on a few casts to choose settings for alignment of the 2 conductivity channels. For the primary conductivity the best results were found with -0.7 records. For the secondary the best results were with a setting of +0.9 records, but they both look pretty noisy in T-S space with all settings.

SHIFT was run on all casts using -0.7 records for the primary conductivity and +0.9 records for the secondary conductivity.

##### DELETE

Acquisition began too early for cast #163 so the initial drop to 10m is included in the file.  A text editor was used to remove those data so that DELETE will select the 0-10m data from after the soak.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning concerned a noisy section at the surface of cast #203.

##### Other Comparisons

Previous experience with these sensors –

Conductivity, pressure and dissolved oxygen sensors have been used on 5 previous cruises (and one cruise that followed) since they were last recalibrated in late 2013 or early 2014. There was no sampling from the 1st and poor comparisons for the 3rd, 4th and 5th. For the 2nd cruise both T/C sensor pairs produced salinity within 0.001 of bottles. Problems were found in the secondary pump after the 4th cruise.

The pressure sensor was found to have drifted lower by about 1.25db during the first 3 of these cruises so the offset was increased by that amount after that.

The dissolved oxygen sensor corrections from the first 5 cruises varied from about 1.5% to 3% with no obvious drift pattern.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Temperature and salinity values were mostly within the climatology below 100m. There were some cases of temperature being slightly high, mostly near the surface, but occasionally deeper in offshore waters to the north and in the northern Strait of Georgia. Since unusually warm waters have been reported offshore over the past year, higher temperatures are not surprising. Salinity between 30 and 70m showed considerable variability with values below the climatology minimum at some offshore stations with depths >1000m, but they were frequently at or above the maximum for some casts between 100 and 200m depths. The salinity excursions are unusual, but since they go in both directions this does not look like evidence of calibration problems.

Repeat Casts – There are no repeat casts but when the same equipment was used during 2015-09 which followed this cruise, there was excellent repeatability in deep water.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

Overall the primary temperature and salinity channels look best for archiving since the secondary has many obviously bad sections. However, there are some bad primary data too, so a decision will have to be made in the course of editing about which is more useful.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. Salinity was cleaned lightly. Some salinity points were removed – generally these are associated with rapid deceleration of the CTD. All files required some editing. Special steps were needed for the following:

Cast #40 – The pumps were turned off between 312.4 and 580.2db. Pumped channels were replaced with pad values from 312.4db to 584.8db using a text editor. Pressure and transmissivity were left.

Cast #67 – The pumps were not turned on until the CTD was at about 6.8db.

Cast #89 – Secondary sensors were selected for this cast due to the primary T/S data looking bad.

##### Recalibration

During 2015-09 which followed this cruise, there was a lot of deep sampling and analysis was done with less delay than for 2015-20. While the standard deviations in the comparison were much lower for 2015-09 (~0.0015), the differences between bottles and CTD were very close to the ones found for 2015-20 with the primary salinity being high by 0.0005 and the secondary low by 0.0014. This suggests that any evaporation/desorption effect is small and that the poor precision and large scatter is more likely to be related to sampling or analysis. Nonetheless, the bottles may be reading slightly higher, on average, so that the CTD salinity is likely a little higher than the comparison suggests. For 2015-09 the primary salinity was chosen for the archive and was not recalibrated since it was likely within 0.002 and this looks like the best choice for these data as well. For the one cast for which secondary channels were selected (#89), the salinity should be recalibrated by adding 0.0016 to bring it into line with the primary channels.

File 2015-20-recal1.ccf was created to add 0.0016 to Salinity:T1:C1 and to correct the Oxygen:Dissolved:SBE channel using:

CTD DO Corrected = CTD DO * 1.0235 + 0.0248

This correction was first applied to the SAMCOR1 and MRGCOR1 files. COMPARE was rerun and the results confirm that the recalibrations were applied properly. The average of differences in the DO fit once outliers were removed was -0.0006mL/L.

(See file 2015-20-DO-comp2.xlsx for details.)

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed, the CTD DO was higher than the bottles by an average of ~0.018mL/L and standard deviation of 0.025mL/L; the results vary with depth with all differences <0.05mL/L below 500m. In the top 100m there are outliers in both directions, but between 100 and 400m most outliers are cases of the CTD DO being higher than bottles. These come from areas with a high DO vertical gradient so even a slight delay in the sensor response would lead to higher DO from the sensor and even slight inefficiency in the flushing of Niskin bottles would lead to lower DO in the bottles. The rough nature of the comparison and the high variability do not support further recalibration.

See 2015-20-dox-comp3.xlsx for details.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases. Since the fluorometer was not working for casts #85-115 no files were prepared for those casts. Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some slightly unstable features in Saanich Inlet but these are likely real as there is no evidence of shed wakes or poor alignment.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run as follows:

- On all casts except #89 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

- For cast #89 the following channels were removed:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

- Fluorescence:URU:Seapoint channel was removed from casts #85-115.

- The PAR channel was removed from the casts for which the instrument was not mounted as found from the log book.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Although the bottle salinity format was set to F8.3 for this cruise due to the precision study and the scatter in the comparison with the CTD, the CTD salinity itself looks ok, as evidenced by the average difference from bottles being virtually the same as the results for 2015-09. If there is a problem, it appears to be with the bottles.

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and unedited except that

some records were removed in editing temperature and salinity.

The fluorometer cable was installed incorrectly for events #85-115, so the

Fluorescence:URU:Seapoint channel was removed from those files.

For many casts between events #24 and 81 there are frequent spikes in

fluorescence below 60m, but near the surface the values look reasonable.

A graphical editor was used to remove records from obviously bad sections of

temperature and salinity data, and some salinity data were smoothed to

remove small spikes. Where there was corruption of salinity data that

was not seen in temperature data, only salinity points were removed.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

Downcast Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±1.0 mL/L   from 0 to 20db

±0.5 mL/L   from 20 to 100db

±0.3 mL/L   from 100db to 250db

±0.08 mL/L from 250db to 400db

±0.04 mL/L below 400db

For details on the processing see document: 2015-20_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

A cross-reference list was produced.

The sensor history was updated.

The track plot looks fine.

A comparison between the two files for event #8 shows similar shapes but the upcast data are offset somewhat which is the usual difference between downcast and upcast as the upcast travels through the wake.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values ranged from 80% in Haro Strait to 185% at LC03. The Haro Strait value is typical of that well-mixed area. For deep offshore casts the values ranged from 93%-110% range, as did most casts in Queen Charlotte Strait, with values close to 105% being most common in those areas. This looks reasonable as we usually see offshore values of 100%-105%. Higher saturations were found in shallower waters with most between 110% and 130% except for some very high values very close to shore. These values do not indicate a problem with DO calibration, but this is a very rough check.

##### Final Bottle Files

The MRGCOR2 files were put through SORT to order on increasing pressure.

REMOVE was run as follows:

- On all casts except #89 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

- For cast #89 the following channels were removed:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

- Fluorescence:URU:Seapoint channel was removed from casts #87-115.

- The PAR channel was removed from the casts for which the instrument was not mounted as found from the log book.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check and a header check were run on all files. A problem was found in file 2015-20-0216; the only bottle fired was the 24th fired and was Niskin #1 but that is not what was indicated in the addsamp file. This was fixed.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

A cross-reference listing was produced for the CHE files.

Data were exported from the CHE files to file 2015-20-bottles-final.xlsx. The entries were compared with the rosette log sheets to ensure no samples had been missed and no further problems were found.

##### Thermosalinograph Data

There were loop salinity and extracted chlorophyll (both filtered and unfiltered) samples taken, some while stopped and some while underway.

a.) Checking calibrations

The 7 configuration files were identical. One file was renamed as 2015-20-tsg.xmlcon. No errors were found in the calibration parameters.

b.) Conversion of Files

The 7 files were converted to CNV files using configuration file 2015-20-tsg.con.

They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots were produced.

- File #1 – Only 42 records, flow not started – do not process.

- File #2 – First few records low flow – will need editing. One other quick flow drop-out and a section of noisy salinity. May need editing. Flow rate changes early from about 1.1 to 1 and stays at ~1 for rest of file.

- File #3 - Only 4 records – do not process.

- File #4 to 7 – Flow steady, ~1.

The track plot looks fine. The plot was added to the end of this report.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db and exported to a spreadsheet which was saved as 2015-20-ctd-tsg-comp.xls. Cast #8 had an upcast file but the data near the surface are not considered of high enough quality to include in this comparison.

There were 102 casts which overlapped with TSG files.

The 5 TSG files to be processed were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, lab temperature, salinity and fluorescence and the files were reduced to the times of CTD files and loop samples.

The TSG data for the times of CTD stops were transferred to the comparison file. To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude and longitude were all≤0.0008° and only 3 differences were >0.0004°; the medians in both were 0.0000°. The larger differences came from times when the record shows the ship was moving more than usual during a stop. This shows both the times and positions are reliable for both systems.

The loop samples were combined in file 2015-20-tsg-loop-rosette-comp.xlsx. The rosette CHL bottle data were added for loops taken during CTD casts – there were no rosette salinity samples from the surface. TSG salinity and fluorescence (median over 2 minutes) from the time of sampling were added to the file.

These two spreadsheets will be used in step (d).

d.) Comparison of T, S and Fl from Loop & Rosette Samples and TSG and CTD data

- T1 vs T2 The intake temperature sensor worked throughout the cruise. The differences between the intake and lab temperatures were very noisy up to cast #22 and after cast #163, but fairly steady elsewhere. The median difference between intake temperature and lab temperature while the ship was stopped was 0.237C° and using all casts the median was 0.232 C°. We expect that the heating in the loop will increase as the intake temperature decreases, getting farther from the ambient loop temperature. For most of the cruise the intake temperature varied little, at ~10-13°C. However, early in file #7 temperature rose dramatically to as high as 19.5°C as the ship came into the Strait of Georgia. At that point the temperature differences became very noisy with many positive values as the intake temperature was likely higher than the ambient loop temperature. The Rivers Inlet part of file #6 has a similar though less extreme change. This confirms that if intake temperature is not available, finding a proxy for the intake temperature will require a temperature-dependent correction to lab temperature.

- TSG vs CTD The spreadsheet comparing CTD and TSG files was examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and TSG. An initial examination showed much higher variability than usual in the differences. A check was made of the near-surface data and temperature gradients were often high at about 4m. This was striking for the first few casts; the first cast with fairly well-mixed surface waters was #17 and that was the first to have temperature differences between TSG and CTD looking about as expected. Normally making a plot of differences versus standard deviation leads to a trend that is quite clear. That is not the case for the temperature data. Some cases with large standard deviations show small differences; plots without some of the largest standard deviations are more useful.

Another approach was taken by repeatedly sorting the data based on the standard deviations in the variable being studied. For the intake temperature, the median difference was smallest when 27 casts were used. For the lab temperature there was little difference based on how many points were included; there seemed to be a better relationship to the intake temperature standard deviation, though again the difference was not great. For salinity the median difference increased as more casts were included. So the 27 cases with the lowest standard deviation in the intake temperature will be used to study intake and lab temperature and fluorescence. The salinity and fluorescence will be discussed below.

1. Intake Temperature The intake temperature is higher than the CTD temperature by a median of 0.008C° (standard deviation 0.027C°). When differences are plotted against standard deviation in the TSG temperature the differences approach ~0.009C°.

2. LAB TEMP The lab temperature was higher than the CTD temperature by a median value of 0.23C°.

3. SALINITY The TSG salinity data are lower than the CTD salinity by a median value of 0.07 (standard deviation 0.027), but looking at the plot of differences versus standard deviation the difference estimated is 0.02. As usual almost all the outliers are in the direction towards TSG salinity reading lower, prosably due to small bubbles in the loop system.

4. FLUORESCENCE The TSG fluorescence was lower than the CTD fluorescence for all but one cast. The ratio of TSG/CTD is about 0.5 when all data are included, but plots of the ratio versus either CTD or TSG fluorescence shows that the ratio is about 0.5 for low fluorescence but drops steeply as fluorescence rises. This is a pattern often seen when fluorometers are compared with extracted chlorophyll but this is extreme. The CTD fluorometer reads high at low values but compared well with the extracted chlorophyll at high values.

It looks as though either the TSG fluorometer was performing poorly or there is a problem in the loop.

(See 2015-20-ctd-tsg-comp.xls.)

- Loop Bottle - TSG Comparisons

Loop extracted CHL samples were taken in 2 ways – filtered and unfiltered. The ratio of unfiltered values to filtered values ranged from 0.6 to 1.2 with average and median ratios being about 0.9.

The TSG fluorescence is much lower than the loop extracted chlorophyll except for one sample. When 3 cases with high standard deviation in the fluorescence data are excluded, the ratio of fluorescence to loop CHL has a median value of 0.35, with lower values for higher CHL.

The median difference shows the TSG salinity being low by 0.168 but the standard deviation of the differences is 0.11. When cases with high standard deviation in the TSG data are excluded the TSG salinity is low by a median of 0.065 but there are only 7 values in that comparison. There were some concerns about the quality of the bottle salinity data for this cruise.

(See 2015-20-TSG-loop-rosette-comp.xlsx.)

- Surface rosette samples versus TSG

There are 6 cases where we have extracted chlorophyll samples from a surface rosette and only 1 for salinity. The ratio of the TSG fluorescence to Rosette CHL falls from ~0.8 for the lowest CHL to ~0.15 for CHL=6ug/L. The TSG salinity was lower than the rosette sample by 0.10

- Surface rosette samples vs Loop samples

There are only 5 cases where we have CHL samples from both the loop and a rosette and there are none for salinity. The unfiltered samples are reasonably close to the rosette samples (ratio from 0.81 to 1.53). For the 3 filtered samples available the ratio ranged from 0.76 to 1.75, but the largest of those had a loop sample flagged 3 due to high variability.

- Calibration History

The TSG temperature and conductivity were recalibrated in December 2013 and have been used for 5 previous cruises. The Data are available from cruise 2015-09 that followed it as well.

During 2014-21 the TSG salinity was found to be lower than loop samples and CTD salinity by ~0.037 but the difference varied with flow rate which was highly variable.  No recalibration was applied due to the variability in the comparisons and the fact that such a large drift in calibration on its first use seemed unlikely. During 2014-18 the salinity was found to be low by 0.014 but the TSG was so noisy that this was not trusted. During 2014-19, 2014-22 and 2015-01 the TSG salinity was found to be low by ~0.02. During 2015-09 it was found to be low by ~0.03.

During 2014-19 the TSG temperature was found to be higher than the CTD temperature by ~0.005Cº and for 2014-21 it was higher by from ~0.005 to ~0.007Cº. During 2014-09 which followed the TSG temperature was found to be higher than the CTD temperature by ~0.004 to 0.006Cº.

The fluorometer in use for this cruise was used on 2015-01 and 2015-09; for the former it read higher by about 30% than the CTD fluorometer but all CHL values were low. For the later cruise the sensor performed badly with values increasing until they went off-scale in waters where high fluorescence is most unlikely.

Conclusions

1. The TSG clock worked well.

2. The TSG flow rate was very steady at a rate of ~1 except for an initial section of the first file when it was ~1.1.

3. The temperature in the loop increases by about 0.23Cº on average which looks reasonable for the time of year. In the Strait of Georgia and Rivers Inlet where surface temperatures were high, the temperature in the loop often increased or showed little difference.

4. The TSG intake temperature was higher than the CTD temperature by ~0.008Cº which is slightly higher than for 2014-19 or 2015-09. While this could be a real difference, it is more likely due to a small mismatch in depth for TSG and CTD readings combined with slightly higher near-surface temperature gradients than for the other cruises. No recalibration is justified.

5. The TSG Salinity is lower than the CTD salinity by a median of ~0.07 but plots of differences versus the standard deviation in the TSG suggests that a difference of 0.02 is more likely in quiet waters. Where the variability is high it is possible that there are also more bubbles which tend to lower the TSG salinity. The TSG salinity is lower than loop samples by a median of 0.065, when some outliers are excluded. There are some doubts about the bottle results that might have led to slightly high values though we would expect that error to be < 0.01; it is possible that the higher differences are partly due to bubbles. The cruises that preceded and followed this one had TSG salinity closer to CTD salinity (0.02 and 0.03), quite likely because surface gradients were lower, on average. If the temperature is high because water is coming from higher in the water column, we would also expect salinity to be lower for the same reason. In general, the differences are higher close to shore. Recalibration using +0.025 looks reasonable; this correction may not reflect calibration drift since it is quite likely that small bubbles are influencing the results. It seems reasonable to correct for whatever we see in quiet, well-mixed waters.

6. Part of the salinity error could be due to the TSG temperature being a little too high. We have too little evidence to judge this. It is more likely due to a slight mismatch of sampling depths.

7. The TSG fluorescence is too low. It is lower than the CTD fluorescence, loop chlorophyll and rosette samples. It is particularly low at the high end of the CHL range. This cannot be explained by assuming the water came from a little higher; a few casts were checked and there was CTD fluorescence of ~10 to 12ug/L right to the surface, yet the TSG fluorometer read about 3ug/L. Even the shape of the fluorescence trace is not useful because of the extremely low values it gives at CHL peaks. This sensor appeared to be ok during 2015-01. It failed during the cruise that followed but in that case values were too high and kept climbing higher. What malfunction could lead to low values in one case and high in another is unknown, but these data do not appear to be worth archiving.

8. There are too few points of comparison between the loop and the rosette to comment, except to note that the unfiltered loop CHL samples are closer to the rosette CHL than the filtered ones.

9. The overall quality of data was high with little of the spiking that has plagued this equipment in the past.

f.) Editing

The ATC files were copied to *.EDT.

Each file was opened in CTDEDIT and the following editing was applied:

- File #2 – CTDEDIT was used to remove Temperature:Intake, Temperature:Lab, Fluorescence:URU;Seapoint and Salinity:T0:C0 points from records 1-17 and records 2370-2374 because the flow rate was very low.

- File #4 – Cleaned a few single-point salinity spikes.

- File #5 – No editing applied.

- File #6 – No editing applied.

- File #7 – No editing applied.

g.) Recalibration

File 2015-20-tsg-recal.ccf was prepared to add 0.025 to the salinity channel and was applied to all *.EDT files.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Difference, Fluorescence:URU:Seapoint, Conductivity:Primary, Flag and Position:New channels.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH, fix the geographic description and add the depth of sampling to the header and to change channel names to standard names and formats.

Those files were saved as TOB files.

The TSG sensor history was updated.

A header check was run and no problems were found.

As a final check plots were made of the cruise track and it looks fine. x

# Particulars

1. Test firings, no Niskin sampling.

2. Niskin 1 sampled, all others closed just to test.

8. Acquisition did not start until 50m of downcast.

12. Note in log suggesting pumps might have been off, but they were actually on.

14. Archiving started at 24metres.

15. Loop taken with 5m rosette firing for event 14.

40. Altimeter did not kick in.

42. Bottle at 250m for NH4 blank – no sample #, no CHE file needed.

46. Deck unit changed to s/n 0508

48. PAR on.

74. PAR off.

87-115. Fluorometer cable changed due to spikes– gain not changed, but cable put in backwards – remove FL from casts 87-115.

89. PAR on.

102. Niskin 1 closed but no CHE file needed.

105. PAR off.

114. Bottle fired but no sample #, no CHE file needed.

117. Fluorometer cable changed –gain unchanged.

123. Log note says pump on after archiving, but appears to have been turned off while still recording and in water. Perhaps it was turned on again after acquisition stopped.

132. Bottom rose during cast from 1515 to 1380m.

134. After 500db Niskin was closed CTD taken down to 526db to fix winch spooling issue.

147. Stop at 894m on downcast to check wire. Stopped at 1359 on upcast and unspooled a bit.

155. PAR on.

163. Started archiving before down to 10 and back. Records from first drop removed in processing.

199. Two bottles fired at surface – only 1 sample # - assumed Niskin #1 sampled.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 31Jan2013 | Factory |  |  |
| Conductivity | 1763 | 1Jan2014 | Factory |  |  |
| Secondary Temp. | 5013 | 27Feb2013 | Factory |  |  |
| Secondary Cond. | 3394 | 3Jan2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 1438 | 3Jan2014 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| SeaPoint Fluor. | 3640 | n/a |  |  |  |
| Pressure Sensor | 506 | 30Dec2013 | Factory |  |  |
| Altimeter | 62354 | n/a | Factory |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3363       Cruise ID#:	2015-20

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3363 | 28Dec13 | Factory |  |  |
| Conductivity | 3363 | 28Dec13 | Factory |  |  |
| SeaPoint Fluorometer | 2356 |  |  |  |  |
| Temperature:Secondary | ? | ? |  |  |  |
| Flow meter | ? | n/a |  |  |  |