## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2015-39

Agency: PBS, Salmon And Freshwater Ecosystems, Nanaimo, B.C.

Party Chief: Zubkowski T.			Location: Strait of Georgia

Project: High Seas Salmon			Platform: Neocaligus

Date: 25 May 2015 – 28 May 2015

Processed by: Germaine Gatien

Date of Processing: 9 May2016 – 10 May 2016

Number of original HEX files: 6		Number of CTD files: 6

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (s/n 0334) was used with temperature sensor #2968, conductivity sensor #2173, WetLabs ECO Fluorometer #2215, pH sensor #0852, oxygen sensor #766 and pressure sensor 0464.

# SUMMARY OF QUALITY AND CONCERNS

No log book was provided but there was a scan of the bridge Document which provided the necessary information. There were only 6 shallow casts.

The CTD was soaked above 1db, generally for about 1.5 minutes. A soak of 3 to 5 minutes is recommended by the manufacturer but that is to allow the instrument’s temperature to equilibrate. If the temperature difference between air and water is not large, a 1 minute soak may be sufficient.

There was no calibration sampling during this cruise or during other cruises that used the same sensors since they were last calibrated at the factory.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data were available. Calibration is required for each cast to get absolute values although general trends within a cast are likely real.

There was no calibration sampling.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea. The file names were non-standard.

2. Preliminary Steps

The Daily Log was obtained.

The cruise summary sheet was completed.

A single configuration was file used.

The calibration parameters were checked and no problems were found.

The configuration was saved as file 2015-39-ctd.xmlcon.

There has been no known use of the pressure, temperature, conductivity and dissolved oxygen sensors between the last factory calibration and this cruise. They were used during 3 cruises after this one that have been processed, but none included calibration sampling and no recalibrations were applied to data from those cruises.

## 3. Conversion of Raw Data

All casts were converted using file 2015-39-ctd.xmlcon.

Plots show that the channels all have reasonable profiles. The soak time was about 100s.

The file names were missing one digit in the event # part. A zero was added to each name.

The times in the headers agree with those in the bridge log and appear to be in PDT.

4. WILDEDIT

This step was skipped as there were no obvious spikes to remove from pressure, temperature or conductivity except at the surface. WILDEDIT would not work well on those spikes.

5. FILTER

There were only 6 shallow casts, so the choice of filter was based on the tests run for cruise 2015-26 during which the same sensors were used.

WFILTER was run with a cosine filter width of 5 for Pressure, Temperature and Conductivity.

6. ALIGNCTD

Tests were run using advancements of 2 to 3s for the dissolved oxygen voltage channel to align it with temperature. ALIGNCTD was run on all files using +2.5 seconds.

7. CELLTM

CELLTM was run using parameters (α, 1/β) = (0.04, 8) as recommended by SeaBird for the SBE 25.

8. DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

CLEAN was run to replace pad values in the pressure channel with interpolated values.

All times in the file headers are in PDT, so ADD TIME CHANNEL was used to add 7 hours to header times of the SBE25 files.

10. Checking Headers

Some header information was missing: latitude, longitude, station name and water depth. A spreadsheet was provided with that information enabling preparation of file 2015-39-merge.csv. IOS SHELL program “MERGE:CSV file to headers” was used to add information from the csv file to the IOS headers.

A cross-reference listing was produced and checked against the log entries and no errors were found.

A header check was run and shows that there are many values that must come from spikes; these will likely disappear when the surface is clipped and DELETE is run. One error was found and corrected and no further errors were found on a rerun.

A track plot looks fine and was added to the end of the report.

The surface check shows an average value of -0.015db. This check is not very useful for this type of sampling as there may be records that are in-water but with the pump not on. However, plots show conductivity rising when pressure is very close to 0db.

##### 11. SHIFT

pH

Normally pH needs alignment, but there is so little variation from top to bottom that there is no noticeable difference when various alignment shifts are applied. Any differences are within the noise level. This step was skipped.

Conductivity

Tests were run to see what conductivity shift produced the best results as judged by looking at T-S plots.

A shift of +0.7 records improved the data most.

Fluorescence

The fluorometer was not pumped, so a shift in alignment is expected to be small or unnecessary. Profile plots of temperature and fluorescence were examined and show that alignment is quite good. No further alignment adjustment will be made.

Oxygen

The dissolved oxygen channel was aligned earlier in processing, but was checked again. As usual, the results vary from feature to feature as local vertical gradients and descent rates affect the alignment. No further shift will be applied.

##### 12. DELETE, CLIP and CTDEDIT

DELETE was run on all casts using the following parameters:

Remove Surface Records in the top 10m.

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

Plots were made to see what editing was needed. It was clear that most data above 1.2db were poor, being from the soak period or the time when the descent was just starting. There were large spikes and obviously unstable features.

CLIP was run to remove data with pressure <1db.

All files were examined in CTDEDIT and some editing was applied to 4 casts. This involved removing a few records from the top and bottom of the casts and very light cleaning of salinity and for cast #30 the removal of some negative fluorescence and dissolved oxygen values in the top 7m.

All CLIP files were copied to EDT and then the CTDEDIT output files were copied to EDT to produce a complete set of files.

14. Inter-comparisons

Sensor History – There is no useful history for these sensors since they were last calibrated at the factory. They were used later on a few other cruises with no calibration sampling.

Comparison of repeat casts –There were no repeat casts.

Historic Ranges – Plots of temperature and salinity with the local climatology superimposed show salinity within the normal range, but a lot of the temperature data are above the maximum including the Cowichan Bay cast and the 2 more southerly Sturgeon Bank casts. While this might just mean the climatology is too severe for near-shore data, it should be noted that temperatures above the climatology maximum have been noted over much of the top 100m in the Strait of Georgia in 2015.

Post-cruise calibrations – None available.

15. CALIBRATE

There was no calibration sampling and no useful history to enable estimates.

16. Fluorescence Filter

This step will be skipped as the fluorescence is not very spiky.

17. Bin Average and REMOVE

The files were bin averaged using 1db bins.

T-S plots were examined look fine.

REMOVE was run to remove Scan_Number, Conductivity:Primary, Oxygen:Voltage, Descent Rate and Flag channels from all casts.

Dissolved Oxygen was derived in mass units and surface saturation was then calculated. Values ranged from 87% to 124%. Given the location and season the range is reasonable, but nothing more than that can be concluded.

18. HEADER EDIT and final checks.

Header Edit was used to fix headers, fix formats and channel names and to add the following note to the headers:

Data Processing Notes:

----------------------

Fluorescence data are nominal and unedited, except that some records

were removed in editing temperature and salinity.

There was no salinity or dissolved oxygen calibration sampling and no other

information  available on which to base a recalibration.

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing. Some data

were removed in editing temperature and salinity and where the pH

sensor had clearly malfunctioned, but the data are otherwise unedited.

Calibration is required for each cast to get absolute pH values,

although general trends within a cast are likely real.

For details on the processing see processing report: 2015-39-proc.doc.

A cross-reference listing was produced.

A header check was run on the CTD files and no further errors were found.

The sensor history was updated.

## CRUISE SUMMARY

| Cruise ID#: 2015-39 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Dates: Start: 25 May 2015 End: 28 May 2015 |  |  |  |  |
| Location: Strait of Georgia |  |  |  |  |
| Party Chief: Zubkowski T. |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? |
| 1 | SEABIRD | 25 | 0334 | No |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0334	Cruise ID#:		2015-39

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 23Dec2014 | Factory |  |  |
| Conductivity | 2173 | 19Dec2014 | Factory |  |  |
| pH | 0852 | 31Dec14 | IOS |  |  |
| ECO Fluorometer | 2215 | May 2015 | ? |  |  |
| SBE43 Oxygen | 0766 | 23Dec2014 | Factory |  |  |
| Pressure | 464 | 23Dec2014 | Factory |  |  |