## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2015-41

Agency: PBS, Salmon And Freshwater Ecosystems, Nanaimo, B.C.

Party Chief: Zubkowski T.			Location: Strait of Georgia

Project: High Seas Salmon			Platform: Neocaligus

Date: 16 July 2015 – 19 July 2015

Processed by: Germaine Gatien

Date of Processing: 10 May2016 – 12 May 2016

Number of original HEX files:	7	Number of CTD files: 7

# INSTRUMENT SUMMARY

SeaBird Model SBE-19+ CTD (s/n 5299) was mounted with oxygen sensor #47.

# SUMMARY OF QUALITY AND CONCERNS

There was no logbook available but there was a scan of the Bridge log.

There were errors in the calibration file used at sea for the SBE19. The most recent values had not been used. The computer clock had a date that was wrong by more than 4 months.

The CTD was soaked above 1db, generally for 1.25 minutes. A soak of 3 to 5 minutes is recommended by the manufacturer but that is to allow the instrument’s temperature to equilibrate. If the temperature difference between air and water is not large, a 1 minute soak may be sufficient.

There was no calibration sampling.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea. The file names were in standard format.

##### Preliminary Steps

There was no log book available but there was a scan of a bridge document with the necessary information. There were only 5 shallow casts.

The cruise summary sheet was completed.

A single configuration was file used for the 5 casts.

There were errors in the SBE19+ calibrations. There was a more recent calibration for the temperature and conductivity and the pressure values had the right date, but the parameters were incorrect. The correct values were entered and the configuration file was saved as file 2015-40-ctd.xmlcon.

The configuration file was saved as file 2015-41-ctd.xmlcon.

These sensors were recalibrated in December 2014 or February 2015 and were used for 4 short cruises in May, June, July and August 2015, with no calibration sampling. As of May 2016 no other data have been processed that used these sensors.

##### Conversion of Raw Data

All casts were converted using file 2015-41-ctd.xmlcon.

Plots show that the channels all have reasonable profiles.

The file names had standard format.

The times in the headers for the SBE19 are in PDT but the dates are off by more than 4 months.

The CTD was soaked for about 1.25minutes. This is shorter than recommended by SeaBird but is probably adequate for the conditions.

##### WILDEDIT

This step was skipped as it does not usually work well on SBE19 data. The only large spikes noted are right at the surface, so will be removed by other routines.

##### FILTER

SeaBird recommend applying a low-pass filter with a 0.5s time constant to the conductivity to force it to have the same response as the temperature. This setting does improve the salinity as judged by T-S plots. There may be more adjustment needed, but this will done later using SHIFT.

Pressure does not appear to need a filter as it is quite smooth.

Filter was run using 0.5s setting for conductivity.

##### ALIGNCTD

There was run to advance the DO signal by +2.5 seconds which did brought the DO traces into better alignment with the temperature traces.

Temperature was advanced by 0.5s relative to pressure.

##### DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

CLEAN was run to replace pad values in the pressure channel with interpolated values.

Note that all times in the file headers are in PDT. So ADD TIME CHANNEL was used to add 7 hours to header times of the SBE25 files. The dates are wrong in the SBE19+ files so 3055 hours were added (127 days + 7 hours).

##### Checking Headers

Some header information was missing: latitude, longitude, station name and water depth.

A spreadsheet was provided and that was used to prepare file 2015-41-merge.csv. IOS SHELL program “MERGE:CSV file to headers” was used to add information from the csv file to the IOS files with output MRH.

A cross-reference listing was produced and checked against the log entries and 1 error was found and corrected.

A header check was also run and shows some negative values that likely come from surface spikes; these will likely disappear when DELETE is run and the surface is clipped.

A track plot was produced and added to the end of this report.

The surface check is not very useful for this type of sampling as there may be records that are in-water but with the pump not on. The average surface pressure was -0.09db. The downcast plots suggest the pressure is slightly low and the upcast plots suggest it is slightly high. Both are within the resolution of the instrument.

##### SHIFT

Conductivity

The alignment to temperature applied earlier improved the stability on a T-S surface, but tests were run to see if some improvement can be achieved by shifting the conductivity slightly. All settings tried reduced stability, so no further shift was applied.

Oxygen

The dissolved oxygen channel was aligned earlier in processing, but was checked again. As usual, the results vary from feature to feature as local vertical gradients and descent rates affect the alignment. No further shift will be applied.

##### DELETE, CLIP and CTDEDIT

DELETE was run on all casts using the following parameters:

Remove Surface Records in the top 10m.

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

Plots were made to see what editing was needed. It was clear that most data above 1.2db were poor, being from the soak period or the time when the descent was just starting. There were large spikes and obviously unstable features.

CLIP was run to remove data with pressure <1.2db.

All files were examined in CTDEDIT and some editing was applied to all 7 casts. A few records were removed due to corruption by shed wakes. Salinity was cleaned to remove the small excursions that appear related to variations in alignment caused by descent rate variations.

##### Inter-comparisons

Sensor History – There is no useful history for these sensors since they were last calibrated at the factory. They were not used on any other cruises with calibration sampling that have been processed as of May 2016.

Comparison of repeat casts –There were no repeat casts.

Historic Ranges – Plots of temperature and salinity with the local climatology superimposed show that all data fell within the normal range except for the 2 casts in Cowichan Bay. There, temperature and salinity were both above the climatology maximum down to about 40db. The 3-standard deviation climatology is not appropriate for near-shore casts. High near-surface temperatures have been observed during other 2015 Strait of Georgia cruises.

Post-cruise calibrations – None available.

##### CALIBRATE

There was no calibration sampling and no useful history to enable estimates.

##### Bin Average and REMOVE

The files were bin averaged using 1db bins.

T-S plots were examined and look fine.

REMOVE was run to remove Scan_Number, Conductivity:Primary, Oxygen:Voltage, Descent Rate and Flag channels from all casts.

As a final check dissolved oxygen surface saturation was calculated. Values ranged from 72% to 105% with the higher values coming from the last two stations. These values are reasonable for this region, but they do not shed much light on the quality of the DO calibration.

##### HEADER EDIT and final checks.

Header Edit was used to fix headers, fix formats and channel names and to add the following note to the headers:

Data Processing Notes:

----------------------

There was no salinity calibration sampling and no other information

available on which to base a recalibration.

For details on the processing see processing report: 2015-41-proc.doc.

A cross-reference listing was produced.

The standards check was run and no problems were found.

A header check was run on the CTD files and no further errors were found.

The sensor history was updated.

## CRUISE SUMMARY

| Cruise ID#: 2015-41 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Dates: Start: 16 July 2015 End: 19 July 2015 |  |  |  |  |
| Location: Strait of Georgia |  |  |  |  |
| Party Chief: Zubkowski T. |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? |
| 1 | SEABIRD | 19+ | 5299 | No |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/5299	Cruise ID#:		2015-41

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5299 | 13Feb2014 | Factory |  |  |
| Conductivity | 5299 | 13Feb2014 | Factory |  |  |
| SBE43 Oxygen | 47 | 20Dec2014 | Factory |  |  |
| Pressure | 5299 | 19Dec2014 | Factory |  |  |