## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2015-42

Agency: PBS, Salmon And Freshwater Ecosystems, Nanaimo, B.C.

Party Chief: Zubkowski T.			Location: Strait of Georgia

Project: High Seas Salmon			Platform: Neocaligus

Date: 21 August 2015 – 24 August 2015

Processed by: Germaine Gatien

Date of Processing: 11 May2016 – ?? May 2016

Number of original HEX files:	6	Number of CTD files:

# INSTRUMENT SUMMARY

SeaBird Model SBE-19+ CTD (s/n 5299) was mounted with oxygen sensor #47.

# SUMMARY OF QUALITY AND CONCERNS

There was no logbook available but there was a scan of the Bridge log.

There were errors in the calibration file used at sea for the SBE19. The most recent values had not been used. The computer clock had a date that was wrong by more than 4 months.

The CTD was soaked above 1db for more than 2 minutes. A soak of 3 to 5 minutes is recommended by the manufacturer but that is to allow the instrument’s temperature to equilibrate. If the temperature difference between air and water is not large, a 2 minute soak should be fine.

There was no calibration sampling.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea. The file names were non-standard.

##### Preliminary Steps

There was no log book available but there was a scan of a bridge document with the necessary information. There were 6 casts, all shallow.

The cruise summary sheet was completed.

A single configuration was file used for the 6 casts.

There were errors in the SBE19+ calibrations. There was a more recent calibration for the temperature and conductivity and the pressure values had the right date, but the parameters were incorrect. The correct values were entered and the configuration file was saved as file 2015-42-ctd.xmlcon.

The configuration file was saved as file 2015-42-ctd.xmlcon.

These sensors were recalibrated in December 2014 or February 2015 and were used for 4 short cruises in May, June, July and August 2015, with no calibration sampling. As of May 2016 no other data have been processed that used these sensors.

##### Conversion of Raw Data

All casts were converted using file 2015-42-ctd.xmlcon.

Plots show that the channels all have reasonable profiles.

The file names had standard format.

The times in the headers for the SBE19 are in PDT but the dates are off by more than 4 months.

##### WILDEDIT

This step was skipped as the only spikes are at the surface and will be removed later.

##### FILTER

A low-pass filter with a 0.5s time constant was applied to the conductivity to force it to have the same response as the temperature.

The pressure is smooth and does not require a filter.

##### ALIGNCTD

There was run to advance the DO signal by +2.5 seconds which did brought the DO traces into better alignment with the temperature traces.

Temperature was advanced by 0.5s relative to pressure.

##### DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration.

Plots show the filters and alignment worked well.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers. This was done separately for each CTD to get the correct information in the headers.

CLEAN was run to replace pad values in the pressure channel with interpolated values.

Note that all times in the file headers are in PDT. So ADD TIME CHANNEL was used to add 7 hours to header times of the SBE25 files. The dates are wrong in the SBE19+ files so 3055 hours were added (127 days + 7 hours).

##### Checking Headers

Some header information was missing: latitude, longitude, station name and water depth.

A spreadsheet was provided and that was used to prepare file 2015-42-merge.csv. IOS SHELL program “MERGE:CSV file to headers” was used to add information from the csv file to the IOS files with output MRH.

A cross-reference listing was produced and checked against the log entries and no errors were found.

A preliminary header check was also run and no errors were found.

A track plot was produced and added to the end of the report.

The surface check is not very useful for this type of sampling as there may be records that are in-water but with the pump not on. The average surface pressure was -0.1db. The pressure looks close to zero based on both downcasts and upcasts.

##### SHIFT

Conductivity

No shift was applied to conductivity since all values tests made the data worse.

Oxygen

The dissolved oxygen channel was aligned earlier in processing, but was checked again. As usual, the results vary from feature to feature as local vertical gradients and descent rates affect the alignment. No further shift will be applied.

##### DELETE, CLIP and CTDEDIT

DELETE was run on all casts using the following parameters:

Remove Surface Records in the top 10m.

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: The only warnings were from the upcast of one file.

Plots were made to see what editing was needed. It was clear that most data above 1.2db were poor, being from the soak period or the time when the descent was just starting. There were large spikes and obviously unstable features.

CLIP was run to remove data with pressure <1.2db.

All files were examined in CTDEDIT and some editing was applied to 5 casts. This involved removing records corrupted by shed wakes and light cleaning of salinity. Some fluorescence spikes near the top of the water column were removed from the SBE404 files.

All CLIP files were copied to EDT and then the CTDEDIT output files were copied to EDT to produce a complete set of files.

##### Inter-comparisons

Sensor History – There is no useful history for these sensors since they were last calibrated at the factory. They were not used on any other cruises with calibration sampling that have been processed as of May 2016.

Comparison of repeat casts –There were 2 casts close together but the first was extremely shallow, so no useful comparison could be made.

Historic Ranges – Plots of temperature and salinity with the local climatology superimposed show all salinity within the normal range. The temperatures were all high in Cowichan Bay and Satellite Channel and through most of the profile on Roberts Bank. On Sturgeon Bank and North Fraser Arm temperatures were within the climatology.  While the 3-standard-deviation climatology is too severe for near-shore data, it should be noted that temperatures above the climatology maximum have been noted over much of the top 100m of the Strait of Georgia waters in 2015.

Post-cruise calibrations – None available.

##### CALIBRATE

There was no calibration sampling and no useful history to enable estimates.

##### Bin Average and REMOVE

The files were bin averaged using 1db bins.

T-S plots were examined and no problems were found.

REMOVE was run to remove Scan_Number, Conductivity:Primary, Oxygen:Voltage, Descent Rate and Flag channels from all casts.

##### HEADER EDIT and final checks.

Header Edit was used to fix headers, fix formats and channel names and to add the following note to the headers:

Data Processing Notes:

----------------------

There was no salinity calibration sampling and no other information

available on which to base a recalibration.

For details on the processing see processing report: 2015-42-proc.doc.

A cross-reference listing was produced.

A header check was run on the CTD files and no further errors were found.

The sensor history was updated.

## CRUISE SUMMARY

| Cruise ID#: 2015-42 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Dates: Start: 21 August 2015 End: 24 August 2015 |  |  |  |  |
| Location: Strait of Georgia |  |  |  |  |
| Party Chief: Zubkowski T. |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? |
| 1 | SEABIRD | 19+ | 5299 | No |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/5299	Cruise ID#:		2015-42

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5299 | 13Feb2015 | Factory |  |  |
| Conductivity | 5299 | 13Feb2015 | Factory |  |  |
| SBE43 Oxygen | 47 | 20Dec2014 | Factory |  |  |
| Pressure | 5299 | 19Dec2014 | Factory |  |  |