## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27 Mar 2025 | Updated channel names & formats in TOB files. G.G. |

## PROCESSING NOTES

Cruise: 2015-48

Agency: MEAD

Location: North-East Pacific

Project: Seamount Ecology

Party Chief: Curtis J.

Platform: John P. Tully

Date: July 4, 2015 – July 21, 2015

Processed by: Germaine Gatien

Date of Processing: 17 February 2016 – 3 May 2016

Number of original HEX files: 1

Number of CTD files: 22	Number of bottle casts:	14

Number of Dive CTD files: 21 	Number of Dive files processed: 17

Number of TSG hex files: 1	Number of TOB files: 1

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0506) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#1438), a SeaPoint Fluorometer (#3640) and an altimeter (#62354).

SeaBird Model SBE25 Sealogger CTD (#0464) was mounted on a camera and attached were a SBE43 DO sensor (#0766) and a WetLabs ECO_AFL/FL fluorometer (#2215).

A thermosalinograph (Seacat 21 S/N 3363) was mounted with a WetLabs fluorometer (#WS3S-953P), a remote temperature sensor and a flow meter.

The data logging computer was #2.

The data acquisition program was Seasave.

The CTD deck unit was an SBE model 11+, serial number 0508.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log book was in good order.

The file names were non-standard. It is helpful to save the files with the event number in them.

There are two sets of CTD files – the SBE911+ files which were mounted on a rosette frame and the SBE25 files which were mounted to the drop-camera frame. The event numbers overlap, so to avoid confusion, the SBE911+ files were named in the usual way with event numbers matching those in the Daily Science Log, ex 2015-48-0001. A leading “9” was included in the SBE25 files so that drop 4 became event 2015-48-9004.

The dissolved oxygen analyst noted concerns about the quality of the titrated oxygen samples from late in the cruise. The comparison with CTD dissolved oxygen confirmed that there was a change so that bottles from events #30-38 look different from earlier in the cruise and different from other cruises using the same CTD DO sensor. So the Oxygen:Dissolved data are reported with only 2 decimal points rather than the usual 3 to emphasize the quality concerns. Values from those casts were also flagged.

The Oxygen:Dissolved:SBE data from the SBE911+ are considered, very roughly, to be:

±0.5 mL/L from 0 to 100db

±0.2 mL/L from 100db to 500db

±0.07 mL/L from 500db to 1500db

±0.05 mL/L below 1500db

There were no salinity calibration samples. A post-cruise calibration was available but there was one cruise between this one and the factory visit so some of the drift may have occurred after this cruise.

A thermosalinograph was in use. There were no loop samples. The external thermistor worked throughout the cruise. TSG Salinity was recalibrated based on a comparison with CTD data and the history of the sensor. The fluorescence data are raw with volts as units. Based on a comparison with CTD fluorescence, it appears that TSG fluorescence in ug/L would be roughly 5 times the raw values.

An SBE25 was mounted on a dive camera. The data appear to have been averaged over 1s on acquisition, though the configuration file does not indicate that was planned. It is assumed there must have been some setting on the deck unit that caused this to happen. The CTD data have been processed, but not all the usual steps applied. No data were removed.

# PROCESSING SUMMARY

NOTE: Details about the processing of the SBE25 that was mounted with the Dive Camera may be found in §25.

SBE911+ Processing

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

The file names are non-standard; the station names are included in the file names but the event numbers are not, so fixing the file names requires checking each entry against the log book.

The data from the initial soak period at 10m were included in the files. These data must be removed before running DELETE.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Nutrients, dissolved oxygen and extracted chlorophyll data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked.

The XMLCON files did not change through the cruise. The calibration constants were checked for all instruments and the only correction made was to the date format for the transmissometer.

The corrected file was saved as 2015-48-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All SBE911 hex files were converted using 2015-48-ctd.xmlcon to create CNV files.

The files were renamed in standard format with event numbers taken from the log book.

A few casts were examined.  The stops for bottles lasted at least 30s.

All expected channels are present.

The upcast and downcast temperature and conductivity channels track in the usual way, with the upcasts generally noisier and the channel pairs further apart than during downcasts. Dissolved oxygen, fluorescence, altimetry and transmissivity profiles all look normal. The descent rate was kept high but was often very noisy.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2015-48-ctd.xmlcon. The file names were corrected to standard format.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

A preliminary header check was done and no problems were found. Fluorescence did not go off-scale.

Temperature and salinity were plotted for all BOT files to check for significant outliers. Both salinity channels had some noisy patches. CTDEDIT was used to clean salinity from casts #5, 7, 17, 24 and 38.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine. Those files were bin averaged on bottle number.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2015-48-bot-hdr.txt which will be updated as needed during processing.

EXTRACTD CHLOROPHYLL

The extracted chlorophyll data were provided in spreadsheet QF2015-48chl.xlsx which includes flags and comments. There were no duplicates. The spreadsheet was simplified and saved as 2015-48chl.csv.

That file was converted to individual *.CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2015-48oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2015-48oxy.csv.

That file was converted into individual *.OXY files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2015-48nuts.xls.

Then the file was simplified, reordered on sample numbers and saved as 2015-48-nuts.csv. The file was converted to individual NUT files.

The CHL, OXY and NUT files were merged with CST files in 3 steps. After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

CLEAN was run on the MRG files to add 0 flags to empty flag channels and to update header limits.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. Nutrients were missing from cast #38 because it had been given the event #58 in the nutrient spreadsheet. The rosette log entry was hard to read. CHL was missing from events #40 and #42. The samples had been run with other cruises and got misplaced. They were found. The spreadsheet was corrected and the MRG files recreated.

##### Compare

Fluorescence

There were only 2 samples, both from the surface. There are not enough data to reach any conclusions about the performance of the fluorometer.

Event #40 – The CTD fluorescence during the bottle stop was 0.535ug/L and the extracted CHL was 0.21ug/L. The downcast CTD Fluorescence was 0.516ug/L.

Event #42 – The rosette log showed the pressure for this sample to be 90db but it was really at 2.8db. The CTD fluorescence during the bottle stop was 0.408ug/L and the extracted CHL was 0.52ug/L. The downcast CTD fluorescence was 0.39ug/L.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There is a lot of scatter in the comparison with the points seeming to fit roughly into 2 groups. The analyst reported a problem with replicate samples from about the halfway point in the data set. So there is likely a problem with all samples from the latter half of the cruise. No reason for the problem was found.

Because there were some deeper casts late in the cruise, the hysteresis setting could cause some scatter. A check was made to see if points below the OMZ stand out in a plot of (CTD DO – Bottle DO) vs CTD DO, but they do not.

Next, plots were made that include only data from above the OMZ. Those data show a large scatter throughout the DO range. The scatter is larger near the surface, as expected, but is still significant down to 600m. When plotted against file pair number it is clear that the scatter increases late in the cruise, which matches what the analyst noted. Those were the deepest casts, but the scatter is in the data above 800m. A plot of differences below 750m shows much less scatter, but that could be because the values are lower there and whatever is causing the problem is just less significant there.

So the problem does seem to lie with the extracted DO, meaning we cannot trust this comparison, at least late in the cruise.

When only events #3-24 are used excluding outliers based on residuals & CTD DO standard deviation >0.02, the fit is:

CTD DO Corrected = CTD DO * 1.0153 +0.0685   (R2 = 0.87)

For the cruise that preceded this one (2015-18) the fit used to recalibrate this sensor was

CTD DO Corrected = CTD DO * 1.0234 +0.0260 (R2 = 0.78)

For the cruise that followed this one (2015-45) the correction applied was:

CTD DO Corrected = CTD DO * 1.0231 +0.0295   (R2 = 0.61)

So the fit from the early casts are close in slope to the cruises that bracketed 2015-48, though the offset is larger.

The fit from events #30-38 with a similar approach to removing outliers was:

CTD DO Corrected = CTD DO * 1.0604 +0.0256 (R2 = 0.93)

The slope is high while the offset is closer to the other 2 cruises. 
The net effect is similar at the low end of the DO range, but very different at high DO. 
The analyst had already noted concerns about the data from these casts, and with this comparison as further evidence, he recommended archiving Oxygen:Dissolved with only 2 decimal points to draw attention to the lower data quality.

The fits always vary somewhat, perhaps depending upon how well bottles flushed and local vertical gradients. Because there are lots of bottles with very low DO, we should be able to trust the offset, and the lower slope may reflect lower DO gradients.

A fit was done for these data that excluded DO values <2ug/L to see if the result looked like the earlier ones. The fit is quite flat from about 2ug/L to 5.5ug/L. Above that there are many outliers that have bottle values much higher than the CTD. The fit looks nothing like what we usually see. Forcing an offset by choosing the one from 2015-46 gave a reasonable slope, but the fit was very poor. The 2015-48 fit is reasonable if we exclude casts #30-38.

For more detail see document 2015-48-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

For the cruises run before and after this one a setting of +3s was sued to align the DO signal with temperature. ALIGNCTD was run on all casts using +3s. Plots were made to check that it worked well and it did improve the alignment, but the noisy descent rates make it difficult to judge as the temperature response to shed wake corruption is so much quicker than for the DO.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. Tests on previous cruises using these sensors showed the default setting of (α = 0.0245, β=9.5) did the best job and it does improve the data for both conductivity channels for these data.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few deep casts to examine differences between sensor pairs.

Differences from some of the deeper casts are entered in the table below. For comparison, differences from earlier cruises and 1 later cruise using the same equipment are also shown with dark shading.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2015-20-0111 | 1000 | -0.0007 | 0 | +0.0010 | High, Noisy |
| 2015-09-0018 | 1000 | -0.0009 | -0.0001 | -0.0006 | High, V.Noisy |
| 2015-09-0104 | 1000 | -0.0007 | -0.0001 | +0.0006 | High, Moderate |
| 2015-18-0008 | 230 | -0.0017 | -0.0010 | +0.0008 | High, Noisy |
| 2015-18-0056 | 345 | -0.0024 | -0.0002 | +0.0002 | F.High, F.Steady |
| 2015-18-0085 | 350 | -0.0025 | -0.0003 | -0.0005 | High, F.Steady |
| 2015-48-0003 | 350 | -0.0026 | -0.0002 | -0.0008 | High, Noisy |
|  | 1000 | -0.0014 | -0.0004 | -0.0026 | High, Noisy |
| 2015-48-0011 | 350 | -0.0029 | -0.0001 | -0.0017 | High, X.Noisy |
|  | 1000 | -0.0013 | -0.0003 | -0.0019 | High, X.Noisy |
|  | 2500 | -0.0006 | -0.0004 | -0.0037 | High, X.Noisy |
| 2015-48-0038 | 350 | -0.0028 | -0.0003 | -0.0016 | High, X.Noisy |
|  | 1000 | -0.0011 | -0.0004 | -0.0039 | High, X.Noisy |
|  | 2500 | -0.0008 | -0.0004 | -0.0043 | High, X.Noisy |
| 2015-46-0002 | 300 | -0.0008 | -0.0005 | -0.0045 | High, V. Steady |
| 2015-46-0023 | 350 | -0.0014 | -0.0007 | -0.0068 | High, Steady |
| 2015-46-0026 | 500 | -0.0009 | -0.0007 | -0.0072 | High, V.Steady |
| 2015-46-0027 | 500 | -0.0009 N | -0.0016 | -0.0170 | F.High, V.Stead |
| 2015-46-0150 | 350 | -0.0014 | -0.0009 | -0.0088 | High, Noisy |
|  | 500 | ~ -0.004XN | -0.0010VN | -0.0087 |  |

The differences in temperature in shallow water are high for 2015-48 but that may be due to the very heavy corruption by shed wakes caused by the extremely noisy descent rate with many complete reversals of direction. At 1000m vertical gradients are lower so that shed wake corruption is not as significant; there the results are similar to 2015-20 and 2015-09. There is little variability in conductivity differences over time. Salinity differences suggest that there might have been some drift in calibration during 2015-46, but given few deep casts and very noisy traces, this is not clear. The differences near the end of 2015-48 look similar to the differences found at the post-cruise factory check. The differences during 2015-46 look erratic and some odd shifts in salinity were noted during that cruise, so there were likely sensor problems that were not due to calibration drift. There was some suspicion about the secondary channel during 2015-46, so the primary was selected for 2015-48.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. The times in the headers are generally about 5 to 10 minutes earlier than the log times and positions differ somewhat as the ship would have moved between those two times. The header time is captured when the file is started which is sometimes just before the ship stops so positions vary somewhat as well, though it may also due to shift drift while on station.

An initial header check was run. There are clearly some spikes in at least one cast, but they are likely in the initial soak stage, so will be removed in processing.

The cruise track was plotted and added to the end of this report. The initial test cast in Saanich Inlet was not included in the plots. A blow-up of the Douglas Channel stations was added since it is hard to see this area on the full map.

Surface check was run and shows an average surface pressure for the cruise was 2.7db with a range of 1.6-4.9db. Most casts haven no pressures recorded above 1db. For one that had data from ~0.4db of the upcast with pumps running, the salinity was >31psu, so the pressure looks about right.

Plots of altimetry near the bottom were examined to see if the header entries are reasonable. Despite many spikes, the algorithm worked well except for cast #38 where it appears that the CTD got within 0.5m of the bottom rather than 11.5m. The header entry was changed for that one and the comment edited to reflect that it was not determined by the usual algorithm. The entry for cast #26 was removed as the CTD clearly never got close to the bottom.

Header entries of bottom depth and altimeter values at the bottom of the cast were exported to spreadsheets for the full files and for the bottle files. The bottom depths were checked against the log book. There were many discrepancies. An estimate of bottom depth was made by using:

Bottom depth = CTD Max Pressure * 0.99 + Altimeter Reading

Those values were compared with the header and log entries to see if it is clear which is more accurate. In some cases neither looked right. The log entry was used to change the entries for casts #5 and 7, and estimates based on the equation above were used to change the entries for cast #9 and 33. These casts are in an area where depth may vary through a cast, so using the best estimate for the time the CTD is at the bottom is a compromise.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and as usual, there is a lot of variability with the up and downcast traces sometimes closer than temperature and sometimes further apart. This is likely due to varying vertical gradients. Where there are distinct features in T and DO profiles, they line up well. Overall the choice made earlier looks appropriate, so no further alignment will be applied.

Conductivity

Tests were run on a few casts to see if what shifts to conductivity do the best job of improving salinity as judged by removing unstable features from T-S plots. There was not much difference between various choices, but overall -0.9 records looks best for the primary and +0.9 records looked best for the secondary. SHIFT was run on all casts applying those shifts. Salinity was recalculated.

##### DELETE

Before running DELETE the data acquired during the soak period had to be removed by using CLIP. Plots were made to see how many records needed to be removed; this varied from cast to cast, so this step had to be done individually for each cast.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There was a warning for cast #26 – the CLIP routine had been applied incorrectly, so CLIP and DELETE were rerun on that cast. Cast #5 was also found to be wrong, though there was no warning since it went a little deeper than 10m. CLIP and DELETE were rerun for that cast.

##### Other Comparisons

Previous experience with these sensors –

Conductivity, pressure and dissolved oxygen sensors were all recalibrated in late 2013 or early 2014. They were used for 1 cast during 2014-19 and all of 2014-50, 2015-01, 2015-03, 2015-17, 2015-20, 2015-09 and 2015-18 and they were used after this cruise during 2015-46.

- Both T/C sensor pairs produced salinity within 0.001 of bottles for 2015-20 and 2015-09 while from 2015-18 it appeared that both salinity channels were low by ~0.003, but the comparisons had more scatter than usual. Repairs were made to the secondary pump after 2015-17. The best comparison was from 2015-09 with many deep casts; it showed the primary to be high by an average of 0.0005 while the secondary was low by 0.0013. During 2015-46 the primary salinity was found to be low by about 0.011 and the secondary by 0.017, but there were many problems with the comparison so recalibration was postponed until a post-cruise calibration was available. See below for those results.

- The pressure sensor was found to have drifted lower by about 1.25db during 2014-50 and that offset has been used since then.

- For 2014-50 the dissolved oxygen sensor was corrected using a linear fit of slope 1.0281 since there was too little sampling of waters with low DO values to estimate an offset. For 2015-01 the fit used had a slope of 1.0187 and offset of +0.056mL/L. For 2015-03 the slope was 1.0147 and the offset +0.0647mL/L. The lower slope may be due to incomplete flushing since the sensor drift leads to CTD DO values reading low, but if the samples are from lower in the water column they will generally be reading low too. The secondary pump problem may have affected the fit as well, but it does not appear to have been a large effect. For 2015-17 the comparison had a huge scatter, so the result of 2015-01 was used. For the offshore cruises 2015-20 and 2015-09 the slope/offset values applied were 1.0235/0.0248 and 1.0246/0.0452. For 2015-18 in Juan de Fuca and Strait of Georgia it was 1.0234/0.0260.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All data fell within the climatology (where climatology was available).

Repeat Casts – There were no repeat casts. Nearby casts were plotted together on a T-S surface and look reasonably close.

# Post-Cruise Calibration

The factory drift reports are rough and show the primary conductivity cell error leading to salinity being low by ~0.0041 which was partly offset by the primary temperature reading low by ~0.0008 deg C, for a net salinity error of something like -0.0033. The drift was also examined by reconverting a few 2015-46 files using the post-cruise calibration parameters. Those suggest the primary salinity is low by 0.0036 and primary temperature is low by 0.0013Cº.  That is reasonably close to the drifts found above given that the drift statement of 0.0002/month could be anywhere between 0.00015 or 0.0024 per month. For the secondary the drift was found to be smaller for both temperature and salinity with the temperature being low by ~0.0007 and salinity high by ~ 0.0013 at 500m.

##### DETAILED EDITING

The differences between the 2 salinity channels near the end of 2015-48 look similar to the differences found at the post-cruise factory check. The differences during 2015-46 look erratic and larger than the differences found post-cruise. Some odd shifts were found in both salinity channels during that cruise, so there were likely sensor problems that were not due to calibration drift. The secondary salinity looked most suspicious and upcasts were worse than downcasts, so the primary was selected for 2015-48. Flow rate variability was suspected as the source of trouble, but no cause was found for erratic flow.

For this cruise there is no evidence of the sort of shifts seen in the cruise that followed, but the descent rates are so noisy that they might not be obvious. Since the differences between salinity channels seen in these data are close to the post-cruise factory results, it looks like either channel pairs could be used.

The primary T and S channels were selected for editing as these were selected for the cruise that followed. CTDEDIT was used to remove large spikes, remove or clean smaller salinity spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and to remove records corrupted by shed wakes including some records from near the top and bottom of the casts. Some bad salinity points were also removed that are associated with decelerations of the CTD. A few casts (9, 11, 38, 40) needed heavy editing due to very noisy CTD descent rates. All files required some editing.

##### Initial Recalibration

Pressure will not be recalibrated.

The dissolved oxygen data will be recalibrated using the results of §5.

Salinity and temperature channels were recalibrated based on post-cruise calibrations.

CALIBRATE was run using file 2015-48-recal1.ccf to add 0.0036 to channel Salinity:T0:C0, to add 0.0018 to Temperature:Primary and to correct the Oxygen:Dissolved:SBE channel in the SAM and MRGCLN2 files using:

CTD DO Corrected = CTD DO * 1.0153 +0.0685

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly.

See file 2015-48-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are always expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data. This cruise is especially prone to the flushing and variability problems.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. The scatter in the comparison is high in the top 100m. The CTD DO is generally higher than bottles down to 800m which is what we expect if flushing is incomplete as the bottles contain water from deeper in the water column where DO is lower. Below 800m the CTD DO is close to or lower than bottles which is again expected below the Oxygen Minimum Zone. These differences are most likely due to incomplete flushing and no further calibration of CTD DO is justified.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. No problems were found.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity and Fluorescence data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5 mL/L from 0 to 100db

±0.2 mL/L from 100db to 500db

±0.07 mL/L from 500db to 1500db

±0.05 mL/L below 1500db

For details on the processing see the report: 2015-48-proc.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. All except 1 offshore cast had surface saturations of between 103% and 104%, as expected. There was 1 offshore cast with 107%, the nearshore cast at station LG02 was ~80% and in Saanich Inlet it was ~150%. None of these values suggest any problem with the DO calibration.

##### Final Bottle Files

CALIBRATE was run on the MRGCLN2 files and then the MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data from the CHE files were exported to a spreadsheet and compared with rosette sheets. No problems were found. The spreadsheet was saved as 2015-48-bottles_final.xlsx.

Plots of each file were examined. No problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

##### Thermosalinograph Data

There were no loop samples.

a.) Checking calibrations

The parameters in the configuration file are correct. The file was saved as 2015-48-tsg.xmlcon.

b.) Conversion of Files

There were 2 HEX files, but one was a test file.

There is some confusion about what fluorometer was mounted during this cruise. During the cruise that followed this one, 2014-46, an error was found in the configuration file; a SeaPoint fluorometer was entered where there should have been a WetLabs. So it was expected that this cruise would have the same error. However, both the log and the configuration files indicate that there was a WetLabs fluorometer. So it is hard to understand how the error crept in between the 2 cruises. Looking at the history of the instrument shows that there was a SeaPoint fluorometer on this TSG in June 2015, but the TSG was moved to the Vector where there was no fluorometer mounted. So it makes sense that a different fluorometer was mounted when the TSG was moved back to the Tully for this cruise. Just to be sure a check was made against a CTD cast for which fluorescence was ~1ug/L. If the SeaPoint configuration is used, the result is ~0.2ug/L. We do not have a calibration for the WetLabs so we only get a voltage of about 0.07, but in previous uses a comparison with CTD fluorescence suggests a scale factor of roughly 10. So 0.7ug/L would compare better with the CTD fluorescence than the 0.2ug/L found with the SeaPoint configuration. This is not conclusive evidence, but taken with the log and configuration file, it does appear that the TSG fluorometer was a WetLabs.

The HEX file was converted to a CNV file.

They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels. (*.ATC)

The track plot and time-series plot of data both looked.

The flow rate was very steady.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db. The data were exported to spreadsheet 2015-48-TSG-CTD-comp.xlsx.

The differences in latitude and longitude between the CTD casts and the TSG record when times were matched had median differences that were <0.0001°, and the maximum differences were 0.0008°, so the clock appears to have worked well and the matches made to construct the files were done correctly.

d.) Comparison of T, S and Fl from TSG and CTD data and rosette data

- T1 vs T2 The intake temperature sensor worked throughout the cruise. The heating in the loop averaged ~0.020Cº with a standard deviation of 0.03Cº.

- TSG vs CTD There were 20 casts that overlapped with TSG data.

The intake temperature was higher than the CTD temperature by an average of 0.0135Cº and standard deviation of 0.046Cº. But the median difference was 0.007Cº. When 3 outliers are excluded the TSG is high by an average of 0.008Cº. The average of the 7 cases with the lowest standard deviation in the intake temperature also gives the TSG to be high by 0.008Cº

The TSG salinity was found to be low by a median value of 0.095psu and standard deviation of 0.023. The results do not show any significant dependence on standard deviation in the TSG data.

The TSG fluorescence in volts is about 20% of the CTD fluorescence in ug/L; the ratio decreases as CTD Fluorescence rises. Fluorescence did not vary greatly through this cruise.

(See 2015-48-ctd-tsg-comp.xls.)

- TSG fluorescence versus Rosette CHL data

There were only 2 CHL samples taken. The CHL value was about 6X the TSG reading for 1 and about 3X for the other.

- Calibration History

The temperature and conductivity sensors were recalibrated in December 2013 and have been used for 7 previous cruises since then and for 4 that have occurred since then. The fluorometers used have varied and some of the cruises were on the Vector with no fluorometer.

- During 2014-21 the TSG salinity was found to be lower than loop samples and CTD salinity by ~0.03 but the difference varied with flow rate which was highly variable.  No recalibration was applied due to the variability in the comparisons and the fact that such a large drift in calibration on its first use seemed unlikely.

    - During 2014-19 the TSG salinity was found to be low by ~0.02, and the TSG temperature was found to be higher than the CTD temperature by ~0.005Cº.

    - 2014-22 results were not trusted.

    - 2015-01 salinity data were recalibrated by adding 0.02. Intake temperature was found to be higher than CTD temperature by from 0.005 Cº to 0.007 Cº.

    - During 2015-20 the salinity was found to be low by 0.025, the intake temperature high by ~0.008Cº.

    - During 2015-09 salinity was found to be low by about 0.03 and the intake temperature was higher than the CTD temperature by ~0.004Cº to ~0.006Cºf

    - During 2015-18 the TSG was on the Vector with no loops, no intake temp, and high gradients making comparisons with CTD data unreliable.

    - No useful comparison could be made using 2015-46 results.

    - For 2015-10 the TSG was lower than loops by 0.17 and lower than the CTD by 0.1 to 0.18. Recalibration was done by adding 0.18 to the salinity.

    - For 2015-21 the intake temperature was higher than the CTD temperature by a median of 0.011Cº. The salinity was low by ~0.18 but the comparison was noisy.

    - During 2014-61 the intake temperature malfunctioned. The TSG salinity was lower than the CTD by between 0.04 and 0.1 and lower than loops by 0.05. Recalibration added 0.014 to salinity.

Conclusions

1. The TSG clock worked well.

2. The TSG flow rate was steady and high.

3. The TSG fluorescence is given in volts. For a very rough estimate of fluorescence you can multiply by 5, but understand that the factor varies.

4. The history of this TSG since it was last calibrated shows a lot of variability, some of which may be explained by it being used on different ships. It is thought that bubbles in the Tully loop may explain low salinity values; evidence for this includes spikes that are always towards lower salinity values. The TSG set-up on the Tully varies from cruise to cruise, and it is possible that other factors such as weather conditions may be factors.

5. The two offshore cruises that bracket this one found very different results when comparing TSG salinity with loop samples and CTD data. The salinity comparison for this cruise falls between the two with salinity low by 0.095. While there is a lot of scatter in the fit, it is clear that the salinity is low and 0.095 looks like a good number to use. There are spikes in the salinity record that do suggest there are some bubbles. There may well be many small ones that are not obvious accounting for the overall low values.

6. The TSG temperature is close enough to the CTD temperature that recalibration is not justified given that minor mismatches in depth may account for at least some of the difference.

f.) Editing

The salinity has some obvious spikes to low values. CTDEDIT was used to reduce a few records at the beginning of the file as flow was being established. Single-point spikes in salinity were cleaned where they were not matched in the temperature trace.

g.) Recalibration

File 2015-48-tsg-recal.ccf was prepared to add 0.095 to the salinity channel and was applied to the *.EDT file.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan_Number, Temperature:Difference, Conductivity:Primary, Position:New, Flag

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats.

The file was saved as a TOB file.

A Standards Check and Header Check were run and no problems were found.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

##### SBE25 Data

The SBE25 was mounted on a dive camera. In some cases there were multiple files for a single dive.

These data are not suitable for the OSD Data Archive, but for use of the chief scientist in conjunction with the dive camera.

The RAW file names were copied to a different folder and renamed in standard format, but with a leading 9 in the event number, ex 2015-48-9004.hex. This is to distinguish this set of casts from those taken with the SBE911+ with the rosette.

- CONVERT

The configuration file was correct and was saved as 2015-48-dive-ctd.xmlcon.

All casts were converted using configuration file 2015-48-ctd.xmlcon.

Plots show that the channels all produced reasonable values except for the descent rate which is extremely high. Investigation could find no problem in the headers or configuration file to explain this – the sampling rate was correct. The data must have been averaged on acquisition, though it left no trail in the header. So the descent rate looks 8 times as high as it really is. Many of the usual processing steps will be inappropriate for 1s-averaged data.

- WILDEDIT

This step was skipped since any spikes will have been removed by the averaging already applied.

- WFILTER

The averaging on acquisition means these data do not have a lot of reversals but there are some. Also since the header indicates that the sampling rate is 8 hertz, whereas the file only contains 1 record per second, it is likely that the usual settings would have a much greater effect than usual. Applying them does remove a few spikes, but those spikes look like they are due to shed wakes and so it is better to use DELETE to remove them, rather than smooth them.

This step was skipped.

- ALIGNCTD

Tests were run on a few casts to see what alignment made the offset between the upcast and downcast DO traces resemble that for the temperature traces. An advance of 1s produced the best results; given the header error this is assumed to be effectively 8s.

ALIGNCTD was run on all casts to advance the DO channel by 1s.

- CELLTM

The usual setting for this step did not work well on these data and none was found that made any significant difference. The upcast and downcast traces were already in good correspondence, so this step is not needed.

- DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included).

- Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

There are two events with more than 1 file:

- Cast 9 had 4 files but only the 1st had any useful data. There is no upcast data.

- Cast 11 had 2 files. The first has nothing useful and the 2nd has downcast data with a big gap and no upcast. So 11 was renamed 11a and 11b was renamed 11.

There were other casts with gaps: 12, 13, 14 & 15.

Spreadsheet 2015-48_merge_headers.csv was prepared to add station names, event numbers, latitude and longitude. These steps are finicky – format has to be just right.

Program MERGE:CSV FILE TO HEADERS was run to add that information to the headers of the IOS files.

The MRG files were edited to change the time interval from 0.1250 to 1.0000.

ADD TIME CHANNEL was run to add 7 hours to time to get UTC (since the original times appear to be in PDT) and to add times to every record. A few files were checked against the DIVE log records and the end times compare well with the log.

- Checking Headers

A cross-reference listing was produced and the entries look reasonable, though no leg was available for these casts.

A track plot was produced and added to the end of this report.

HEADER CHECK was run. There are some negative pressures that clearly correspond to a time when the CTD was out of the water.  The descent rate is much too high, but this is believed to be because the data were averaged on acquisition. Dividing the descent rate by 8 produces believable values and the header shows there were 8 records per second.

The surface check shows that most casts started below 10m. Of the 3 that did not, 2 have negative pressures (-0.8 and -0.5db) but also very low conductivity so that the CTD might have been out of water or pumps were not working. The upcasts are not useful in assessing the accuracy of the pressure as most do not include near-surface records and some have negative values but the data suggests the CTD was out of water. There is no evidence to enable a judgment on the accuracy.

- DELETE

A test run of DELETE demonstrated that the record would be fragmented and need heavy editing; the upcast sections may well be of interest to the researchers as may other data that would be removed.

- Final Steps

REMOVE was run to remove channels Conductivity:Primary, Oxygen:Voltage:SBE and Flag.

DERIVED QUANTITES was run to calculate Depth.

HEADER EDIT was run to change the format for Pressure and Depth and to add the following comments to the headers:

Data Processing Notes:

These data were processed for the use of the researchers and are not intended

to go into the IOS DATA ARCHIVE. They will be archived elsewhere.

The CTD was an SBE25 mounted on a drop-camera frame.

There were frequent interruptions in the data acquisition.

No data were removed in processing.

File names have a leading 9 in the event number section to distinguish them from

CTD casts that occurred at other sites and deployed an SBE911+.

---------------------------------------------------------------------------------

A cross-reference listing was produced. .

These files were then sent to Katie Gayle for her use. The files are not suitable for the OSD Data Archive, given low and highly variable descent rates and many stops and gaps.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? |
| --- | --- | --- | --- | --- |
|  | SEABIRD | 911+ | 0506 | Yes |
|  | SEABIRD | 25 | 0464 | No |

| Calibration Information SBE 911+ CTD |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 31Jan2013 | Factory |  |  |
| Conductivity | 1763 | 1Jan2014 | Factory |  |  |
| Secondary Temp. | 5013 | 27Feb2013 | Factory |  |  |
| Secondary Cond. | 3394 | 3Jan2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 1438 | 3Jan2014 | Factory |  |  |
| SeaPoint Fluor. | 3640 | n/a |  |  |  |
| Pressure Sensor | 506 | 30Dec2013 | Factory |  |  |
| Altimeter | 62354 | n/a |  |  |  |

| Calibration Information SBE25 CTD |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 23Dec2014 | Factory |  |  |
| Conductivity | 2173 | 19Dec2014 | Factory |  |  |
| SBE 43 DO sensor | 0766 | 23Dec2014 | Factory |  |  |
| WetLabs Eco Fluor. | 2215 | 2July2015 | IOS |  |  |
| Pressure Sensor | 0464 | 23Dec2014 | Factory |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3363       Cruise ID#:	2015-48

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3363 | 28Dec13 | Factory |  |  |
| Conductivity | 3363 | 28Dec13 | Factory |  |  |
| WetLabs Fluorometer | WS3S-953P | ? |  |  |  |
| Temperature:Secondary | ? | ? |  |  |  |
| Flow meter | ? | n/a |  |  |  |

Event Numbers for SBE25 files mounted with Dive Camera.