## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27 Mar 2025 | Updated channel names & formats in TOB files. G.G. |
| 20 October 2017 | Corrected PROJECT entry in Underway files. Roy Hourston. |

## PROCESSING NOTES

Cruise: 2015-54

Agency: Ocean Sciences Division

Location: Douglas Channel

Project: WC – Douglas Channel

Party Chief: Wright C. & Vagle S.

Platform: Vector

Date: October 13, 2015 – November 9, 2015

Processed by: Germaine Gatien

Date of Processing: 5 February 2016 – 12 April 2016

Number of original HEX files: 284

Number of CTD files: 284	Number of bottle files:	62

Number of TSG files: 5

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0443) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#0997) on the secondary pump, a SeaPoint Fluorometer (#3640) on the primary pump, a Biospherical QSP-400 PAR sensor (#4565) and an altimeter (#1253).

The data logging computer was Vector CTD Laptop.

The data acquisition program was Seasave.

The CTD deck unit was an SBE model 11+, #425.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log book was generally in good order and included many notes including cases where header entries were in error. There were a few errors in event numbers in the Sampling/Rosette log sheets and some duplication of sample numbers, but given the complexity of the cruise the records were good.

As has occurred on many cruises this year, the station name entries in the SeaBird headers had the wrong format, STN instead of STATION. This leads to a problem in conversion of the files to IOS Header format.

Comparison of CTD salinity to bottle samples suggests that the CTD salinity is low, but it is likely that the bottle values are high for 2 reasons. The samples were analyzed within 3 weeks of the end of the cruise, but the cruise lasted 4 weeks, so there was likely some evaporation and/or adsorption of samples, which would both raise salinity. Incomplete flushing of Niskin bottles is also likely due to very steady descent rates and high vertical salinity gradients for many casts; this would lead to the Niskin bottles containing water from lower in the water column where salinity is higher than the ambient salinity at the bottle stop level. The best evidence from the history of the sensors suggests salinity is within 0.002. No recalibration is justified.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.6 mL/L from 0 to 20db

±0.3 mL/L from 20db to 100db

±0.1 mL/L from 100db to 300db

±0.06 mL/L below 300db

As usual for these instruments, the SeaPoint fluorometer values were higher than the extracted CHL samples for CHL<1ug/L, then gradually fell relative to CHL settling at about 50% of the extracted chlorophyll values for CHL>4ug/L.

The Vector TSG system is new and we have much to learn about it. There were no loop samples taken so we have no indication about how well the TSG itself is performing. The absence of a flow meter and an intake temperature sensor limits how these data can be interpreted. Comparison with CTD casts is useful, but we will need a lot of experience before we can use the comparison to recalibrate TSG salinity or to make an estimate of ambient temperature based on TSG temperature. Taking lots of loop samples (and analyzing them promptly) and using the system in a variety of environments will speed up the process. Installation of a flow meter would help greatly by removing from comparisons sections with very low or unsteady flow. Zero flow is usually obvious in the record, but an erratic or very low flow rate is not.

No recalibration was applied to the TSG data. A rough estimate is that the TSG temperature is higher than ambient temperatures by ~0.1Cº, but that would vary with intake temperature and flow rate in the loop. No estimate can be made of salinity except that TSG salinity was lower than CTD salinity for 90% of the casts.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

Some file names had the wrong format – missing a digit in the event number section. This was corrected in the converted files.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. There were many comments including details of errors in station names in some files. There were many comments about problems with sounder readings that were out of sync with the altimeter readings and with the bridge sounder. So the bottom depths entered in the headers may be of questionable value.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked.

The XMLCON files did not change through the cruise. The calibration constants were checked for all instruments and the only corrections made were to factor E in the dissolved oxygen parameter (based on 2015-10 tests).

The corrected file was saved as 2015-54-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2015-54-ctd.xmlcon to create CNV files.

The station names were corrected in the files with errors noted in the log book.

A few casts were examined.  All expected channels are present. The primary and secondary temperature and conductivity channels are reasonably close during downcasts, but as usual the upcasts differ more due to noise in both channels. During bottle stops once the initial noisy patch is over, the channels pairs are close again.  Dissolved oxygen, fluorescence, PAR, altimetry and transmissivity profiles all look normal. The descent rate was generally high and varied from extremely steady to very noisy with some complete reversals of direction.

There is an error in the station name format, STN instead of STATION, which will make the conversion routine skip that line. So Roy Hourston ran a routine to change the header name in the converted files, both CNV and ROS. A change to IOS SHELL is being investigated so that either entry will work since this is a frequent error.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2015-54-ctd.xmlcon.

The ROS files were converted to IOS format.  They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and no problems were seen that required editing.

A preliminary header check was done and no problems were found. Fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv.

Before adding sample numbers it was necessary to sort out duplicated sample numbers, as follows:

- The sample numbers used for event #40 had all been used during cast #38, so a leading 9 for the sample #40 samples will be needed. This will apply only to DIC and SPC

- Samples 95 and 96 were used in both events #41 and #38, so 995 and 996 will be used for event #41. Some analysts have already made that adjustment, but samples coming in later will need to be adjusted – TOC, DOC, CDOM, DELTA18, PHYTO.

- Some samples for event #364 had sample numbers 9364-9374 but the rosette log shows 538 to 549. In this case the sample numbers in the analyses will have to be adjusted.  The nutrients have the right sample numbers and there doesn’t seem to be any salinity though there was supposed to be one. Chlorophyll and oxygen analysts were informed and they changed the sample numbers in the bottle files and the analysts’ spreadsheets.

Sample numbers were added to the file based on the rosette log records.

A few problems were encountered:

- No sampling from events #1, 16, 95. Removed from list.

- 19. Called 20 on the rosette sheet, but 19 in the main log. 20 was not a 911+ event.

- 28. Called 27 in the rosette log but 28 in the CTD log. #27 was a mooring event, The bottle file name was changed to 28.

- Similarly, 28 should be 29.

- Event #64 has a rosette cast with all bottles fired at the surface -no sampling. No mention in log about this but presumed a test. No processing needed so dropped from list.

- Event #95 – 2 bottles fired but no sampling. No processing needed so dropped from list.

- Event #348 – Rosette sheet said event #478 which doesn’t exist. CHL sample misnamed – others ok. CHL spreadsheet was corrected.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine. Those files were bin averaged on bottle number.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2015-54-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2015-54chl*.xls. The file included comments, flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2015-54chl.csv and the file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2015-54oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2015-54oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2015-54SAL.xls. The analysis was done within 21 to 42 days of collection. The files were simplified and saved as 2015-54sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2015-54nuts.xls.

Then the file was simplified, reordered on sample numbers and saved as 2015-54-nuts.csv. The file was converted to individual NUT files.

All files for event #41 were reordered on increasing sample number.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. A number of problems were found:

- File #7 – drop bottle 14 later – no sampling

- File #17 – Only DO had flag 9 for sample 27, but all other sampling was moved to another bottle, so this is fine.

- File 19 – Many lines should be dropped later – only bottles 1- 4 sampled.

- File 38 – CHL was missing because it was identified as from event 40 and should be 38. The spreadsheet was updated.

- File 41 – Nutrient #95 should be 995 because there was a sample # duplication.

- File 150 – SAL sample #349 – not on rosette log and the value 32.8085 is completely out of line.  – think this should be sample 539 which is missing from event 364. Some samples were named wrong for that cast which probably led to this sample going astray. Analyst made these changes.

- File 349 – The only sample from this file was supposed to be duplicate CHL but it is missing. A CHL file for that cast was created with flag 1 and comment.

- File 364 – As noted above some samples were given the wrong sample numbers for this cast.

CLEAN was run on the MRG files to add 0 flags to empty flag channels and to update header limits. (MRGCLN2)

##### Compare

Salinity

Compare was run with pressure as reference channel.

There is a lot of scatter in the comparison especially in the top 5m. All of the major outliers are from the top 4m with 2 exceptions. Those shallow outliers are all associated with very high standard deviation in the CTD salinity. The near-surface salinity gradients are very high for these casts. The major outliers below 5m are:

- Event 41, Sample #995 – The bottle value is much lower than the CTD and judging by the bottle values for salinity, dissolved oxygen and nutrients, this bottle was a misfire, having fired near the surface. All values should be replaced with pad values and flag 5.

- Event 150, Sample #349 – The rosette log sheet indicates that only one salinity sample was to be gathered from this cast, #348, at the surface; that sample compared reasonably well with the CTD. The salinity spreadsheet included a sample from bottle #349 which was also a surface bottle. But that sample had salinity higher than the CTD by 23psu and higher than the other surface sample by ~29. The value looks like a near-bottom bottle. Suggest pad value for sample #349, flag 5, comment “either mislabelled or misfired. No such sample indicated on rosette log”

Both CTD salinity channels read lower than the bottles with the exception of some surface bottles and one bottle that is believed to have misfired. The differences are clearly dependent on pressure in a way that suggests that flushing was a major problem. Above 250m the differences get larger as we near the surface, likely due to increasing salinity gradients. When cases are excluded that involve bottles fired near the bottom, bottles above 260m and bottles for which the CTD salinity standard deviation >0.0008, there are only 12 bottles left. The primary CTD salinity was lower than bottles by an average of 0.0080 and the secondary by 0.0055; the standard deviation in both was ~0.003.

Examination of a few of those 12 cases focused on whether poor flushing could account for them reading low – was the gradient sufficiently high? Many of them were easy to explain because salinity matching that of the bottles was found during the bottle stop as shed wakes passed through. A few others have values that match that of the CTD about 10m deeper, so if the flushing is incomplete, this would be a possible explanation. In another case there were two bottles at the same depth; one bottle value was easy to explain while the other was out of line with the other bottle and required that the Niskin contain water from about 30m deeper than the CTD. So shed wake corruption of the contents of the Niskin and the likelihood that the Niskin bottles carry water from deeper in the water column explains much, but probably not all, of the difference between the CTD and bottles for these 12 outliers.

There may also be a little adsorption and evaporation of samples. While salinity analysis was done within 3 weeks of the end of the cruise, it was a long cruise. So some samples waited more than 5 weeks and the differences that are hardest to explain are from the earlier casts. A plot of differences against time also suggests this could be a factor though interpretation is not simple given variations in personnel, sampling conditions and geography.

There were no duplicates from this cruise but there were frequently two bottles fired at the same depth, so an analysis of these was done. The differences between bottles at the same depth were calculated and displayed with pressure and standard deviation in the CTD data. Those from 0 to 5m differed greatly and had very high standard deviations in the CTD salinity. This is expected due to large temporal and spatial variation and inefficient flushing of Niskin bottles in the presence of high near-surface salinity gradients.

The differences below 5m are much smaller and may be due to a combination of small errors such as salinity analysis error, evaporation or adsorption of samples and Niskin bottle flushing inefficiency. Removing the severe outliers from casts #41 and 150, the following results were obtained:

| Differences between bottles fired at same depth |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | average | median | Std dev | # of pairs |
| all | 0.1351 | 0.0075 | 0.3786 | 39 |
| Press<5db | 0.4445 | 0.2369 | 0.6287 | 11 |
| Press>5db | 0.0135 | 0.0044 | 0.0324 | 28 |
| Press>100db | 0.0060 | 0.0041 | 0.0068 | 13 |

There were 3 cases below 100db that were a little out of line from the others.

- Sample #201 from cast #57 had been flagged by the analyst due to salt crystals on the bottle. The flag should be changed to 4 with an extra comment.

- Sample #195 from cast #57 is higher than any salinity recorded by the CTD while the other sample from that depth compared well with the CTD. Suggest flag 3.

- Samples #46 and 47 from cast #28 are both higher than the CTD, but the standard deviation in the CTD salinity is high at 0.0018 and there is clear evidence of many shed wakes passing through during the stop. So the bottle values are likely good, reflecting the contents of Niskin bottles which probably came from somewhat lower in the water column.

These results are good. When the 3 outliers above are removed the results below 100m are even better:

| average | median | Std dev | # of pairs |
| --- | --- | --- | --- |
| 0.0028 | 0.0012 | 0.0029 | 10 |

For full details for the COMPARE run see file 2015-54-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There was only 1 noteworthy outlier, sample 995 from cast #41, which has already been identified as a misfired bottle.

There are no bottle samples with DO<2mL/L, so getting a fit with a reliable offset is a problem.

- With the offset free

CTD DO Corrected = CTD DO * 1.0222 +0.0376   (R2 = 0.66)

- With the offset forced to equal to -0.0309 which was the offset found during 2015-21 which used this equipment and included many low DO samples, the fit was:

CTD DO Corrected = CTD DO * 1.0242 +0.0309   (R2 = 0.65)

The fit from 2015-21 was mostly in open waters, providing the best comparison available for this equipment:

CTD DO Corrected = CTD DO * 1.0225 +0.0309

Since poor flushing is apt to lead to bottles being lower than CTD DO, there is likely to be some overcorrection if we use the fit from the 2015-54 comparison. If we have a CTD reading of 6mL/L, using fit 1 will produce a value of 6.1708mL/L and fit 2 will produce 6.1755mL/L. The 2015-21 fit would produce a corrected value of 6.1659mL/L.

For 2015-19 (the cruise previous to this one) the same equipment was used in an area with some flushing problems. The offset was fixed to match 2015-21 and it resulted in a slope very close to that of 2015-21, so the 2015-21 fit was used. For 2015-54 the correction is a little higher for both fits, which is likely due to flushing inefficiency since so many of the casts were in protected areas with very smooth descent rates.2015-19 had a mix of conditions.

For more detail see document 2015-54-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. A few more outliers were investigated:

- Event 45- sample 133 - surface DO looks a bit high but the salinity is low compared to the CTD, so likely there is a large gradient at the surface and the physical distance between the CTD and the Niskin bottle is significant. No flag to be added.

- Event 330 – sample 444 - The 250m bottle DO value is high by about 0.7mL/L compared to the CTD data (after recalibration) and looks quite close to the 100m bottle. All other samples and CTD data for that Niskin look closer to the 513db data than the 100m data. So this does not look like a mis-fire and it is unlikely that it came from the wrong bottle. Flushing would have the opposite effect, and the local gradients are not high, so there does not seem to be any nearby source for such water. In the absence of an explanation, flag 3 looks appropriate.

- Event 333 – sample 459, 20m – DO value is somewhat out of line in the comparison to CTD DO. DO, Extracted CHL and nutrients all look like the 10m values, but the CTD data do not. There is no suggestion that poor flushing can account for this or a high DO gradient just above the CTD. So it looks like a mis-sample, but given that the gradients are not very high, there is some room for doubt, so a 3 flag looks appropriate.

- Event #352 sample 507 – the DO at 75m looks like that at 20m but the nutrients and temperature have similar profiles, so this is likely fine. No flag suggested.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor. There was a small range of CHL values, 0.02 – 1.75ug/L, and only 3 of the 113 samples had CHL >1ug/L.

A plot of FL/CHL versus CHL has the usual shape. The ratio at the low end is somewhat higher than usual, but that is likely due to some samples with very low CHL values. At the high end the ratio settles to about 0.5 which is typical of this type of sensor.

A plot of FL versus CHL forced through the origin has a slope of 1.08, but the R2 value is very low with low CHL data all falling above the trend line. When the offset is left free the fit has a slope and offset of 0.70 and 0.59, but the fit does not look convincing. Excluding 1 outlier provides a better fit with slope 0.76, offset 0.2 and R2=0.66.

For full details of the comparison see file 2015-54-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Based on 2015-21 results ALIGNCTD was run using a setting of +3.5s to advance the DO signal. A few casts were checked to ensure the results were good. As usual, some features were overcorrected and others not corrected enough, but the setting looks appropriate overall.

ALIGNCTD was run on all casts using +3.5s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. Tests on previous cruises using these sensors showed the default setting of (α = 0.0245, β=9.5) did the best job and it does improve the data for both conductivity channels for these data.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. While none are very deep, it does allow us to check for sudden changes. Differences from earlier cruises using the same equipment are also shown with dark shading.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2015-10-0039 | 500 | -0.0012 | -0.00034 | -0.0024 | High, Moderate |
|  | 1000 | -0.0010 | -0.00015 | -0.0005 | “ |
|  | 1900 | -0.0011 | -0.00007 | +0.0003 | “ |
| 2015-10-0099 | 500 | -0.0011 | -0.00016 | -0.0009 | High, Very Noisy |
|  | 1000 | -0.0010 | -0.00006 | +0.0004 | “ |
|  | 1900 | -0.0013 | -0.00003 | +0.0010 | “ |
| 2015-19-0075 | 350 | -0.0011 | -0.00003 | +0.0008 | High, Moderate |
| 2015-19-0093 | 350 | -0.0010 | -0.00003 | +0.0007 | High, V.Steady |
|  | 400 | -0.0012 | -0.00002 | +0.0008 | “ |
| 2015-19-0097 | 350 | -0.0004 | -0.00003 | +0.0008 | High, V.Steady |
|  | 400 | -0.0003XN | -0.00006VN | +0.0015VN | “ |
| 2015-54-0057 | 450 | -0.0011 | +0.00003 | +0.0013 | High, Moderate |
| 2015-54-0124 | 450 | -0.0007 | +0.00016 | +0.0024 | High, Steady |
| 2015-54-0127 | 450 | -0.0008 N | +0.00015 | +0.0022 | High, Steady |
| 2015-54-0236 | 450 | -0.0009 | +0.00019 | +0.0030 | High, Moderate |
| 2015-54-0330 | 450 | -0.0010 | +0.00020 | +0.0030 | High, Steady |
| 2015-54-0373 | 450 | -0.0008 | +0.00025 | +0.0033 | High, Steady |

The temperature differences look like those in other recent cruises, but the conductivity and salinity differences are increasing with time.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. A few problems were found in station names – some were missing since the fix to “STN” to “STATION” must have missed a few casts. Those were fixed in the CNV file and conversion and CLEAN were rerun. A few other corrections were needed:

- 171-173 were changed from MI** to GI**.

- 183 SC59W changed to SC59

- 269 was changed from GC74 to GC68

- 336 was changed from VP39 to VP39A

The log time entries were generally a little later by up to 8 minutes, but most differences were  < 5 minutes.  It has been noted during some cruises that the CTD computer was started much earlier than the beginning of the cast, so that the start time and position in the file may not be reliable. So, it is good to see that was not the case for this cruise.

A header check was run. No off-scale fluorescence values were found. There are many “unreal” values in the primary T and C sensor ranges as well as other evidence of spikes. Cast #214 was examined based on these results and the pumps were off for the first few metres which may account for some of the bad values.

The cruise track was plotted. Because there are so many casts in a small area, it was broken up into 3 sections. These were added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 1.05db which is reasonable for this project. There were a number of pressure readings while the CTD was on deck and they are mostly between 0.6 and 0.8; there is no obvious systematic difference between readings done before the cast and after, but there were not many taken after. However, there was one case where the “after” reading was 2.6db. Examination of data suggests the CTD was in the water at 2.6db for that cast, so there was likely an error in the log entry. There are a number of cases when the CTD appears to be in water with the pressure being ~0.6db but with pumps turned off it is a little hard to judge except by looking at transmissivity and that picture is not completely clear since the values close to the surface may be either very high or very low and the transmissometer calibration may be incorrect.

SeaBird recommend that pressure tests be run after the CTD has had the power on for about 5 hours somewhere with temperature fairly constant. Those conditions would not apply to these deck readings so the results may not be reliable, though the errors may not be significant. Overall, there does seem to be evidence that the pressure sensor is reading a little high at the surface, so subtracting 0.6db is appropriate The CTD was often started just as it entered the water, so this makes sense.

The altimeter and depth readings from the headers of the CLN and SAMAVG files were exported to spreadsheets to check that they are reasonable. A rough check was made by subtracting the altimeter reading from the depth entered in the header to find the maximum depth sampled. That was then compared with the maximum pressure recorded in the files using a rough estimate that depth is ~99% of pressure. When those two estimates differed by more than 5m, a plot was made to see if the altimetry reading was reasonable. Despite many spikes near the bottom, the header entry for the altimeter was always reasonable. Given many comments in the log about problems with sounder readings, the differences are more likely due to bad depth readings or the ship drifting between readings.

In some cases there was no altimetry header entry. In about half of these cases, plots hint that the CTD reached about 7m from the bottom, but the traces are very messy and the appearance is unusual. It is suspicious that in each of these cases the estimate seems to be 7 to 7.5m.

There are many discrepancies between the depth entries in the log and those in the file headers. In some cases both look wrong. From the many notes in the log book it is clear that the sounder was not providing reliable data. In some cases second readings were entered below the initial one, but it appears that these are based on adding the altimeter readings to the maximum pressures, so for deep casts there could be a substantial error. A spreadsheet was prepared with maximum pressure, altimeter header reading, log bottom depth entry and then a calculation was made as follows:

Difference = Pressure *0.99 (a rough conversion to depth) + altimeter header – depth header

Where this differences was >4m, the cast was examined in more detail to see if a better estimate could be made. There is reason to expect some difference as the altimetry header is an average over the bottom 2m and the bottom depth can vary through a cast. Some discrepancies were found that led to changing the file entries:

- In some cases there were differences between the log and header entries that look like typos made in entering the depth in the files.

- In the case of the many casts at station MS1 the header depth was not changed but the log entries do change. The log entries look better.

- Where the estimated water depth based on maximum pressure and altimetry looks more reasonable than the log or header entry, the derived value was used.

- In some cases the altimetry was too noisy at the bottom, but a reading could be taken further up that allowed for bottom depth estimate. This allowed a bottom depth estimate and where those look more appropriate than the header entry, it was adjusted.

File 2015-54-altimeter-ctd.csv contains a record of evidence and suggested changes tow water depth.

File 2015-54-mergc.csv contains just the file name and the water depth (either original or corrected, as needed.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and as usual, there is a lot of variability with the up and downcast traces sometimes closer than temperature and sometimes further apart. This is likely due to varying vertical gradients. Overall the choice made earlier looks appropriate, so no further alignment will be applied.

Conductivity

Tests were run on 5 casts using a variety of shifts and the best results were with the same parameters used during 2015-19: -0.2 records for the primary and -0.8 records for the secondary looked best.

SHIFT was run twice on all casts using those settings.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

The DEL files were copied to *.EDT.

##### Other Comparisons

Previous experience with these sensors –

The sensors were used during 2015-10, 2015-21 and 2015-19 which immediately preceded this cruise. The primary and secondary salinity were both found to be low by~ 0.002 for 2015-10. They were low by 0.0045 and 0.0032, respectively, for 2015-21, but delayed salinity analysis reduced the reliability of those results. For 2015-19 salinity was low by 0.0051 and 0.0076 but analysis was delayed and salt crystals were found on some bottles and incomplete flushing of Niskin bottles was likely a problem. No problems were found with the pressure for any of the cruises. Dissolved oxygen fits were reasonably similar given very different pressure and DO ranges included in them.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. In the Strait of Georgia the temperature was above the maximum between 40 and 80m. In the Douglas Channel area there was no appropriate climatology for most casts; temperature and salinity fell within the climatology where it did except for some slightly high temperatures at the bottom.

These excursions are not considered evidence of calibration problems.

Repeat Casts – There were some repeat casts but they were too shallow to expect good repeatability.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The primary channels were selected for archiving during 2015-10 and 2015-21 and the secondary were chosen for 2015-19.  The secondary channels have slightly less spiking in these files, so they were selected for editing and archiving.

These files start close to the surface so in the top 1 or 2m the pumps are off and editing would be easier if the pump status were displayed. It is preferred to minimize the number of channels displayed, so instead CLEAN was run to remove the secondary temperature wherever the pumps are off. This makes it clear in the editor while having the full salinity trace is useful. (Extension CLNT)

CTDEDIT was used to remove data collected while the pumps were off, remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. All files required some editing.

##### Initial Recalibration

It appears that the pressure is slightly high, so the offset will be adjusted by subtracting 0.6db.

It is harder to judge salinity calibration, but based on the comparison and history, it is likely that the sensors are within ±0.002. No recalibration will be applied.

The DO recalibration used for cruise 2015-21 looks reasonably close to that found for the current cruise and since it was based on a wider DO range including many bottles with DO<2mL/L values, the offset is likely more reliable. The parameters from 2015-21 were also selected for recalibration of the 2015-19 DO data.

CALIBRATE was run using file 2015-54-recal1.ccf to subtract 0.6db to the pressure and to correct the Oxygen:Dissolved:SBE channel in the SAM and MRGCLN2 files using:

CTD DO Corrected = CTD DO * 1.0225 +0.0309

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly. The average of differences in the fit once outliers were removed was +0.006mL/L, but the standard deviation is 0.035mL/L.

See file 2015-54-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed, the CTD DO was higher than the bottles by an average of ~0.04mL/L and standard deviation of 0.08mL/L; several methods of removing outliers led to very similar averages. The results vary with depth with the CTD reading close to bottles right at the surface and below 100db and higher than bottles between 10db and 100db. This pattern is likely due to incomplete flushing of the Niskin bottles in the presence of significant DO gradients between 10db and 100db; slow response time of the CTD sensor could partly account for this pattern as well. No further recalibration of DO is justified.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some small unstable features, but that is normal for this region in near-shore and well-mixed casts. There are 2 casts with some zero values for transmissivity (#28, #29). This is likely due to some biological material temporarily stuck on the lens, but no editing was applied; these data are always considered nominal.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

Merge CSV file to Header was run to correct water depths, using file 2015-54-merge.csv as input.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and unedited except

that some records were removed in editing temperature and salinity.

The deployment method used for the CTD did not include a soak at 10m.

Acquisition began as the CTD entered the water with pumps off. Pumps

were turned on at approximately 2m and the CTD was generally kept at

that depth for about 1.5 to 2 minutes before the full cast was run.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision.

There were many problems with the sounder so many water depth entries were

corrected based on log records or maximum pressure sampled and altimetry.

For details see file 2015-54-water_depth_study.xlsx.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.6 mL/L from 0 to 20db

±0.3 mL/L from 20db to 100db

±0.1 mL/L from 100db to 300db

±0.06 mL/L below 300db

For details on the processing see the report: 2015-54-proc.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and it was found that one cast was mis-named as 2015-54-0061; it was changed to 2015-54-0062. No further problems were found.

A file list was produced.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values were highly variable ranging from 65% to 100%, with many at ~90%. The casts furthest offshore had the highest saturation rates between 85 and 98%, with the highest at cast #304 which was furthest offshore. It is not unusual to see lower values in inlets and well-mixed waters. Checks were made to ensure the low values were not due to an error in calibration so bottle files were re-examined and the CTD dissolved oxygen values at the surface are generally higher than the bottles. It is possible that the method used in deployment may have depressed the CTD DO slightly at the surface as the pumps had not been on long before the CTD was lowered and there was no soak at 10m.

##### Final Bottle Files

CALIBRATE was run on the MRGCLN2 files and then the MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

Data from the MRGREO files were exported to a spreadsheet and compared with rosette sheets. A few problems were found:

- No sample numbers were assigned to most of the bottles from event #19 so those lines were removed. There were 4 bottles sampled.

- There is no chlorophyll sample for cast #349 and that was the only sample for that cast. The cast was removed. It can be created later should the sample be found. KEEP MRGREO – cast removed from cast list.

- There were a few other cases where there were no samples corresponding to ticks on the rosette log but there is no sign of them in the analyst’s spreadsheets so it is assumed they were skipped.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check and a header check were run on all files and no errors were found.

The track plot looks ok.

File 2015-54-0183.MRGREO contains no samples so no CHE file was created.

Plots of each file were examined to ensure no problems had crept in. No problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

##### Thermosalinograph Data

There were no loop samples, flow meter or intake thermistor.

a.) Checking calibrations

There were 5 TSG files. They had non-standard names. All contain some data.

The configuration file was renamed as 2015-54-tsg.xmlcon. No errors were found in the calibration parameters.

b.) Conversion of Files

The HEX files were converted to a SeaBird CNV file.

The CNV files were converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots were produced.  At the beginning of files #2 and 4 there are sections during which it appears there was no flow to the TSG. Otherwise the records look good. There is a lot of variability in the salinity but this is expected from a cruise moving in and out of inlets. Where the data have less variability there is no sign of the sort of spiking often seen in TSG salinity and thought to be due to bubbles.

The track plot looks fine. The plot was added to the end of this report.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 3db and exported to a spreadsheet which was saved as 2015-54-ctd3-tsg-comp.xlsx.

Later 2db data were selected for comparison – there are fewer casts available for comparison at that depth but it is closer to the intake depth.

There are some gaps in the TSG records.

There were 270 casts which overlapped with TSG files and had data in the 3m bin. Of these, 4 occurred at a time when it appears that the flow to the TSG was off, so they were removed from the comparison.

The TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for lab temperature and salinity and the files were reduced to the times of CTD files.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The median differences in latitude and longitude are ~0.0001° and all differences were ≤0.0009° with most much smaller.

d.) Comparison of T and S from TSG and CTD data

- TSG vs CTD

There is a lot of noise in the differences between the CTD and TSG temperature and salinity. This is not surprising given the nature of the cruise – small differences in time and imperfect matches in the depth from which the data were selected can account for this. When the data were ordered on the standard deviation in the TSG temperature data over a 2-minute window and the 50 casts with the lowest standard deviation selected, the median difference shows the TSG temperature to be high by about 0.1Cº. Reducing the number of casts included does not make a lot of difference, with 0.0988 for 10 casts, 0.0938 for 20, 0.0973 for 30 and 0.0995 for 40.

A similar analysis was done for salinity but the number of casts included was much more significant with a trend towards smaller differences as larger standard deviations were allowed, with the TSG salinity always reading lower than the CTD salinity, with a median difference of 0.0784 with 10 casts and 0.0936 with 50casts.

The median differences using all casts shows the TSG temperature to be high by 0.076 Cº (standard deviation 0.364 Cº) and the salinity is low by 0.019 (standard deviation 2.75)

A second analysis was done using CTD data from 2db. The comparison was remarkably similar particularly for temperature, likely because the heating in the loop is much greater than the change in temperature over 1db. The TSG salinity is closer to the 3db CTD salinity than that from 2db, but that is not necessarily a reflection of the level from which the loop is drawing water. There may be errors due to small bubbles or calibration drift.

(For more details see 2015-54-ctd3-tsg-comp.xlsx.)

- Calibration History

The TSG temperature and conductivity were recalibrated in February 2013 and have been used for 3 other cruises since then. During 2015-17 in the same region the temperature was found to be high by 0.1Cº and the salinity was close to the CTD, on average. During 2015-18 the TSG lab temperatures were much higher than the CTD temperatures at any depth for which we had data. The intake water had a temperature close to air temperatures so it was not expected that there would be a lot of heating in the loop. Similarly, the TSG salinity was much lower than the CTD salinity and there was no evidence of bubbles. There could have been some problem with fresh water getting into the system or else the water was being drawn from very close to the surface. A TSG malfunction might be the explanation. For 2015-19 there were sections of data with evidence that the flow had stopped and other sections where salinity data looked bad but temperature seemed ok. That could be due to fresh water getting into the loop, flow rate variations reducing heating in the loop or malfunction of the TSG itself.

Conclusions

1. The TSG clock appears to have worked well.

2. The salinity differences rise steadily as the standard deviation in the TSG salinity rises. Perhaps the high variability is associated with areas that have more bubbles; bubbles in the loop water could lower the salinity measured by the TSG. There is no evidence of bubbles in the traces, though perhaps they are present but do not stand out in sections of noisy salinity. Another possible explanation is that the highest variability is found in areas with high near-surface gradients, so that slight mismatches in depth have a larger effect on differences. For the lowest standard deviations the salinity is low by ~0.07psu.

3. The temperature differences do not vary as strongly with standard deviation in the TSG temperature, with the TSG higher than the CTD by roughly 0.1Cº. This is the same as seen during 2015-17. The heating in the loop looks reasonable based on experience from the Tully. While the loop routing and length are different on the 2 ships, and they tend to sample regions with different temperature ranges, it seems reasonable that the heating on the Vector is a little lower than typically seen on the Tully (~0.15 Cº to 0.22 Cº).

4. The calibration history for this CTD is not very useful since the geography of this region is so different from the Strait of Georgia, and we have no direct information about flow rates on any of the cruises.

3. There is no reliable evidence from this cruise on which to base a recalibration. The best evidence about salinity is from cruise 2015-17 when the TSG salinity was found to be close to the CTD salinity. For temperature the heating in the loop depends on flow rate and intake temperature. Overall we might estimate that ambient temperatures are lower than the lab temperature by about 0.1Cº but with no information about flow rate variations it seems unwise to recalibrate, thus suggesting more confidence in the result than justified. No recalibration will be applied these data.

4. The Vector TSG system is new and we have much to learn about it. There are no loop samples so we have no indication about how well the TSG itself is performing. Taken together with the absence of a flow rate meter and an intake temperature sensor, we will need a lot of experience before we can justify interpreting differences from CTD casts to calibration drift. Frequent recalibration of the TSG, taking lots of loop samples and using the system in a variety of environments will speed up the process.

f.) Editing

The ATC file was opened in CTDEDIT and data were removed that were clearly bad. These included:

- File 2015-54-0001 had a section at the beginning where flow appears to have been zero. Temperature and salinity data were removed from records #1-589.

- File 2015-54-0004 had a section of bad temperature where flow is clearly zero. Temperature and salinity data were removed from records #2515-3432.

Editing notes were made in the headers of affected files.

In the comparison with CTD casts 4 cases were rejected as extreme outliers. These CTD casts all occurred during the time periods found to be bad independently in the editing process, and there were no other CTD casts during those periods.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Conductivity:Primary, Flag and Position:New channels.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats.

Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series plots; they look fine.

# Particulars

Leg 1:

1. All bottles tripped as test and no problems– no bottle files needed.

6. Deck pressure check 1.0 start; 0.8end.

10. Station name wrong in file- LS197 should be LS97 – FIXED in converted file.

16. 4 bottles tested – no problems with them. There was no sampling – no bottle file needed.

17. Bottle #18 fired too early, not 30s wait. Bottle #13 fired but no water – samples from bottle #12.

19. Test all bottles – all normal. Some sampling.

29, Winch lost pressure feed at ~69m for a few minutes.

33. Deck test 0.68db.

40/41. Flat calm, raining

51. Deck test 1.1 start, 0.6 end.

60. CTD hit bottom. Clean afterwards

62. File saved as event 61 –changed to 62.

63. Bottle #1 didn’t seal

73. 0.8m on deck

77. Our depth read 328m, Bridge depth read 488m, Altimeter did not kick in until 450m. Chart depth 500m. Stopped at 450m.

83. Bottle 2 was closed so rosette reversed, bottle opened and rosette redeployed before “real” cast.

87. Pressure 0.8db before.

95. 2 bottles fired at surface, no samples, no bottle file needed.

100. Pressure 0.7 and 0.8db on deck before

105. 2.6db after cast.

112. Altimeter at 5 for pressure 136 but depth 130

113. Waterfall close

115. Bottom with sounder was 436, based on altimeter ~463

124. Bottom hook on bottle 13 was broken and loose on recovery – replaced.

126. Pump was off briefly at the bottom  <1s.

128. Altimeter never kicked in.

130-131. Altimeter said 5m from bottom

134-135. Altimeter said 5m from bottom

141. Originally named 2015-54-0173 – renamed 141

144. False start - pump not on.

153. 0.8db on deck

156. at low tide

157. at low tide. Altimeter shows 7m from bottom.

158. at high tide, very windy, N wind

159. at high tide, very windy, N wind.

160. Station name wrong in data file. Fixed in converted file.

162. May have started to write over, but file still there.

163. deck pressure 0.7db before 0.7/0.8db after.

169. deck pressure 0.7m

175. Station name wrong in file header - fixed in converted file.

181. Station name wrong in file header – fixed in converted file.

183. 1 bottle tripped at bottom for sediment traps, 5m off bottom by altimeter

189. Ship was on sill with depths changing through cast. 20m off bottom by altimeter

189. Pumps did not come on until CTD at 39db.

191. Station name wrong in file header - fixed in converted file

195. Significant wire angle aft.

197. Moderate wire angle aft during downcast.

201. Bottle #9 did not close.

206. Forgot to record initial position in log

Leg 2:

219. Weights on – 45 deg wire angle

222. muddy bottom

237. logging kept on after CTD on deck

239. like other casts tonight bottoms are way off

248. fish trap station

259-266. rosette left in water at surface between casts

292. lots of jellyfish in water

309. lots of jellyfish in water

321/322. Noisy transmissivity

336. noisy transmissivity

349. CHL only at 10m – sample # not in sequence

CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 17Dec2014 | Factory |  |  |
| Conductivity | 2280 | 19Dec2014 | Factory |  |  |
| Secondary Temp. | 2663 | 15Jan2015 | Factory |  |  |
| Secondary Cond. | 2754 | 19Dec2014 | Factory |  |  |
| Transmissometer | 1396DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 17Jan2015 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| SeaPoint Fluor. | 3640 | n/a |  |  |  |
| Pressure Sensor | 0443 | 17Dec2014 | Factory |  |  |
| Altimeter | 1252 | n/a |  |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3363       Cruise ID#:	2015-54

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 28Dec13 | Factory |  |  |
| Conductivity | 2487 | 28Dec13 | Factory |  |  |