## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: CFMETR Nanoose

Agency: Canadian Forces

Project: Canadian Forces Maritime Experimental Test Range

Area: Nanoose

Date: 9 January 2015 – 20 November 2015

Platform: CFAV Stikine (and others)

Processed by: Germaine Gatien

Date of Processing: 25 January 2016 – 27 January 2016

Number of original Hex files: 78  	Number of casts processed: 71

INSTRUMENT SUMMARY

Six Sea Bird Model SBE 19 SEACAT CTDs were used – serial numbers #1449, 1450, 1511, 1512 and 1513. Each CTD had an altimeter attached.

SUMMARY OF QUALITY AND CONCERNS

Five different CTDs were used. The calibration information could not be confirmed. The CTDs are serviced annually.

To assist in analysis event numbers are helpful and the software in use will only allow 4 digits which suits most cruises. However a 7-character system is used for this project. A 4-character format was derived from those IDs based on the month and day of each cast. This means the same event # will appear if there is more than one cast on one day as frequently occurred.

The soak period for the CTD varied greatly from about 40s to 2 minutes. It is recommended that the SBE19+ be soaked for at least 2 minutes to improve salinity quality, though in the absence of large differences between air and water temperatures it is likely that 1 minute is sufficient.

There was no salinity calibration sampling. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors are expected to be as high as 0.05units in areas of high gradients. Away from high gradients errors in salinity are expected to be on the order of 0.005.

The times given in the SeaBird headers do not include a time zone, but the times strongly suggest that local time was used. All times were converted to UTC by adding 8 hours which is suitable if the original was in PST. If PDT was used for some casts, then those will have an error of 1 hour.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

78 hex files, 75 con configuration files and 74 hdr files were delivered.

The file names are non-standard and contain the letter N followed by the date in YYMMDD format and 2 digits to indicate the sequential casts on a single day. There are no event numbers as this is a time-series rather than a cruise.

- Conversion of Raw Data

- The raw data were converted by matching the configuration file to the input file name. This produced 74 CNV files.

- A spreadsheet “NanooseCtdStations2015.xls” was provided with positions for the casts.

- There are some differences between the spreadsheet and the raw files available, such as no position information being available for some files and some naming discrepancies.

- N15012901 -This cast was aborted and file n15012902 was run successfully a few minutes later. So n15012901 can be dropped and the position given for that file can be used for n15012902.

- N15021102 – No HDR or CON file – has same time as N15021101 but fewer records so not to be processed.

- N15021805 – Only CON file – no data

- N15033101 and N15033101_2 – 2nd file larger but has no con file. Converted with con file from the first group. Same upload times for both but 2nd has more data. The extra data suggests that the CTD may have touched the bottom and the depth seems to have been varying. So it is best to use the first and not process the second further.

- N15052801 and N150528_as_sent – The first is a little smaller but judging by the end of the file the CTD was very close to the bottom and there is evidence of shed wake, so having more data is probably not useful.

- N15061001, 1001a, 10002 – all uploaded at the same time. The first is slightly smaller but goes very close to the bottom, so it is likely the best choice.

- N15080701 – No positions provided and some files normally found in the RAW folder are missing. The cast is very shallow so it was likely aborted. Will not be processed.

- N150091501 – First attempt failed to convert because there was no con file. There is a CNV file among the raw files and that was created using a con file for this cast, so it must have been mislaid. Since we need different channels than in the old CNV file, a con file from cast N15091701 that used the same CTD was used to convert this file. The data look good.

- N15102302 – Didn’t convert properly. Time is the same as for N15102301 but there are fewer data.

- Positions are missing from the headers of the files but were provided in the spreadsheet. The spreadsheet was saved as “NanooseCtdStations2015.csv”. A column of event numbers was adding to the CSV file using MMDD part of the file name as event number.

The format of latitude and longitude in the spreadsheet needed to be changed to DD MM.MM N and DDD MM.MM W. Concatenate was used to do this. For example for column LOC:Latitude the entries are found using:  =CONCATENATE(INT(C2),"  ",TRUNC(((C2)-INT(C2))*60,2)," N ") where C2 is latitude. There needs to be a space before the latitude and around the N and W. (There is no need to add “! (deg min)” and the event number format does not need a leading zero where months only have 1 digit – those matters will be handled later in the MERGE routine.)

A column LOC:Station was added with every row having “Nanoose” entered.

- The soak period ranged from 40s to 1.8 minutes. A 2-minute soak is generally advised for the best salinity data, but in most cases the conductivity seems to have reached equilibrium after 1 minute.

- Temperature and conductivity generally look ok.

- The altimetry looks useful near the bottom when the CTD got close to the bottom.

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature.  Tests were run on a few casts to ensure this was the best choice.

The pressure was filtered with a time constant of 0.2s to increase the pressure resolution.

Plots of a few casts before and after this step show that reversals in pressure were removed without oversmoothing. A setting of 2s is often used for SBE19 CTDs but that looks too high for these data. The pressure is generally very smooth with just a few examples found that show signs of needing a filter.

- ALIGNCTD

An initial run with no alignment applied to the temperature channel produced poor results with instabilities in T-S plots. Tests were run on a few casts with different CTDs and the best results were found when temperature was advance by 0.73s relative to pressure. That is close to the setting used in 2012 (0.7s). There are small instabilities left in many casts. Alignment can be fine-tuned later using SHIFT if further adjustment is found useful.

- DERIVE

Program DERIVE was run to calculate salinity

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers.

As mentioned earlier, this program does not have event numbers, but such numbers can be helpful in finding particular data based on such things as maps or reports on extreme values. Using the MMDD part of the date to make up an event number is not distinct because there was often more than 1 cast on a single day. Including more than 4 digits does not work well in some reports and plots, which defeats the purpose of having event numbers.

So an event number was created for each cast based on the MMDD part of the date. This will result in casts from a single day all having the same event number, but that appears to be the best that can be done. (Attempts were made to use 5 or 6 digits in the event number but Header Check could not handle more than 5 digits and Track Plot could only handle 4.)

The program MERGE CSV FILES TO HEADERS was then used to add positions, event numbers and project name to the headers.

8. Checking Headers

Track Plot was used to ensure positions look normal – they were close but with the usual level of variability. The plot was added to the end of this report.

The times are said to be UTC but are clearly in local time, so ADD TIME CHANNEL was used to add 8 hours. It is not known if PST was used throughout the year, so the times may be off by one hour if and when PDT was used at sea.

Header Check was run and makes it clear that there were some large spike. There were 5 different CTDs in use (S/N 1449, 1450, 1511, 1512, 1513).

Surface check was run and the average surface pressure was 0.9db.There were 4 slightly negative values. For CTD #1450 for one cast when pressure was -0.08db the salinity looks like the CTD was right at the surface, but at another when it was -0.62db it appeared to be a bit deeper, but surface salinity is highly variable and the two casts were 20 days apart.

For CTD #1511 and 1449 the salinity at -0.4db and -0.3db respectively show that the instrument was clearly right at the surface or out of the water. There were no negative pressures for the other 2 CTDs.

These pressure sensors are considered ±1db and often show different surface pressure from one cast to another, so these results do not suggest a need to recalibrate.

- The altimetry header information from each file was exported to a spreadsheet. In about half of the casts the CTD did not get within 15m of the bottom, so there is no header entry. About half the files with an altimetry header were checked and the altimetry header values look appropriate.

9. SHIFT

These data were aligned earlier in the process, but tests were run to see if a further shift might reduce unstable features and applying a setting of -1.5 records to the conductivity did do that.

SHIFT was run on all casts with that setting.

10. CLIP

CLIP was used to remove the first 50 scans; by eliminating the early soak period editing will be simpler.

11. DELETE

Then DELETE was run on the clipped files using the following parameters:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

There were warnings for events #1023 and 1113. The CLIP files were examined to see if some editing is needed before running DELETE.

For event #1023 there was a jump of 2.5db. There was clearly some sort of problem with pressure, but no reversal. DELETE handled that appropriately.

For event #1113 the data at the bottom of the cast look bad and DELETE did not work well. The CLIP file for that cast was edited by removing records at the end and then DELETE was rerun. There were no warnings on that run of DELETE.

12. CTDEDIT

The DEL files were copied to EDT.

DEL files were opened in CTDEDIT. Most of the casts required light editing, mostly to smooth cases where unstable features appeared to be due to minor misalignments of T and C in the presence of temperature gradients. Some files had groups of 3 records that appear corrupted – temperature increases abruptly and then snaps back to values close to values from earlier. Such records were removed where they seemed obviously bad. There were also a few cases of shed wake corruption near the bottom and poor data at the surface. 7 casts required no editing.

After the first run the data were examined in T-S plots.

- For file n15012801 there was an unstable feature between 25 and 45db that started and ended with odd temperature (small spikes) and salinity shifted to high values not seen in any other casts, including one the next day at a close position.  A second attempt at editing found no simple adjustment that could fix this, so all records in that range were removed. The cast was one with many small spikes in temperature.

- File n15080401 has some unstable features, but they are small and possibly real. No justifiable editing will make them go away, so they were not edited further.

The edited files were copied to EDT.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure      	Averaging interval = 1.000

Minimum bin value = .000	Average value will be used.

Interpolated values are NOT used for empty bins

Profile plots were examined on-screen and no problems were found.

Page plots were examined and look ok with only a few very slightly unstable features.

14. Intercomparisons

Historic ranges – The only excursions from the local climatology were near the surface where temperature was often a little high down to between 20 and 50m between May and September and there was low salinity at the surface for one cast in June. The 3-standard deviation climatology is both old and too severe for near-surface data nears shore. There is no indication of calibration problems.

Repeat casts – Casts were all close in space, and on a few occasions there were 2 or more casts on one day but not necessarily very close together. Two casts that were very close in space and about 3 hours apart were examined. At about 300db temperatures were within 0.02Cº of each other (and generally much closer than that) along sigma-T lines and salinity values were within 0.005psu. The CTDs used were different. There was considerable temperature variability in each of these casts even near the bottom, so the agreement is considered good.

15. CALIBRATION

No recalibration was applied as there was no salinity calibration sampling.

16. REMOVE and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity, Descent_Rate Altimetry and Flag.

The HEADEDIT routine was used to fix formats and to add the following comment to the headers:

SeaCat salinity is prone to error due to a mismatch of conductivity and

temperature response times, especially when the descent rate is non-uniform.

Such errors are likely to be larger than calibration errors. SeaCat salinity

errors are expected to be as high as 0.05 units in areas of high gradients,

but on the order of 0.005 where gradients are low and descent rate steady.

There were no salinity calibration data available for this cruise.

The CTDs are serviced annually.

The standards check routine was run and no problems were found.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing and header check were produced. No errors were found.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	Nanoose Test Range

Dates: 	Start:	9 January 20155	    End:	20 November 2015

Location: Nanoose Test Range 	Vessel:    Stikine (and others)

NOTE: Five SEACATs were used for this cruise.

The calibration information used in conversion can be found in the headers of the CTD files.