## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27 March 2025 | Updated channel names & formats in TOB files. G.G. |

## PROCESSING NOTES

Cruise: 2016-02

Agency: National Research Council of Canada & DFO-Ocean Sciences Division

Location: Douglas Channel

Project: World Class – Douglas Channel

Party Chief: Greer C.

Platform: W.E. Ricker

Date: 9 March 2016 – 14 March 2016

Processed by: Germaine Gatien

Date of Processing: 15 June 2016 – 21 June 2016

Number of original HEX files:	26	Number of CTD files: 26

Number of bottle files: 	25		Number of bottle casts processed: 23

Number of original TSG files:  	4	Number of processed TSG files:	 4

# INSTRUMENT SUMMARY

SeaBird Model SBE 25+ CTD (#1091) was used for this cruise. It was mounted in a mini-rosette and attached were a Wetlabs CSTAR transmissometer (#983DR), a SBE 43 DO sensor (#1483), an umpumped WetLabs ECO Fluorometer (#2216) and an altimeter.

A thermosalinograph (Seacat 21 S/N 2487) was used with sampling interval 30s.

Seasave version V7 23.2 was used for acquisition.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

The oxygen kit was Arctic SIO Kit B.

An IOS min-rosette with 6 5L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log and analysis logs were in good order. It was especially appreciated that there was information included about the new mini-rosette and the TSG. The rosette sheets had one error in event number and a few discrepancies between the sampling indicated and what was delivered to analysts, which may be partly due to samples being taken for 2 groups to be analyzed in different labs. Otherwise the rosette logs were in good order; they included notes about whether bottle files were required when there was no sampling by DFO scientists.

There were only 3 salinity calibration samples. Two indicated that the CTD salinity was lower than bottles by less than 0.001 while the third was rejected as an extreme outlier. There was a 6-week delay in analysis so bottle values may have been a little too high which would imply that the CTD could be reading slightly high. However, there was likely a small compensating error due to incomplete flushing of Niskin bottles. The temperature and conductivity were recently calibrated at the factory, so little drift in salinity is expected.

The dissolved oxygen comparison between CTD data and bottle samples was not reliable enough to make an estimate of the accuracy of the Oxygen:Dissolved:SBE data.  There were only 20 bottles of which 8 were flagged due to bubbles or poor replication. The casts were taken in a channel where poor flushing of bottles is expected; the 4 bottom bottles would be affected in the opposite way to the others. The vertical DO gradients were very complex with significant reversals. The CTD DO data in the bottle files were recalibrated using a linear fit based on a 2014 cruise that used the same sensor.

The pressure calibration used at sea was not the most recent. It was accidentally chosen in converting the files. Fortunately, the resulting error was found to be small (~0.3db) and the near-surface values actually look better than in data converted with the later calibration. Strain gauge pressure sensors initial accuracy is quoted as 1db so this result is not a great surprise.

The data from the thermosalinograph look good, but there was insufficient information to enable an assessment of its accuracy. There are large gaps in the record.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets for dissolved oxygen, extracted chlorophyll and salinity.

The BE time in the log corresponds to the time at the end of the soak period.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The pressure sensor was new with no history. This was the first known use of the temperature and conductivity sensors since they were last calibrated. The history of the dissolved oxygen sensor was checked.

##### Conversion of Full Files from Raw Data

It was intended that all hex files be converted using 2016-02-ctd.xmlcon to create CNV files, but an error was made in conversion and the pressure calibration information was incorrect in the configuration files used. A study described in section 17 determined that the error had a small effect and the results actually looked better than when the correct values were used. The pressure sensor is a strain gauge sensor which is subject to lower accuracy.

A few casts were examined and all expected channels are present and look reasonable.

The descent rates were steady and high.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2016-02-ctd.xmlcon.

It was noted in the log that all bottles for cast #8 were fired at the surface. Because they were fired quickly, few records from the CTD sensors are captured for bottles 1-5 in the rosette files. This could affect the reliability of comparisons; bottle 6 will be most reliable. There are other cases where several bottles are fired at a single depth, so in general the best comparisons will come from the last bottle in a set in each case. This is not a significant issue except that the standard deviations in the COMPARE files are useful in deciding which data are most suitable for comparisons. With only a few points variability may be misjudged. Most OSD sampling was done from the last bottle fired in each set.

Note that bottles were fired during events #1 and 67 but only as tests or to get water, so no sample numbers were assigned and no bottle files are needed. They were not processed further.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files to check for outliers and none were found.

A preliminary header check was run and no problems were found.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

One bottle was removed from the list where there was no sampling and no sample number.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Those files were then bin-averaged and called SAMAVG.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2016-02-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2016-02chl*.xls which included comments, flags and a precision study. The spreadsheet was simplified by removing some columns and the file was saved as 2016-02chl.csv, which was then converted to individual CHL files. The event number was incorrect for CHL sample #6; it should be event #2 – this was wrong on the rosette log.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2016-02nuts.xls. This includes a precision study. The file was simplified, reordered on sample numbers and saved as 2016-02-nuts.csv. The csv file was converted to individual NUT files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2016-02oxy.xls which includes flags and comments. There was no precision study because there were few duplicate samples. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2016-02oxy.csv which was converted to individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2016-02SAL.xls. Analysis was done within 42-44 days of collection. The file was simplified and saved as 2016-02sal.csv which was then converted to individual SAL files. There were only 3 samples and no duplicates.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. A few discrepancies were found:

- CHL sample given as sample #24 on the rosette sheet but #23 in the QF file. The sample label had been changed to #23, so the QF file appears to be correct. Both were surface samples.

- CHL sample #75 is found in the QF file but not on the rosette log on which there are markings and a comment to indicate the sample did not get taken. However, there is sample label clearly marked Doug16/ bottle 6/ sample #75. The sample is presumed to be correctly identified.

- SAL sample #76 from event 45 Niskin 1 was not found in the QF file or on the salinity analysis log sheet. There were obviously some changes made to the log sheet, so it assumed that the sample was never gathered or was lost.

##### Compare

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There is a lot of scatter and trend-lines have very large offsets. Setting the offset to 0 produces an unconvincing fit. Removing points with bubbles and 1 other outlier produce a fit with a slope close to that found in 2014 but a much larger offset.

CTD DO Corrected = CTD DO * 1.0301 + 0.1959

Dropping the only remaining bottom bottle produces the fit:

CTD DO Corrected = CTD DO * 1.0407 + 0.1351

A similar fit is produced by excluding bottom bottles but using everything else. This fit makes it clear that the bottom bottles are quite different. Bottles fired near the bottom will have different flushing characteristics from other bottles and/or may be affected by bottom currents or sediment.

CTD DO Corrected = CTD DO * 1.0428 + 0.1213

The slope of these fits is close to the 2014 value (1.04) but the offset is much higher. Errors due to calibration drift are generally found to be roughly linear with a small offset. Explanations for larger offsets include:

- Bubbles were found in many samples. Those were excluded from the comparison but perhaps small ones were present but not detected. There were problems experienced in drawing samples from the new mini-rosette on the Ricker.

- Mismatches due to poor flushing of Niskin bottles. This error would be greatest where DO gradients are highest. But they would lead to the correction to the CTD being smaller, rather than larger as seen here. We expect poor flushing in narrow channels where CTD ascent rates tend to be very steady, indicating quiet conditions so bottles won’t get shaken. So there likely is some error due to poor flushing, but something else must be happening that produces larger errors in the opposite direction.

For the 4 casts with DO samples there are complex gradients with mid-depth reversals that are sometimes small, sometimes large, and there are high gradients near the surface and at mid-depths. Given few bottles that have not been flagged, the comparison must be considered of very limited value.

For recalibration of the downcast data we want to correct for calibration drift. Neither poor flushing nor bubbles in samples are factors for the downcast. But for the bottle files we want a good reflection of what is actually in the bottles. If we had a measure of what errors are caused by poor flushing we could attempt to recalibrate to fix that, but we cannot separate that error from a larger error that may be due to bubbles. The fit that excludes bottom bottles and samples with bubbles has a slope very close to the one found during 2014-23.

CTD DO Corrected = CTD DO * 1.04     + 0.0037 (2014-23)

CTD DO Corrected = CTD DO * 1.0407 + 0.1351 (2016-02)

Since we suspect that bubbles are the major problem, then just dropping the offset may be justified as bubbles are likely to be present throughout the DO range. We could consider doing a fit that forces zero offset; that would lead to a slope of 1.05 to 1.065 depending on how outliers are identified. The quality of the bottle data renders this approach less reliable than using an old fit.

For more details see document 2016-02-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No outliers were found.

Salinity

Compare was run with pressure as reference channel.

There were only 3 samples and they came from the bottom of casts. One is clearly an outlier. The two other samples are well within 0.001 of the CTD salinity. There was a delay of 6 weeks before analysis so we expect a little evaporation and/or adsorption of sample so that bottle values might be slightly high. From the dissolved oxygen comparison we might expect that flushing is incomplete, but these are all bottom bottles so the effect would be to have salinity from higher in the water column, hence lower salinity, but not much lower since local salinity gradients are very low for 2 of the samples. These two effects are small and offsetting. The outlier sample is higher than the CTD salinity by 0.64. This one came from cast #27 which had a high salinity gradient at the bottom, so if anything we would expect the sample to be much lower than the CTD due to incomplete flushing. The salinity sample value was 32.8862 while the maximum salinity measured by the CTD was 32.2499. One possibility is that the sample is mislabelled and is really the missing sample from event #45. In that case it would higher than the CTD by ~0.004; that is still out of line with the other 2, so we cannot assume this is the missing sample. The analyst decided to flag it 5 with a pad value.

For full details for the COMPARE run see file 2016-02-sal-comp1.xls.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a WetLabs ECO sensor. All samples came from the surface and CHL was in the range of roughly 0.5ug/L to 1.3ug/L. The ratio of fluorescence to CHL ranges from 2.5 to 1.3 with the highest CHL having the lowest ratio. This is typical of these fluorometers.

For full details of the comparison see file 2016-02-fl-chl-comp1.xlsx.

2016-02-ctd-recal1.ccf was prepared to apply the correction: CTD DO Corrected = CTD DO * 1.04

The SAM files were corrected using 2016-02-ctd-recal1.ccf and COMPARE was rerun on the dissolved oxygen and showed those corrections were applied correctly.

For full details see file 2016-02-dox-comp2.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### FILTER

WFILTER was run using a cosine filter, size 5, on the pressure, temperature and conductivity channels. It had little effect on pressure and temperature for the test casts, but conductivity reversals were removed.

##### ALIGN DO

Tests were run on a few casts to see what alignment made the offset between the upcast and downcast DO traces resemble that for the temperature traces. The vertical offset between up and downcast temperature was approximately 4m and for DO, 18m, giving a difference of 12m. Since both downcast and upcast contribute, that is 6m each and the descent/ascent rate is about 1m/s, so a 6s advance was taken as a first guess. Tests confirmed that was the best choice.

ALIGNCTD was run on all casts using +6s.

##### CELLTM

The default setting for the SBE25 is of (α = 0.04, β=8). A few casts were tested to see if this worked well and there were only small differences between results with a wide variety of settings, all improving the correspondence between downcast and upcast in T-S space. The upcast data are often too noisy for a good judgment to be made, so the default value suggested by the manufacturer was chosen.

CELLTM was run using (α = 0.04, β=8) for the primary conductivity.

##### DERIVE

Program DERIVE was run on all casts to calculate primary salinity and dissolved oxygen concentration.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. The only discrepancy was the longitude for event #18 at station DEV1A. The NMEA entry is 128 51.20W. The log entry was originally 128º but was erased and changed to 129, probably because the casts before and after it had 129º and it just looked wrong. If it were actually 129 it would be in Douglas Channel rather than Devastation Channel.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was -0.068db. All the salinity values are near 0, so the CTD was either out of the water or right at the surface. The one deck reading available was from event #46 and ranged from 0 to 0.02db. This pressure calibration looks fine.

A header check was run. There were negative values for many variables but these appear to be right at the surface. Fluorescence is very noisy during the soak period varying from near-zero to values as high as 35ug//L. The CTD is very close to the surface and may have been moving in and out of the water. The transmissivity is also spiky.

The altimeter and water depth readings from the headers of the CLN and MRGCLN2 files were exported to a spreadsheet. Plots were made of altimetry near the bottom for some casts and the headers look appropriate for all files except for casts #8, 10, 11, 24, 61, 63, and 65 which were shallow casts in deep water where altimetry spikes were misinterpreted as being near the bottom. Those altimeter headers were removed from the CLN and SAM files (though the altimetry header was already missing for a few of the SAM files).

Water depth entries were compared with the log book entries. Most were within 2m of the log entry. For event #37 the depth was changed to match the log. There was no depth entered for event #67. The log has 504m while 498m looks more appropriate at the bottom based on the maximum pressure and altimetry. The 504m reading was probably taken at the beginning or end of the cast and there may have been some shoaling, so the entry was changed to 500, that being a rough estimate.  That was corrected in the CLN file; there was no rosette file for that cast.

After the corrections, the SAM files were bin-averaged again and the MRG files were recreated, cleaned and recalibrated.

##### Shift

Fluorescence

The fluorometer was not pumped, so a shift in alignment is expected to be small or unnecessary. Profile plots of temperature and fluorescence were examined and confirm that the alignment is ok.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it looks as good as we can expect to achieve. No further alignment was applied to the DO concentration channel,

Conductivity

Tests were run on 4 casts using a variety of shifts. The best results varied from feature to feature, with choices of +0.3, 0.4 or +0.5 records all working quite well. SHIFT was run on all casts using +0.4.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

This is the first known use of the pressure sensor. There is no record of use for the conductivity sensor since it was last calibrated in August 2014. The dissolved oxygen sensor had been used during 2014-23 when the correction applied was slope 1.04, offset 0.0037. It either malfunctioned or was not really mounted during 2015-26.

Historic ranges – No local climatology was available.

Repeat Casts – There are some repeat casts but the variability in this region means they are unsuitable for testing the repeatability of the CTD.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

There are many unstable features right at the surface; most do not look like real variability.

CTDEDIT was used to remove noisy near-surface records and records corrupted by shed wakes which were mostly at the bottom of casts. Salinity was cleaned lightly where it appeared that small alignment variations led to unstable salinity features. All files required light editing.

For event #2 some bad fluorescence data was removed around 4m. There are other cases of high and low values in the top 1-2m of casts but for those it is not completely obvious that they are wrong, so they were left in place. This channel is not normally edited and the data should be considered nominal.

##### Pressure study

At this point in the processing it was discovered that the wrong pressure calibration had been used in converting the CNV files. There were calibrations in both February 2015 and July 2015 and the earlier one was used at sea. This was noted in processing and a new calibration file was created, but in conversion the box that matches HEX files to XMLCON files was checked by accident.

To see whether the resulting error was significant, a deep cast was selected and pressures compared using the 2 configurations to see if a simple recalibration would be a satisfactory way to correct the pressure. A linear fit was found that worked well with errors <0.02db. A simple offset correction would leave errors of <0.03db. The differences between the two fits were between 0.24 and 0.29db. The initial accuracy of the strain gauge pressure sensor is ~1db.

When the pressures in the files are checked it looks like the original parameters produced better results as pressure looks too low with the later calibration. The surface check described in section 12 showed good results. If a recalibration is applied the results will look worse. This may be due to drift or just that these sensors are not very accurate. So, while an error was made and it is easy to fix, the effect is small and the earlier calibration may be better for these data. A quick check was made cruise 2016-03 which used the same equipment and immediately followed this cruise. It also showed better near-surface results whwen the older calibration was used. No correction will be made.

##### Initial Recalibration

The EDT files were recalibrated using file 2016-02-ctd-recal1.ccf which applied the dissolved oxygen correction described in section 5 and to correct the pressure:

CTD DO Corrected = CTD DO * 1.04

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is usually made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed (the same ones as in 2016-02-dox-comp1.xlsx), the CTD DO was higher than the bottle DO by a median of ~0.01mL/L. However, the standard deviation was 0.17mL/L.

For these data we expect a poorer fit than usual because there appears to be a problem with the bottle data, possibly due to bubbles. We would expect the CTD DO to be lower than the bottles due to calibration drift, though poor flushing could lead it to look higher especially in areas of high vertical DO gradient. Variability is high in the comparison. As well as the possibility of random errors due to bubbles, examination of the DO profiles shows that there are several large reversals in DO gradients for 2 casts and for the other two there are low gradients with small reversals and fairly high gradient between 150m and 225m. These complexities plus the many bottles that were flagged due to bubbles and poor replicates show that the comparison is of little value. The largest difference was 0.35mL/L while below 2m the CTD is within 0.2mL/L of bottles excluding the bottom bottles and flagged samples. We have no reason to believe that the differences are due to poor downcast DO.  We cannot make an estimate of the reliability of the downcast CTD DO.

See 2016-02-dox-comp3.xlsx for the comparison details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was had little effect on spikes and over-smoothed other data. This step was skipped.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots look fine.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Conductivity:Primary, Oxygen:Voltage:SBE, Altimeter, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

---------------------------------------------------------------------------------

This was a joint cruise between the National Research Council of Canada and the

Ocean Sciences Division of the Department of Fisheries and oceans, Sidney, B.C.

The chief scientist was from NRCC.

The OSD contact is Cynthia Wright.

----------------------------------------------------------------------------------

Data Processing Notes:

----------------------

Fluorescence and transmissivity data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Salinity calibration sampling was limited to 3 samples one of which was

rejected. Delayed analysis and poor flushing of Niskin bottles also limit

the reliability of the comparison, but the 2 samples do not suggest any

significant calibration drift.

Dissolved oxygen calibration sampling was limited to 20 samples

and many were flagged by the analyst due to bubbles or poor

replication. The DO profiles were very complex. No estimate

is made for the accuracy of the  Oxygen:Dissolved:SBE channel.

For details on the processing see the report: 2016-02_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

A cross-reference list was produced.

The sensor history was updated.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The surface saturation values were between 82% and 96% with the highest values found closer to open waters. The lowest were associated with cases where the DO gradient reversed direction right at the surface. A test was run to see if including the offset in the recalibration of the CTD DO channel would make a significant difference and it did not. When the same sensor was used on cruise 2016-03 in the Strait of Georgia the saturation rates were higher, mostly between 95% and 105%. In those profiles the DO gradient did not reverse itself at the surface often. So the low values from 2016-02 cannot be presumed to indicate the DO values are too low and even increasing the recalibration factor from 1.04 to 1.06 makes only a small difference. The evidence is too weak to adjust the recaliabration.

##### Final Bottle Files

The MRGCOR2 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Conductivity:Primary, Descent Rate, Altimeter, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check and a header check were run on all files and no errors were found.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

A cross-reference listing was produced for the CHE files.

Data were exported from the CHE files to file 2016-02-bottles-final.xlsx. The entries were compared with the rosette log sheets to ensure no samples had been missed and no problems were found.

##### Thermosalinograph Data

There were no loop samples, no intake thermistor and no flow meter.

The TSG was found to have stopped on several occasions so there will be gaps in the record.

a.) Checking calibrations

The configuration files for the 4 files are identical. One file was renamed as 2016-02-tsg.xmlcon. No errors were found in the calibration parameters.

b.) Conversion of Files

The files were converted to CNV files using configuration file 2016-02-tsg.con. They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots were produced. The data at the beginning of file 2016-02-0001 looks as though the flow was turned off, so some editing will be required. Otherwise the traces look good.

The track plot looks fine though there are gaps. The plot was added to the end of this report.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 3db and exported to a spreadsheet which was saved as 2016-02-ctd-tsg-comp.xls.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found in spreadsheet 2016-02-ctd-tsg-comp.xlsx. The differences in latitude and longitude were all≤0.0003° except for the one case where the TSG was started after the start of the CTD cast. The median differences were <0.0001°. This shows both the times and positions are reliable for both CTD and TSG systems.

d.) Comparison of T and S from TSG and CTD data

The spreadsheet described in the previous section was then examined to find the differences between the salinity and temperature channels for the CTD and the TSG.

There were 10 casts which overlapped with TSG files and had data at 3db, but of those 1 was started during the CTD cast so did not match time well and 1 did not have flow on. Four others had very high standard deviation in the TSG temperature and salinity so that we can’t expect a good match.

For the 4 cases that remain, the TSG salinity was lower than the CTD salinity in each case, which suggests that the TSG was sampling a little higher in the water column. An estimate was made of how much higher by finding the TSG salinity in the CTD profile and depths found were 2.6, 0.1, 0.1 and 2.7db. Of course, there could be calibration drift in the CTD temperature and conductivity but those effects are usually small. Since the temperature is increasing with depth for these casts we expect that the TSG intake temperature should be lower than the CTD temperature based on our observations of salinity. Unfortunately, we only have lab temperature which will have gained some heat in the loop. So the fact that in 3 cases the TSG lab temperature is higher than the CTD and for 1 it is lower is likely due to competing factors. The outlier in the group is from the cast with the sharpest temperature gradient, so it is the least reliable in the comparison.

To try to estimate how much heating there is in the loop the level at which the TSG salinity occurs in the CTD salinity profile was found and the temperature from that level recorded. Those values are lower than the CTD temperature by between 0.002 and 0.044Cº, with an average of 0.02Cº. The 4 estimates of heating in the loop are then -0.07Cº, +0.07Cº, +0.05Cº, 0.07Cº with a median value of +0.06Cº. We have little history from the Ricker loop, but values from the Tully tend to be in the 0.15-0.25Cº and the Ricker loop is shorter, so this value does not look unreasonable. It is not enough information to allow a calculation of a proxy for the intake temperature.

e.) Calibration History

This was the first known use of this TSG since it was last recalibrated in December 2015.

Conclusions

1. The TSG clock worked well.

2. The temperature and salinity traces look fine and values compare reasonably well with 4 CTD casts that overlap with the TSG record and have reasonably low standard deviations in the TSG data.

3. Only the first cast needs editing to remove a section that shows clear signs of flow being turned off.

4. There is insufficient information to recalibrate the TSG temperature and salinity.

f.) Editing

The ATC files were copied to *.EDT.

All files were opened in CTDEDIT but only file 1 contained data that looked bad. Editing was applied to file 1 to remove temperature and salinity values for the first 236 records where flow was either turned off or just getting established.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Conductivity:Primary, Flag and Position:New channels.

HEADER EDIT was used to change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats and to add the following comment.

TSG Processing

*****************************************************************************

There were no loop samples available and no flow rate meter.

There were only 8 CTD casts that overlapped with the TSG record while the flow

was on, and 4 of those were in areas with large variability in the TSG records.

The 4 CTD casts that were suitable for comparison were in reasonable agreement

with the TSG data, but there are insufficient data to enable an assessment of

accuracy or to allow an estimate of the intake temperature.

Temperature:Lab - These are uncorrected TSG temperature data from the lab.

*****************************************************************************

Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

# Notes from Log Book

General

Deployment Method: Surface soak of 2min. Time in log corresponds to beginning of descent, after the soak.

Computer clock 6 minutes behind NMEA clock.

TSG stopped occasionally.

Particulars

1. All bottles fired but no sampling, no sample numbers assigned.

2. Descent rate low in top 20m. Surface trips 1.0-1.5db.

4. Sewage release 15 minutes before cast.

8. All bottles fired at surface very quickly.

11. Possible waste discharge between 10 and 11. Soak time reduced to 90s since conductivity cell should be bit wet.

12. 3 bottles added at last minute so sample #s 934, 935, 936 were used.

27. Trans and DO rather noisy.

35 & 37. Odd oxygen but somewhat matching temp.

45. Sewage dump before cast so waited for current to flush.

46. Deck pressure 0 to 0.02db.

61. Air temp 2.5ºC and snowing.

65. Some wire angle due to current so CTD slowly gets deeper during bottle stops.

67. 6 bottles fired at bottom just to get water, no sample numbers assigned.

# CRUISE SUMMARY-CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 1091 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4700 | 20Aug2014 | Factory |  |  |
| Conductivity | 3321 | 20Aug2014 | Factory |  |  |
| Transmissometer | 983DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 1483 | 17Jan2014 | Factory |  |  |
| WetLabs ECO Fluor | 2216 | n/a |  |  |  |
| Altimeter | ? |  |  |  |  |
| Pressure Sensor | 1091 | 12Feb2015 | Factory |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2016-02

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 7Nov15 | Factory |  |  |
| Conductivity | 2487 | 7Nov15 | Factory |  |  |

Note that there was a test cast in Saanich Inlet that is not displayed on the maps on this page.