## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2016-03

Agency: IOS, Ocean Sciences Division &

PBS, Marine Ecosystems and Aquaculture Division, Nanaimo, B.C.

Location: Strait of Georgia

Project: Strait of Georgia Zooplankton

Party Chief: Gauthier S.

Platform: W.E. Ricker

Date: 14 March 2016 – 20 March 2016

Processed by: Germaine Gatien

Date of Processing: 21 June 2016 – 27 June 2016

Number of original HEX files:	34	Number of CTD files: 34

Number of bottle files: 	11		Number of bottle casts processed: 11

# INSTRUMENT SUMMARY

SeaBird Model SBE 25+ CTD (#1091) was used for this cruise. It was mounted in a mini-rosette and attached were a Wetlabs CSTAR transmissometer (#983DR), a SBE 43 DO sensor (#1483), an umpumped WetLabs ECO Fluorometer (#2216) and an altimeter.

Seasave version V7 23.2 was used for acquisition.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

An IOS min-rosette with 6 5L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

There was no information in the front of the log book. While participants may think it is obviously the same equipment as on the previous cruise, it is important to keep a complete record in the log book including CTD sensors and science crew. Some of the information was obtained from the post-cruise report. The log contained useful notes on problems and the CTD deployment method, which varied.

The rosette log was in good order.

There were only 6 salinity samples but they were well placed near, but not at, the bottom and mostly at levels where the vertical salinity gradients were fairly low. The CTD salinity was lower than bottles by an average of 0.0006 and a median value of 0.0013. The analysis of bottles was done about 5 weeks after collection so errors due to evaporation and adsorption are likely small. There may also be some error due to incomplete flushing but the low gradients should make those errors small. Both these factors could lead to bottle values being a little too high, so the CTD salinity could be reading slightly high.

There was no dissolved oxygen sampling and that done during the previous cruise was not trusted due to sampling problems. Recalibration of channel Oxygen:Dissolved:SBE was based on the results of a 2014 cruise that used the same sensor and included DO calibration sampling.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets for extracted chlorophyll and salinity.

Nutrients, extracted chlorophyll and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The pressure sensor, temperature and conductivity sensors had been used for 1 other cruise since they were last calibrated. The dissolved oxygen sensor was used for 3 other cruises though the only cruise with good calibration sampling was in 2014.

##### Conversion of Full Files from Raw Data

The configuration file used at sea contains pressure sensor calibrations that are not the most recent for this sensor. However, when the same sensors were used for 2016-02 the older calibration was used by accident. The differences were slight, but overall it looked like the older calibration produced better results. So test conversions were done to see which produced the best results. For this cruise there was a soak at 10m for just the first 3 casts. During cast 46 there is a transition for data that looks like it is in air to water values at pressure -0.2db using the July calibration and +0.1db using the February values. So the difference is slight but the February values once again look more reliable.

After conversion the cruise number was corrected in the headers of events #1-5.

All hex files were converted using the xmlcon file used in acquisition to create CNV files. These are identical to file 2016-03-ctd.xmlcon that was used for the bottle files.

A few casts were examined and all expected channels are present and look reasonable.

The descent rates were high and steady.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2016-03-ctd.xmlcon. The cruise number was corrected in the headers of events #1-5.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files to check for outliers and none were found.

A preliminary header check was run and no problems were found.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Those files were then bin-averaged and called SAMAVG.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2016-03-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2016-03chl*.xls which included comments, flags and a precision study. The spreadsheet was simplified by removing some columns and the file was saved as 2016-03chl.csv. The event numbers were incorrect – they were actually station names, so those were corrected. The file was then converted to individual CHL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2016-03nuts.xls. This includes a precision study. The file was simplified, reordered on sample numbers and saved as 2016-03-nuts.csv. The csv file was converted to individual NUT files.

SALINITY

Salinity analysis was obtained in 2016-03SAL.xls. Analysis was done within 31-36 days of collection. The file was simplified and saved as 2016-03sal.csv which was then converted to individual SAL files. There were only 6 samples, all duplicates and the Sp value was 0.0031.

The SAL, CHL and NUT files were merged with CST files in 3 steps.

After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. A problem was found with the samples from station 39 because the event number ws given as 3 on the rosette log sheet, but was actually from event 5. That was corrected and the merges rerun. No further errors were found.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There were only 6 samples but they were well placed near, but not at, the bottom and mostly at levels where the vertical salinity gradients were fairly low.

There was one outlier, sample #11, which had been flagged due to the Niskin failing the integrity test. The CTD salinity was lower than the duplicate bottles by 0.006 and 0.009 for that bottle.

For the other 5 bottles the CTD was within 0.0031 of bottles. The CTD was lower than bottles by an average of 0.0006 and a median value of 0.0013.

The analysis of bottles was done about 5 weeks after collection so errors due to evaporation and adsorption are likely small, but not 0.

Niskin bottles may not flush well in quiet waters which would also raise the salinity of bottle samples.  However, none of the salinity sampling was done in very narrow channels, and one of the 5 casts was taken in fairly open water, where ship motion should help stir up the contents of the Niskin bottles. Nonetheless, the descent rate of the CTD is fairly quiet, so there is likely some inefficiency in flushing. The local gradients were fairly low which limits the errors associated with flushing problems. There is a likely a small effect which would lead to higher salinity in the bottles.

The combined small errors could mean that the CTD salinity is reading a little high, but drift appears to be slight.

For full details for the COMPARE run see file 2016-03-sal-comp1.xls.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a WetLabs ECO sensor. All samples came from the top 20m and CHL was in the range of roughly 0.56ug/L to 11.6ug/L.

The ratio of fluorescence to CHL ranges from 0.5 to 1.5 with the highest CHL having the lowest ratio. This is typical of these fluorometers.

For full details of the comparison see file 2016-03-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### FILTER

WFILTER was run using a cosine filter, size 5, on the pressure, temperature and conductivity channels. The traces are quite smooth but there were places where pressure, temperature and conductivity reversals were removed.

##### ALIGN DO

Tests were run on a few casts to see what alignment made the offset between the upcast and downcast DO traces resemble that for the temperature traces. The vertical offset between up and downcast temperature was approximately 4m and for DO, 18m, giving a difference of 12m. Since both downcast and upcast contribute, that is 6m each and the descent/ascent rate is about 1m/s, so a 6s advance was taken as a first guess. Tests confirmed that was the best choice and it was also found for 2016-02.

ALIGNCTD was run on all casts using +6s.

##### CELLTM

The default setting for the SBE25 is of (α = 0.04, β=8). For 2016-02 a few casts were tested to see if this worked well and there were only small differences between results with a wide variety of settings, all improving the correspondence between downcast and upcast in T-S space. The upcast data are often too noisy for a good judgment to be made, so the default value suggested by the manufacturer was chosen. The same choice was made for this cruise.

CELLTM was run using (α = 0.04, β=8) for the primary conductivity.

##### DERIVE

Program DERIVE was run on all casts to calculate primary salinity and dissolved oxygen concentration.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. No discrepancies were found.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was +0.17db. All the salinity values at the surface are <0.5psu with most much lower than that. So the CTD was either out of the water or right at the surface. The pressure calibration looks fine.

A header check was run. There were extreme values for some variables but these appear to be right at the surface. The CTD was very close to the surface and may have been moving in and out of the water. The transmissivity is also spiky.

The altimeter and water depth readings from the headers of the CLN and MRGCLN2 files were exported to a spreadsheet. Plots were made of altimetry near the bottom for all casts and the header entries look appropriate for all CTD files. However, for the bottle files of cast #42 the only sampling was at the surface, so the reading of 1.4m for altimetry is clearly wrong so it was removed from the header of bottle files from the SAM and SAMAVG files.

Water depth entries were compared with the log book entries and there were only a few significant differences. Where they differed from the log by more than 2m, checks were made by using 99% of the maximum pressure to get a rough idea of the maximum depth sampled. The difference between that and the header depth entry was compared with the altimetry reading. Where the header entry looks closer to the calculated estimate, no change was made. Five casts were investigated and where changes looked appropriate they were made in the SAM and CLN files:

- Cast #6 - the header is 5m higher than the log, but there is a note saying that 5m should be added to the water depth due to “trans.boom offset”.  No change was made to the header entry.

- Cast #10 – 212m in log and 202 in header – not clear which is better but probably the log entry, so header was changed. There was no bottle file for this one.

- Cast #22 the log entry (429) makes much more sense, so that was entered in the header.

- Cast #24 there is a note saying there was shoaling during the cast, so no change was made.

- Cast #46 – the log has 351, header 347. The log entry looks better, so the header was changed.

After the corrections, the SAM files were bin-averaged again and the MRG files were recreated and cleaned.

Plots of pressure versus scan number were used to identify what data needed to be removed from the initial soak that occurred for some casts. This is necessary to ensure DELETE chooses the data from the full cast for the top 10m. CLIP was run to remove the first 800 records from the first 3 files because data was acquired during a 10m soak. It was also used to remove the first 500 records from the other casts as it may simplify editing and will cause no loss of useful data since all were acquired during the soak period.

##### Shift

Fluorescence

The fluorometer was not pumped, so a shift in alignment is expected to be small or unnecessary. Profile plots of temperature and fluorescence were examined but the complexity of the near-surface temperature profiles made the comparison uninformative. For the previous cruise the fluorometer alignment was thought to be ok, so no further alignment was applied to these data either.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked and the alignment looks ok though both traces are complex. No further alignment was applied.

Conductivity

Tests were run on many casts using a variety of shifts. During the previous cruise the best results were found using +0.4 records. For this cruise there is very little effect, but +0.4 makes very tiny improvements on a few casts.

SHIFT was run on the CLIP files using +0.4.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

This is the second known use of the pressure sensor. There was a study of pressures during 2016-02 and it was found that the February 2015 calibration produced better results than that from July 2016. No further correction was made. The conductivity sensor was used for just one cruise, 2016-02, since it was last calibrated in August 2014. The dissolved oxygen sensor had been used during 2014-23 when the correction applied was slope 1.04, offset 0.0037. It either malfunctioned or was not really mounted during 2015-26. During 2016-02 the fit was 1.0407/0.1351 but the 2014-23 correction was used since the fit was not trusted due to sampling problems.

Historic ranges – There were areas of temperatures above the range maximum between 130m and 180m at some stations in the northern half of the Strait of Georgia, mostly towards the west. Such excursions have been common in this area for the past year.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

CTDEDIT was used to remove noisy near-surface records and records corrupted by shed wakes near the bottom of casts. Salinity was cleaned lightly where it appeared that small alignment variations led to unstable salinity features. All files required light editing.

##### Initial Recalibration

There is no evidence that CTD pressure and salinity need recalibration.

It is likely that the dissolved oxygen does need recalibration but we have no calibration samples to guide us in that decision. When the sensor was used in 2014 the correction applied was

CTD DO Corrected = CTD DO * 1.04

The same sensor was used during 2016-02 when there were dissolved oxygen samples. There was a very unusual fit of bottles versus CTD DO which was possibly due to bubbles in samples. Given few bottles, a large scatter in results and a known sampling problem, the comparison was not trusted.

For this cruise file 2016-03-ctd-recal1.cal was prepared to apply:

CTD DO Corrected = CTD DO * 1.04

The SBE Dissolved Oxygen in the rosette files may be a little low but there is no way to determine by how much.

The MRGCLN2 and ED1 files were recalibrated using file 2016-03-ctd-recal1.ccf to apply.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files for a few test plots. Plots showed that the filter was had little effect on spikes and over-smoothed other data. This step was skipped.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots look fine.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Conductivity:Primary, Oxygen:Voltage:SBE, Altimeter, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

----------------------------------------------------------------------------------

Data Processing Notes:

----------------------

Fluorescence and transmissivity data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Salinity calibration sampling suggests that calibration drift is small, but

the evidence is limited.

There was no dissolved oxygen calibration sampling. Recalibration was based

on an earlier cruise. No estimate can be made for the accuracy of the

Oxygen:Dissolved:SBE channel.

For details on the processing see the report: 2016-03_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

A cross-reference list was produced.

The sensor history was updated.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The surface saturation values were between 92% and 110% with all but 2 within 95% and 105%. These are reasonable results although we often see a wide range of values in this region, so it is weak evidence.

##### Final Bottle Files

The MRGCOR2 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Conductivity:Primary, Descent Rate, Altimeter, Descent_Rate and Flag.

A second SBE DO channel was added for the CTD DO with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check and a header check were run on all files and no errors were found.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

A cross-reference listing was produced for the CHE files.

Data were exported from the CHE files to file 2016-03-bottles-final.xlsx. The entries were compared with the rosette log sheets to ensure no samples had been missed and no problems were found.

# Notes from Log Book

General

The deployment method varied through the cruise.

On March 15 the CTD was kept in the water at the surface for 2 minutes, then went down 10m and back before the full cast.

From 16 March to 20 March there was a 1.5 to 2 minutes soak at the surface only

Particulars

1-5. Cruise number wrong in headers. Fixed.

16. Upcast salinity very spiky.

42. Forgot to inform Eng. Dept., ship may have been pumping out. (Shouldn’t affect sampling though high fluor. in upper 10m, not just surface.)

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 1091 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4700 | 20Aug2014 | Factory |  |  |
| Conductivity | 3321 | 20Aug2014 | Factory |  |  |
| Transmissometer | 983DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 1483 | 17Jan2014 | Factory |  |  |
| WetLabs ECO Fluor | 2216 | n/a |  |  |  |
| Altimeter | ? |  |  |  |  |
| Pressure Sensor | 1091 | 12Feb2015 | Factory |  |  |