## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 25 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| ### 16 August 2020 | ### Added HPLC Data. S.H. |
| 5July2018 | Added DIC/ALK to 2016-05-0047.CHE |

## PROCESSING NOTES

Cruise: 2016-05

Agency: Ocean Sciences Division

Location: Strait of Georgia and Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Strait Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: April 21, 2016 – April 26, 2016

Processed by: Germaine Gatien

Date of Processing: 28 June 2016 – 13 July 2016

Number of original HEX files: 90

Number of CTD files: 90	Number of bottle files: 27

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#1119) on the secondary pump, a SeaPoint Fluorometer (#3865) on the primary pump with a 3X cable, a Biospherical QSP-400 PAR sensor (#4565), Surface PAR (#20518) a pH sensor (#0692) and an altimeter (#1253).

The data logging computer was Vector CTD Laptop.

The data acquisition program was Seasave v7.23.1.

The CTD deck unit was an SBE model 11+.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log book and sampling/rosette log sheets were in good order.

There was a discrepancy between the fluorometer serial number recorded in the log book and that in the configuration file used for the first few casts. A change was made to the configuration file later and noted in the log, but having photos of the sensors made it easy to confirm which sensor was in use.

One file was edited at sea to fix the station name. The resulting file was impossible to convert without some adjustments to the hex file as some carriage returns had been removed. It is safer to note in the log that there is an error in the headers rather than try to fix it.

After the initial 10m soak, a wait of about 30s at the surface is recommended. The full cast began a few seconds after data acquisition began. There may have been a wait before acquisition began but it would be useful to have a record of the wait time to ensure the wait times are adequate.

Salinity samples were analyzed 9 weeks after collection so evaporation and adsorption of samples has likely raised the salinity values of samples, with some randomness in the effect. There is also reason to expect incomplete flushing of Niskin bottles due to very calm waters for many of the casts. The errors due to those effects all lead to higher salinity in bottle samples. The comparison with bottles had a lot of scatter with the primary salinity appearing to be low by ~0.007 and the secondary by ~0.006. There are too few samples to tease apart the sources of error or to make a reasonable estimate of CTD salinity accuracy, except to say that the error is likely smaller than it appears. No recalibration was applied to the CTD salinity.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.7 mL/L from 0 to 20db

±0.3 mL/L from 20db to 200db

±0.1 mL/L below 200db

The CTD fluorometer values were generally lower than the extracted CHL with fluorescence averaging about 56% of the extracted chlorophyll. Usually this type of fluorometer reads higher than CHL for low CHL values, but there were few low CHL values.

No TSG data was acquired due to a problem with the loop system.

The pH:SBE data were bad from casts 2-7, 14, 17-21 and 90 so the channel was removed from those files. In 2 cases the cap was left on the sensor during the cast. For the others it looks as though the sensor was not hydrated between casts as the data gradually improved through the casts.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data were available at the time of processing. Calibration is required for each cast to get absolute values although

general trends within a cast are likely real.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

It was noted that the fluorometer serial number in the configuration file did not match the one in the log book. Photos confirm that the log record is correct. Later in the cruise the fluorometer serial number was changed.

There was no history for the conductivity and DO sensors; the DO sensor had been used for 2 other cruises.

The only change to the XMLCON files through the cruise was a correction to the fluorometer serial number. The calibration constants were checked for all instruments and no other problems were found. The corrected file was saved as 2016-05-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2016-05-ctd.xmlcon to create CNV files.

The station number for event #101 was corrected (25→27).

A few casts were examined.  All expected channels are present. The primary and secondary temperature and conductivity channels are reasonably close during downcasts, but as usual the upcasts differ more due to noise in both channels. There are some spikes in primary conductivity that are not seen in the temperature or dissolved oxygen that were on the same pump. Dissolved oxygen, fluorescence, PAR, pH, altimetry and transmissivity profiles all look normal. PAR maximum values were usually lower than Surface PAR, but when the minimum sampling depth was low PAR was higher. The descent rate was mostly fairly high and steady, but the variability was much higher near the mouth of Juan de Fuca Strait.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2016-05-ctd.xmlcon.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and problems were seen in the primary salinity in 6 different casts and in the secondary salinity in 3 casts, with the primary spikes larger.

Examination of a few more CTD profiles suggests there may be serious problems with the primary conductivity, but until the salinity comparison is done, both salinity channels will be processed as usual.

CTDEDIT was used to edit the BOT files for events 3, 26, 43, 69, 74, 81, 88. Comments about editing were added to the headers. The output files were then copied to BOT.

A preliminary header check was done and no problems were found. Fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine. Those files were bin averaged on bottle number.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2016-05-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2016-05chl*.xls. The file included comments, flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2016-05chl.csv, event numbers were added, and the file was then converted to individual CHL files.

The analyst noted that the rosette spreadsheet indicated that sampling was done at 5m for casts 1-18, but the labels indicate it was at 10m.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2016-05oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2016-05oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2016-05SAL.xls. The analysis was done within 62-67 days of collection. The files were simplified and saved as 2016-05sal.csv.

That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2016-05nuts.xls.

Then the file was simplified, reordered on sample numbers and saved as 2016-05-nuts.csv.

The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. No problems were found.

CLEAN was run on the MRG files to add 0 flags to empty flag channels and to update header limits.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There is a lot of scatter in the comparison with differences. The 9-week wait for analysis makes significant evaporation and adsorption of samples likely, with the former tending to be a random effect. Both will raise the bottle salinity making the CTD salinity look lower than it really is. Further, this cruise had many casts with very steady descent rates and that is associated with poor flushing of Niskin bottles; the Niskins would contain water from deeper in the water column and hence have higher salinity than that measured by the CTD at the firing depth. With all these factors causing bottles values to be high, it is not surprising that both CTD salinity channels are reading lower than most bottles with differences up to 0.042psu.

The 2 largest differences come from cast #1 at 75m and #81 at 149m. The cast #1 outlier is associated with a very steady descent/ascent rate and fairly high salinity gradient. The cast #81 outlier is associated with some spikes in salinity and the standard deviation in the CTD salinity is large during the 10s window used in the comparison. So those bottles are excluded from the comparison. Two other bottles were flagged 3 by the analyst. They don’t look particularly out of line, but given the scatter that does not rule out a problem with the samples.

When the 2 outliers and the 2 flagged values are excluded the following data are obtained:

|  | Average | Std Dev | Median |
| --- | --- | --- | --- |
| Sal0 - Bottle | -0.0073 | 0.0063 | -0.0066 |
| Sal1 - Bottle | -0.0066 | 0.0061 | -0.0056 |

The median differences lead to an estimate of Sal1 – Sal0 = 0.0011 which is in good agreement with the differences listed in section 9.

The deepest 4 un-flagged values show the CTD primary salinity low by 0.0057 and the secondary low by 0.0045; both have a standard deviation of 0.0019.

The two casts from near the mouth of Juan de Fuca Strait, where descent rates were fairly noisy and bottles likely flushed well, show the primary CTD lower than bottles by 0.0009 in one case and higher by 0.0014 for the other. If there was significant evaporation or adsorption for those samples then both CTD salinity channels may be reading high. An estimate for the effect of adsorption is +0.0012 psu/month, so about +0.002 for this cruise. Evaporation would be somewhat random depending on how well each bottle was sealed. There are too many sources of error and too few data to enable a reasonable estimate of accuracy for the CTD salinity.

For full details for the COMPARE run see file 2016-05-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The fit of (CTD DO – Titrated DO) vs CTD DO is unusual with a small correction looking appropriate. Examination of CTD DO profiles shows that many of the casts had low DO gradients at mid-depths with many cases of DO reversals.

When all casts were selected, a few notable outliers removed and more data removed based on residuals, the fit found was:

CTD DO Corrected = CTD DO * 1.0142 +0.0378  (R2 = 0.07)

or	CTD DO Corrected = CTD DO * 1.0118	    (R2 = 0.19)

Next casts were examined that had DO decreasing fairly steadily with depth:

CTD DO Corrected = CTD DO * 1.0122 +0.0068  (R2 = 0.42)

or	CTD DO Corrected = CTD DO * 1.0136	    (R2 = 0.42)

In contrast casts which were well-mixed or had reversals in DO:

CTD DO Corrected = CTD DO * 1.0028 +0.0443  (R2 = 0.04)

or	CTD DO Corrected = CTD DO * 1.0114	    (R2 = 0.37)

Next, casts were examined that had very steady descent rates:

CTD DO Corrected = CTD DO * 1.0012 +0.0487  (R2 = 0.006)

or	CTD DO Corrected = CTD DO * 1.0111	    (R2 = 0.38)

And casts with noisier descent rates (none were very noisy):

CTD DO Corrected = CTD DO * 1.0101 +0.0168  (R2 = 0.46)

or	CTD DO Corrected = CTD DO * 1.0136	    (R2 = 0.40)

It makes sense that the correction is larger for the casts with a noisier descent rate since the bottles likely flush better. Where flushing is weak the water in the bottle is likely from a little lower, with a lower DO than would be found at the level of the bottle firing. So the CTD would look closer to bottles when descent rates are very steady. There is also a larger correction for the casts with DO increasing steadily with depth than for well-mixed casts. This is likely related to slower response when gradients are larger.

The correction factors estimated in all these ways vary from 1.011 to 1.014. These are likely all underestimates of the calibration errors due to incomplete flushing of bottles.  A final check was made by using only the 2 casts closest to the mouth of Juan de Fuca Strait; these were deep with fairly noisy descent rates and fairly large DO gradients. The fit was:

CTD DO Corrected = CTD DO * 1.0152	    (R2 = 0.74)

Multiplying all DO values by 1.015 looks like the best estimate we can make.

There were 3 significant outliers but 2 of those came from cast #1 which usually looks out of line because of extreme gradients. The only serious problem is with sample #181 from cast #61, 40m. This was flagged 4 by the analyst due to a serious leak in the bottle and both DO and nutrient values look out of line in profile. The analysts were informed and padded the DO and NUTS values with flag 5.

For more detail see document 2016-05-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor. There was a large range of CHL values with 74 samples between ~0.3ug/L and ~15ug/L and 1 sample at 38ug/L. When the highest values is excluded the fit is: 	 SBE Fluorescence  = 0.54 * CHL

But looking at the fit of CTD/FL vs CHL shows that the highest value is not necessarily an outlier; rather it may be that fluorescence is gradually falling relative to CHL.

Two samples originally flagged 2 by the analyst appear to be bad, so pad values were entered.

Event 9, sample from 75m  - not listed on rosette sheet and value out of line. Flag 5.

Event 35, sample 122 – bottle did not close – don’t know where this sample came from. Flag 9.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

A few casts were checked before and after this step to see if spikes in conductivity were reduced and they were. So the problems noted in the bottle files may be less serious in the profile files.

##### ALIGN DO

Tests were run on ALIGNCTD to determine the best setting to advance the DO signal. The results were difficult to judge because the temperature and DO profiles are so complex. Settings from 2s to 3.5s looked best, with +2.5s best overall.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. Tests on previous cruises using these sensors showed the default setting of (α = 0.0245, β=9.5) did the best job and it does improve the data for both conductivity channels for these data.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. None are very deep.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2016-05-0060 | 350 | -0.0002 | +0.0001 N | +0.0013 | High, Steady |
| 2016-05-0085 | 350 | -0.0001 | +0.0001 N | +0.0012 | High, Steady |
| 2016-05-0101 | 330 | -0.0002 | +0.0001 N | +0.0013 | High, F.Steady |

The differences are small. The conductivity differences are not large, but they are noisier than usual.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. No errors were found.

A header check was run. No problems were found. No off-scale fluorescence values were found.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.4db which is reasonable for the Vector. The shallowest data came from 0.48db during cast #17. The salinity looks reasonable for that depth and PAR is high, but there is no means of judging pressure accuracy.

The altimeter and bottom depth headers were exported from the SAMAVG and CLN files to spreadsheets. The bottom depths were checked against the log book. There were many discrepancies; mostly the log entries had been changed so may be a better reflection of the depth when the CTD was at the bottom. In a few cases there was obviously an error in the header entry.

As a further check the following calculation was made:

Check=maximum depth sampled (Max press*0.99) - water depth +altimeter reading

We don’t expect a 0 reading since the altimetry reading is an average over 2db and the bottom depth is subject to changes during the cast. Some of these differences will be due to errors in the bottom depth entries. Where changing the bottom depth doesn’t resolve the problem, plots were examined to see if the altimetry header is inappropriate.

The following changes were made:

- Bottom depth changed to match the log entry: 6,23,51,59,86,95,100,103

- Altimetry header was removed for: 54

No changes are needed to the bottle files.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and as usual, there is a lot of variability with the up and downcast traces sometimes closer than temperature and sometimes further apart. This is likely due to varying vertical gradients. Overall the choice made earlier looks appropriate, so no further alignment will be applied.

Conductivity

Tests were run on 4 casts using a variety of shifts and the best results varied from cast to cast and feature to feature, but overall -0.4 records for the primary and -0.5 records for the secondary looked best.

SHIFT was run twice on all casts using those settings.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. The last time this sensor was used a setting of +90 records had the best results. Tests were confusing with values between +50 and +80 looking best for some features, but no value looking good for others. Plots of the un-shifted data explain why – much of the downcast pH data looks bad. Casts 2 to 6, 14, 17-20 have low downcast values that gradually move to more believable values. A few casts starting with event #21 have near-surface pH values that are probably a little low – this may be true later as well, but not as notably so. From cast #30 onwards the signal looks reasonable except for event #90 which has very low almost constant low values – the cap must have been left on the sensor.

SHIFT was run on the pH:SBE channel using a setting of +70 records.

Plots were examined after this step and the results look ok.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

The temperature, conductivity and dissolved oxygen sensors had no history of use since the last factory calibration. The pressure sensor was used for 3 Ricker cruises in 2015; no recalibration was found necessary during those cruises.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. As has been seen in all Strait of Georgia/ Juan de Fuca Strait cruises over the past year or two, temperature is frequently above the top of the range, usually between 25 and 75db, but in the southern part of SoG somewhat deeper. There are also some high surface values in the northern section. In the central Strait of Georgia all temperatures were within the climatology. Salinity is mostly within the climatology range except for frequent cases of low salinity in the top 10db. The 3-standard deviation climatology is not recommended for near-shore work. This check is usually done to check for systematic errors in calibration, but the same results are being seen from a variety of sensors over the past year.

Repeat Casts – There was no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

There is little to distinguish between the primary and secondary T and S channels. The secondary channels have slightly less spiking so those were selected, though for cast #45 there were bad sections in the secondary data, so the primary sensors were edited.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. For a few casts there were sections of bad salinity points where temperature was ok; such salinity points only were removed. All files required some editing.

##### Initial Recalibration

Pressure does not appear to need recalibration.

No recalibration will be applied to the CTD salinity. The comparison with bottles had a lot of scatter and while the CTD salinity is lower than bottle samples, this is likely due largely to evaporation and adsorption of samples plus poor flushing of Niskin bottles. Making crude estimates of those errors suggests that the CTD salinity is reasonably close to bottles.

Based on the comparison of section 5, CALIBRATE was run using file 2016-05-recal1.ccf to correct the Oxygen:Dissolved:SBE channel in the SAM and MRGCLN2 files using:

CTD DO Corrected = CTD DO * 1.015

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly. The average of differences in the fit once outliers were removed was +0.012mL/L and the standard deviation is 0.027mL/L.

See file 2016-05-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed, the CTD DO was higher than the bottles by an average of ~0.06mL/L and standard deviation of 0.15mL/L; the results vary with depth with the CTD reading low right at the surface, high between 5m and 150m and close to bottles below 150m. This pattern is likely due to incomplete flushing in the presence of significant DO gradients above 150m and slow response of the CTD DO in the presence of frequent reversals of DO gradient near the surface. No further recalibration of DO is justified.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

The CTD files from rosette casts were clipped to 50m; sigma-T was derived and the data were exported to a single file, 2016-05-SOG.csv.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some small unstable features, but those are from sites where active mixing is expected.

There are some spikes in fluorescence that might suggest pump problems, but there are also spikes in the unpumped transmissivity channel, often at the same levels, so these may well be real features. PAR values are generally lower than PAR:Reference, but the shallowest PAR data are from 2 to 4m the PAR profiles look fine. The pH, dissolved oxygen, temperature and salinity look ok.

The pH profiles for casts 2-6, 14 and 17-21 look bad. The downcasts look very different from the upcasts as though the sensor was very slow to respond to in water values. Perhaps the sensor was not kept hydrated between casts. Cast #1 looks fine. Cast #7 values are low but there is a log note that the cap was left on. It is surprising that cast #15 looks ok, but that may be because it occurred within 30 minutes of the previous cast so it may have stayed wet. For all casts from #22 to the end of the cruise the profiles look normal and during the cruise that followed, 2016-45, no problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts except #45 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

pH was also removed for casts 2-7, 14 and 17-21.

REMOVE was run on cast #45 to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, pH, PAR and PAR:Reference data are nominal and

unedited except that some records were removed in editing temperature and

salinity.

NOTE: While the SBE fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision. The results were unusual

with a smaller correction than usual, which is likely due to

incomplete flushing of Niskin bottles and low dissolved oxygen

gradients at mid-depths. See the processing report for details.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.7 mL/L from 0 to 20db

±0.3 mL/L from 20db to 200db

±0.1 mL/L below 200db

The pH channel was removed from casts #2 to 7, 14, 17 to 21 and 90 as the

downcast data were clearly bad. For most of those casts, it is possible that

the sensor was not hydrated between casts. For cast #14 it is known that the

cap had been left on during the cast and for cast #90 it appears that the cap

was left on.

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see the report: 2016-05-proc.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and it was found that there was one other cast with bad pH, #90. It looks like the cap was left on this one since the upcast and downcast both look alike with virtually no variation, similar to the plot of cast #7.

All the steps in this section were rerun on #90. No further errors were found when the Header Check routine was rerun.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values were highly variable ranging from 85% to 185%, with most <135%. The lowest values were found in Haro Strait and eastern Juan de Fuca Strait and the southern Gulf Islands region. The highest value was in Saanich Inlet while other high values were mostly in the southern Strait of Georgia. Given the wide range in surface saturation, these observations are not helpful in assessing the DO sensor calibration.

##### Final Bottle Files

CALIBRATE was run on all files and then the MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

pH was also removed for casts 3 and 19.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check and a header check were run on all files. An error was found in file #69 – only nutrients had been identified as from cast #69, thus were missed in the merge process. The writing was unclear on the rosette sheet and some analysts thought it was #67.  After that was fixed, Header Check was rerun and no further errors were found.

The track plot looks ok.

Data from the CHE files were exported to a spreadsheet and compared with rosette sheets and no errors were found.

Plots of each file were examined and no problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars (mostly notes from log)

TSG manifold failed. No TSG data were acquired.

7. pH cover left on.

9. Con file changed to correct serial number of fluorometer. No change of equipment.

28. Grey water dumped when CTD at 20m depth.

57. Fluorometer spikey, alt shows 6.8m.

62. Error in hdr and hex files. The headers had been edited at sea and some lines were concatenated.

85. #2 generator exhaust with rosette.

101. Station number wrong in header – should be 27, not 25. Fixed.

108. No water for bottle #1 – did not close.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 18Dec2014 | Factory |  |  |
| Conductivity | 1766 | 20Aug2014 | Factory |  |  |
| Secondary Temp. | 4054 | 20Aug2014 | Factory |  |  |
| Secondary Cond. | 3531 | 6Feb2013 | Factory |  |  |
| Transmissometer | 1396DR | 2Feb2016 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 19Dec2015 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| Surface PAR | 20518 | 21Mar2016 |  |  |  |
| pH | 0692 | 12Jan2016 | Factory |  |  |
| SeaPoint Fluor. | 3865 | n/a |  |  |  |
| Pressure Sensor | 0550 | 17Dec2014 | Factory |  |  |
| Altimeter | 1253 | n/a |  |  |  |