## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 11 Mar 2025 | ### Corrected channel name PAR:1 to PAR1 in CTD cast #1. TSG channel names/units updated. G.G. |
| ### 18 Jan 2023 | ### Corrected Salinity:Bottle format, adjusted some flags and comments. G.G. |
| ### 2 Dec 2021 | ### Corrected the Salinity:Bottle precision that was lost during the HPLC addition. Re-did dmsp comments 70 & 81, re-did carbon addition for 5, 9 &75. , DMPS comments 70, 81.S.H. |
| ### 19 July 2021 | ### Removed previous DIC and TA additions from 5 files, added updated DIC, TA and pH data to 11 files. SH. |
| ### 30 April 2021 | ### Added NUTS comments to phosphate flags event 34. S.H. |
| ### 15 April 2021 | ### Added DIC and TA data and changed some phosphate flags for event 34. S.H. |
| ### 1 April 2021 | ### DMSP clean-up – edited comments to events 70 & 81. S.H. |
| ### 16 Aug 2020 | ### Added HPLC Data. S.H. |
| 6 Feb 2020 | Corrections to comments for DMSP. G.G. |
| 19 June 2019 | Added DMSP data to 6 CHE casts. Only pad values due to contamination. |

## PROCESSING NOTES

Cruise: 2016-06

Agency: OSD

Location: North-East Pacific

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: 5 June 2016 – 21 June 2016

Processed by: Germaine Gatien

Date of Processing: 25 August 2016 – 17 October 2016

Number of original HEX files: 49 	Number of CTD files: 49

Number of bottle files: 48		Number of bottle casts processed: 48

Number of original TSG files: 3		Number of processed TSG files:	 3

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), a SBE 43 DO sensor (#997) on the secondary pump, a SeaPoint Fluorometer (#3640) on the primary pump, QSP-400 PAR sensor (#70613) and an altimeter (#62354). A second PAR sensor (#4615) was mounted for only the 1st cast.

A thermosalinograph (Seacat 21 S/N 3411) was mounted with a Wet Labs WETstar fluorometer (S/N ws-3s 953p), remote temperature sensor and a flow meter.

Seasave version V7.23.2 was used for acquisition.

The data logging computer was the Tully CTD Laptop (Acer).

The deck unit was a Seabird model 11+, serial number 0425.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

The oxygen kit was Scripps kit #1.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log, rosette log sheets and analysis logs were generally in good order with detailed comments on problems. Notes from the Chief scientist were very helpful. Nutrient and salinity samples from X-Niskin casts were mixed with the SBE911+ rosette samples. To avoid time being wasted trying to find CTD data corresponding to such samples, it would be helpful to have a list of X-Niskin casts in the sampling notes or if there is some indication on the label to ensure the results are entered separately in the analysts’ spreadsheets.

The salinity analyst recommended that salinity bottle samples be reported to only the 3rd decimal place due to a poorer pooled standard deviation than expected.

The comparison of salinity samples with the CTD salinity suggested that the primary salinity was low by about 0.006 and the secondary by about 0.0015. There was an unusually large difference between the two temperature sensors, with the secondary lower than the primary by ~0.002Cº. It is not known which temperature sensor is more out of line, but correcting the primary temperature by subtracting the difference and recalculating salinity decreases the difference between primary salinity and bottle salinity. Correcting the secondary temperature by adding the difference would decrease the secondary salinity and hence increase its difference from bottles. Having the primary salinity low by 0.006 seems less likely than having it low by 0.004, so the primary temperature was recalibrated. After the temperature correction was made, a further recalibration was applied by adding 0.004 and 0.0015 to the primary and secondary salinity channels. If a post-cruise calibration shows this choice to have been inappropriate, salinity can be recalibrated later.

The CTD fluorescence shows the usual pattern, reading higher than the extracted CHL for CHL<1ug/L and then falling relative to CHL until it is about 50% of CHL at high CHL values.

Two PAR sensors were mounted for the first cast only as a test. For the rest of the cruise only PAR sensor #70613 was used; this sensor can be left on the CTD for casts as deep as 2000m.

There were some unusual variations in the comparison of titrated dissolved oxygen samples with CTD dissolved oxygen. Some data with low DO values did stand out in the comparison but this was not due to hysteresis. Data from casts at P4 did stand out, especially from event #106. This may be due to P4 being a cast with a shallow and thin oxygen minimum zone compared to casts from further offshore. The DO values at P4 are the lowest from this cruise. Further offshore the OMZ is deeper with higher DO values and a thicker layer where the CTD likely had more time to equilibrate and any flushing errors are likely to be slight. To determine a DO correction a fit was selected that excluded bottles from near the OMZ for casts #14 and #106 plus a few outliers based on residuals.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

The Oxygen:Dissolved:SBE data in the CTD files (downcast) are considered, very roughly, to be:

±0.4 mL/L from 0 to 100db

±0.3 mL/L from 100db to 500db

±0.1 mL/L from 500db to 1000db

±0.05 mL/L below 1000db

The thermosalinograph worked well throughout the cruise with large salinity spikes restricted to limited sections of the record. There were no large spikes between stations P6 and P21 or during the return trip or during stops for casts. Small bubbles may be present especially when the ship is moving. A small correction was made to the TSG salinity based on comparison with the CTD and loop samples; this should improve the overall accuracy but may overcorrect while stopped and where there are no bubbles, and under-corrected where bubbles are more significant. The TSG fluorescence is given in voltage units; for a very rough estimate in concentration units multiply by a factor of 3 west of station P2, but keep in mind that the ratio is highly variable. Closer to shore the factor would be higher and highly dependent on extracted chlorophyll values. The TSG intake temperature was higher than the CTD temperature by a median of ~0.0024Cº which is a good result. DO loop samples taken from the transducer chamber were closer to surface rosette DO values than those from the lab, but were still significantly higher than those from the surface rosette.

There were some problems with matching loop samples to events in the log book. There should have been loop samples during event #45 at P14 and shortly afterwards while underway. The only samples found were labelled P14 full speed, but there is no entry in the log so we don’t know when the samples were taken. The salinity sample suggests that the ship was moving, so the label is correct, but no samples were found for the stop at P14. There were 2 dissolved oxygen samples labelled Loop P4. Since there were casts at P4 on both the outward and inward parts of the cruise this labelling was insufficient. A sample number was determined by comparison with the CTD dissolved oxygen in rosette files and checking when samples were analyzed. The salinity sample labelled LoopP11 does not compare well with either the TSG salinity or the rosette data (CTD and salinity sample) so was likely not gathered during the stop at P11.

Assigning distinctive sample numbers to loop samples may fix the problems with loop sample identification. Using a leading value such as 8, for example sample #8006 for the 6th loop sample, would ensure there is no confusion with non-loop samples. A leading 9 is already commonly used to fix cases of accidentally repeated rosette sample numbers, so that would not be a good choice.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Nutrients, extracted chlorophyll, dissolved oxygen, DMS and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked. All were used during 5 previous cruises.

All configuration files were the same. File 2016-06-ctd.xmlcon was prepared by making two changes to the file used at sea.

- The PAR sensor is a new one and the factory calibration includes an offset that was not entered in the con file, so that was added.

- An adjustment of -0.6db to the pressure offset was determined to be necessary during processing of cruise 2015-54 (October) when there was sampling very close to the surface. During 2016-12 which followed this cruise there were many deck measurements that were all between 1.1db and 1.3db. So, the offset may be increasing. It was changed to -1.0 which is the factory setting from the last servicing (0.05243db) minus ~1.05db. This can be fine-tuned later if it looks necessary.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2016-06-ctd.xmlcon to create CNV files. The hysteresis and Tau functions were selected. There are some deep casts. There were 2 PAR sensors in the configuration file throughout the cruise but sensor #4615 was only mounted for the first cast as a test, so channel PAR1 was only converted for file 2016-06-0001.

Three files had the wrong format in the file names, so that was corrected after conversion. (Ex. 2016-06-0038hex.cnv was changed to 2016-06-0038.cnv.)

A few casts were examined and all expected channels are present.

The transmissivity was very low at the bottom of cast 13 which confirms the observation at sea that the CTD might have hit bottom.

There are spikes in both conductivity channels, but mostly in the primary. Primary temperature has some odd excursions and spikes.

The altimeter appears to have failed for one cast, but looks ok for others. Fluorescence and transmissivity look ok.

The 2 PAR channels are reasonably close for cast #1 and the traces look normal.

Dissolved oxygen looks normal.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2016-06-ctd.xmlcon. The naming format was changed for 3 casts.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. For casts 5, 10, 12, 19, 20, 84 and 85 a few spiky points were removed from one or the other salinity channel using CTDEDIT.

The output files were copied to *.bot.

A preliminary header check and no problems were found. CTD fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

One error was identified:

- File 2016-06-0105 should be named 106 as event 105 was a BOT cast not a ROS. There is no file for event 106 which was a ROS cast and the time in the file is right for 106. The name was corrected for both the rosette and full file.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2016-06-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2016-06chl*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2016-06chl.csv. The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2016-06oxy.xlsx which includes flags, comments and a precision study. Draw temperatures are available. Filtered samples that were run for comparison were removed.  The spreadsheet page with the final data was simplified and was then saved as 2016-06oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in file 2016-06SAL.xlsx which included a precision study. The analysis was done between 9 and 28 days of collection. There were samples from the X-Niskin mixed in with samples from the rosette. It is time-consuming to separate these when there is no hint in the sample numbers. If there is some indication on the labels, then it would be helpful to enter that information in the spreadsheet and then make an individual sheet for those samples as is done for loop samples. The files were simplified and saved as 2016-06sal.csv. File 2016-06sal.csv was then converted to individual SAL files. Note that Salinity:Bottle is only given 3 decimal places due to poor results from the precision study.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2016-06_nutrients*.xlsx. This includes a precision study. The file was simplified, reordered on sample numbers and saved as 2016-06-nuts.csv. The file was converted to individual NUT files.

DMS

DMS data were obtained in file DMS summary (2016-06).xls. Values given as < were changed to 0 and the comments that will go into the header will explain that 0 means below minimum detectable level. DMS: was entered before comments. The file was then saved as 2016-06DMS.csv and converted to individual DMS files. There was a separate report on analysis techniques and problems.

The SAL, CHL, OXY, NUT and DMS files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. There were a number of cases where dissolved oxygen sampling was indicated on the sampling log sheets but no sample value was found in the analyst’s spreadsheet, so it is assumed this sampling was done for a non-IOS group. An error was found in file #50. Sample #265 came from Niskin 3, not Niskin #1. The ADDSAMP file was edited and the CST and SAM files rebuilt. The MRG file was reviewed and looks correct. The line for Niskin #1 was removed.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There was one extreme outlier – the sample from cast #25. This is clearly a case of a misfire as mentioned by the analyst. The dissolved oxygen sample and draw temperature are both way out of line as well. This bottle clearly closed at the surface. The sample values were replaced with pad values and COMPARE was rerun for salinity and dissolved oxygen.

There were another 5 outliers, but 2 were near-surface samples and there was a high standard deviation in the CTD salinity.  The other 3 were:

- Sample 131, event 22 –Outlier in comparison with CTD salinity; bottle salinity is higher than any salinity seen in the profile by about 0.06. There was no liner on the bottle. The analyst replaced the value with a pad value.

- Sample 140, event 25 – CTD salinity low compared to bottle by ~0.03, depth ~800m. Salinity that high is found about 20m lower in the water column and it is unlikely that flushing was as poor as that. The analyst added a 3 flag.

- Sample 254, event #46 –CTD salinity low compared to bottle by ~0.12. That salinity is much higher than any sampled by the CTD in the cast. The analyst added a 3 flag.

When these outliers and all values above 200m were excluded the primary salinity was found to be low by an average value of 0.0065 and the secondary low by 0.0020. The standard deviations were 0.0020 for both. The 22 bottles fired during the calibration cast (#75) had very low standard deviations in the CTD salinity channels. The primary was found to be lower than bottles by an average of 0.0072 and the secondary by 0.0016 with standard deviations in both being 0.0019. The samples were analyzed quickly, so evaporation and adsorption would be expected to have little effect on samples, likely no more than 0.0005, so CTD primary salinity might be low by ~0.006 and the secondary by 0.0015.

During 2016-47 there was a large difference was found between the temperature channels, with the secondary reading lower than primary by ~0.002 which would account for a difference of ~0.002 between the two salinity channels. For this cruise the temperature differences are similar but salinity differences are a little higher. It is not known if the primary temperature is reading too high, the secondary too low or a combination of both. However, if the secondary is corrected by adding 0.002, the secondary salinity would be even further from bottles. If the primary is corrected by subtracting 0.002 the primary salinity would move closer to bottles. So it looks more likely that the primary temperature is the principal source of the difference. Correcting primary temperature and recalculating primary salinity would make it look low by about 0.004.

The analysis precision was poor at Sp=0.002 which the analyst considered poor given the quick analysis. As a result the analyst gave results with only 3 decimal places.

Plots of differences against time and against salinity do not suggest any trend that cannot be explained by different depths and local gradients.

For full details for the COMPARE run see file 2016-06-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There are 3 significant outliers. One of those was an outlier for all bottles and was determined to be due to a misfire in the salinity comparison. It also looks like a misfire in this comparison. For the other 2 cases the bottle value was higher than the CTD DO which cannot be explained by poor flushing of bottles.

- Cast 25, 149m – sample 147 – Both bottle and CTD profiles have slight reversals but in opposite directions. Bottle 4.20, CTD 3.18. Examination of the data just before and during the stop show that the bottle could contain water higher than the CTD would have measured by up to 0.2mL/L, but that doesn’t come close to explaining the difference seen in the comparison. The problem turned out to be due to a mix-up in samples. Once corrected this case was not an outlier.

- Cast 34, 399m - sample 196 - The local gradient was very high so the CTD may not have equilibrated well. There was a reversal above the bottle, but no obvious sign of trouble during the bottle stop. The nutrients show the bottle likely didn’t close late. Bottle 2.26, CTD 1.71. There is no evidence of problems with sampling or analysis. Not enough evidence to justify a flag.

The fit looks odd, with some casts looking out of line. When looking at the low DO values, this suggests that the problem might be hysteresis. When all data are examined there is no sign of hysteresis. What does stand out is cast #106 and to a lesser extent cast #14, both at P4. This is likely because the oxygen minimum zone has a different shape at P4 with lower values (<0.19mL/L from uncorrected CTD values) and less time to fall from 0.2 to the minimum. The DO sensors have slower response at very low DO, so a small difference is expected if there is less time to respond. The differences between cast #14 at P4 and most other casts is small, it just stands out more in plots near the origin as the scatter is much smaller there. The difference is somewhat greater for cast #106 and this maybe because the oxygen minimum was right at the level sampled, whereas the minimum was below the 800m bottle for #14. So the variation at the DO minimum is likely not significant, but does affect the fit.

The scatter is also unusual for event #106, not just at the minimum but for all bottles except the 2 right at the surface. The bottles at 25 and 50m are out of line, but that is often the case at those depths. All bottles below 50m have CTD lower than the bottles by more than expected for this fit or any recent fits from this equipment. This is always a site of rapid change and particularly between 250m and 400m there are large differences, especially in dissolved oxygen, with rapid reversals. In one 37m section DO changes from 2.2 to 2.4 and back to 2.2mL/L. Between sensor response errors and small differences between the depth of water sampled by the bottle and CTD, it is easy to account for the scatter for this cast. Casts 2 to 12 are also a little out of line in the fits but they are shallow with high DO vertical gradients.

The bottles flagged by the analyst do not look out of line in the comparison, but with so much scatter that does not necessarily mean there was no problem with the values.

Fits are highly dependent on how outliers are excluded. When all casts are used, but the bottles from near the OMZ for casts #14 and 106 excluded plus a few outliers based on residuals, the fit was:

CTD DO Corrected = CTD DO * 1.0163 + 0.041 (1)

If we force the offset found for 2016-47 we get:

CTD DO Corrected = CTD DO * 1.0169 + 0.038 (2)

The R2 values are very similar for both.

Fits were also tried that chose a variety of casts and levels to see if these fits are suitable for most casts and all choices led to similar results.

These fits result in corrections that are smaller than those found for other recent cruises, with differences of 0.003mL/L at 0mL/L, 0.012mL/L at 4mL/L and 0.02mL/L at 7mL/L.. Examination of individual files shows that a lot of the higher DO data came from casts that had a sub-surface DO maximum at depths of 40m to 75m. The complications from sub-surface DO maxima and deep DO minima, plus frequent small reversals may explain differences from other cruises, but it is not obvious what to do about recalibration. The largest error is going to be due to calibration drift and our best guess for that is to look at other cruises. The fit for this cruise is very close to that of 2015-10 when the sensor was freshly calibrated and for 2016-41, but the slope is lower than for 2016-21 and 2016-47.

No fit is going to cope well with these severe gradient changes – having both DO minima and maxima makes a linear fit unreliable. There are errors due to the DO sensor having a slow response and errors due to the vertical distance between the Niskin and CTD and likely some errors due to incomplete flushing of the Niskin bottles, and all those errors are dependent on the DO gradient. We can hope that the errors above and below the maxima and minima largely cancel each other out.

In the absence of a clear “best” choice, fit #1 will be applied.

For more details see document 2016-06-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor. Except for one high value (12.53ug/L) all extracted CHL was <1.3ug/L. For low CHL this type of fluorometer tends to read high as it did in this case. For the 1 high CHL value, fluorescence was about 44% of CHL, which is typical as well.

The plot of FL/CHL vs CHL below shows the usual pattern with CTD fluorescence reading higher than the extracted CHL for CHL<1ug/L and then falling relative to CHL until it is about 50% of CHL at high CHL values.

A fit forced through the origin has a slope of about 1.8 when the one high CHL sample is excluded.

For full details of the comparison see file 2016-06-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on a few casts but the results were hard to judge because the temperature traces were noisy. But a setting of +3s looked reasonable and was the choice for the past 2 cruises that had this equipment. ALIGNCTD was run on all casts using +3s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. During other recent cruises, the default setting of (α = 0.0245, β=9.5) looked better for both conductivity channels than others tested. One cast was checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. The shaded values are from 2 previous cruises that used these sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2015-21-0034 | 1000 | -0.0013 | -0.00014 | -0.0002 | High, Mod |
|  | 1900 | -0.0014 | -0.00006 | +0.0008 | Med, Noisy |
| 2015-21-0068 | 1000 | -0.0013 | -0.00008 | +0.0002 | High, VNoisy |
|  | 1900 | -0.0012 | -0.00004 | +0.0009 | “ |
| 2015-21-0099 | 1000 | -0.0014 | -0.00008 | +0.0004 | High, XNoisy |
|  | 1900 | -0.0013 | -0.00002 | +0.0012 | “ |
| 2016-47-0039 | 1000 | -0.0019 | +0.0001 | +0.0032 | High, V Noisy |
|  | 1900 | -0.0018 | +0.0002 | +0.0040 | “ |
| 2016-47-0069 | 1000 | -0.0019 | +0.0001 | +0.0036 | High, V Noisy |
|  | 1900 | -0.0017 | +0.0001 | +0.0039 | “ |
| 2016-47-0094 | 1000 | -0.0015 | +0.0002 | +0.0039 | High, Mod |
|  | 1900 | -0.0017 | +0.0002 | +0.0045 | “ |
| 2016-47-0105 | 1000 | -0.0017 | +0.00015 | +0.0038 | High, Noisy |
|  | 1900 | -0.0018 | +0.0002 | +0.0045 | “ |
| 2016-06-0029 | 1000 | -0.0021 | +0.0003XN | +0.0045 VN | High, Moderate |
|  | 1900 | -0.0017 | +0.0003 | +0.0050 | “ |
| 2016-06-0045 | 1000 | -0.0018 | +0.0002 | +0.0046 | High, Moderate |
|  | 1900 | -0.0018 | +0.0003 | +0.0054 | “ |
| 2016-06-0068 | 1000 | -0.0016 | +0.0002 | +0.0045 | High, Moderate |
|  | 1900 | -0.0016 | +0.0003 | +0.0052 | “ |
|  | 3800 | -0.0020 | +0.0003 | +0.0059 | “ |
| 2016-06-0079 | 1000 | -0.0016 | +0.0003 | +0.0048 | High, Moderate |
|  | 1900 | -0.0017 | +0.0003 | +0.0055 | “ |

The temperature differences are similar to those from 2016-47. The conductivity and salinity differences are slightly higher but there is no evidence of drift through 2016-06.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. An error in format for the station name in event #5 was found and corrected.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.6db which is reasonable.

The header check showed that there were some negative pressures. Two casts were found with negative pressures. In one case the pressure went slightly negative just for 1 record while the pumps were on; the CTD had obviously been bumping around at the surface. The transmissivity was 0 and salinity low. This suggests that the pressure offset is appropriate. For cast #5 there is a section at the end of the cast when pressure becomes negative just after the pumps are turned off. There are readings at pressure -0.03db and ~0 transmissivity suggesting that the surface is very close to 0db. Then the pressure goes very briefly to 0 then moves to ~-01db with transmissivity high enough to suggest that the CTD was out of water. The salinity was very low. This evidence suggests that the surface is within 0.1db of zero. There were no other anomalies in the header check.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. Water depths were compared with the log book entries. For cast 46 the depth was changed to match the log (acquisition had started early) and for casts 81-84 the depth was changed to that found in the log for events 82 to 84 – all 4 were at P26 and very close in position. Cast #32 had no depth entered so it was added using the log book entry.

For most casts the CTD did not get within 15m of the bottom so there is no altimetry header entry. Where there are entries a calculation of (Pressure*0.99 -Depth +Altimeter Reading) was made and where the result is <5m the entry was considered good. There were a few larger differences so the near-bottom altimetry was examined and found to be ok. Most of these occurred for very deep casts – likely the 0.99*Pressure estimate does not work well at those depths. The altimetry signal looks reliable and the water depth agrees with the log, so no changes were made. There were a number of shallow casts in deep water that had altimetry headers, so those were removed as the CTD was obviously never within 15m of bottom. For cast #58 the depth in the log is greater than that in the header; it is much higher than seen at this site in February and lower than for 2 cruises in 2015, so the header entry was kept as is.

The final merge step and CLEAN run were repeated for the bottle files.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 5 casts to see what shift in conductivity does the best job of removing noise in the salinity channels The best primary shift was -0.4, while -0.6 worked best for the secondary.

SHIFT was run on all casts using -0.4 records for the primary conductivity.

SHIFT was run on all casts using -0.6 records for the secondary conductivity.

Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

The sensors were used during 4 cruises in 2015 and 2 just before this cruise in 2016. The pressure offset was updated based on near-surface sampling on the previous 2 cruises. Salinity calibrations were mostly not trusted due to flushing problems, delayed analysis and/or noisy comparisons for most of the cruises. The secondary salinity has been closer to bottles than the primary for all cruises. There was a good comparison for 2016-47 for which the primary salinity was found to be low by ~0.006, but after a correction to temperature it was low by 0.004. The secondary salinity was low by 0.002.  The dissolved oxygen comparisons have been reasonably consistent.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All temperature and salinity values were within the climatology although surface salinity values at P4 were very close to the minimum.

Repeat Casts – There were repeat casts though the deepest were not close together in time. Casts #82 and #93 occurred about a day apart at P26. Differences along lines of constant σt were about 0.001 for temperature and 0.0002 for salinity at σt= 27.6 (near 1800m) for both T-S pairs. This shows good repeatability.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

An initial examination to decide which channels to archive showed that the primary was noisier with more unstable features.  So the secondary T and S channels were selected for archiving.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. Some bad salinity points were removed. All files required some editing.

After editing T-S plots were examined for all casts. The only unstable features in the T-S plots were for the cast in Haro Strait where such features are expected.

##### Recalibration

There is no evidence to suggest that the pressure channel needs further adjustment.

The primary salinity was found to be lower than bottles by ~0.007 and the secondary by 0.0017. There may be a slight effect on bottle samples due to evaporation and/or adsorption, but it would be small given that analysis was done within a month; allowing for an effect of 0.0005 is reasonable, so that would imply that the primary salinity is low by ~0.0065 and the secondary by ~0.0015. The difference between those estimates is 0.005 which is in reasonable agreement with the differences noted in section 9. If the primary temperature is lowered by 0.002 then the primary salinity will increase, bringing it within ~0.0045 of bottles.

Dissolved oxygen recalibration was discussed in section 5.

File 2016-06-recal1.ccf was prepared to subtract 0.0020 from the primary temperature, to recalculate primary salinity and to apply the following correction to channel Oxygen:Dissolved:SBE:

CTD DO Corrected = CTD DO * 1.0163 + 0.041

A second run of CALIBRATE using file 2016-06-recal2.ccf added 0.0045 to the primary salinity and 0.0015 to the secondary salinity.

These corrections were first applied to the SAM and MRGCLN2 files. COMPARE was rerun and the results confirm that the two salinity channels are within 0.001 of bottles. For dissolved oxygen using roughly the same points as in the fit used for recalibration, the average of differences in the DO fit was +0.0018mL/L and the standard deviation was 0.029mL/L. In the top 250m the SBE DO data are roughly within ±0.08mL/L and below that ±0.05.

(See file 2016-06-DO-comp2.xlsx for details.)

CALIBRATE was then run on the EDT files using 2 runs.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When all data were included the CTD DO was higher than the titrated samples by an average of ~0.02mL/L with a standard deviation of 0.13mL/L. Removing some outliers based on residuals reduced the standard deviation but had little effect on the average difference. The CTD DO is generally higher than the titrated DO samples above the Oxygen Minimum Zone. It is close to bottles near the minimum and reads lower than bottles below that. There are two known sources of error that do not involve calibration drift: slow response of the DO sensor in high gradients and poor flushing of Niskin bottles. Both will lead to the downcast DO from the sensor reading higher than the bottles above the DO minimum (found between 800 and1500m) and lower below the OMZ. The level of the oxygen minimum varies among the casts, being deeper further offshore, and the minimum DO values vary, being higher offshore. The local gradients in the OMZ are lower far offshore and errors due to flushing and response time should be low; for the 2 bottles taken at the minimum at P20 and P26, the CTD DO was higher than the bottles by an average of 0.002mL/L. This offers some evidence that the calibration error is small.

No further recalibration is justified. See 2016-06-dox-comp3.xlsx for details.

##### Fluorescence Processing and special files for Dr. Peña

The COR2 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

The PAR channel was removed from casts 34, 53, 68, 82 and 93.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and unedited except that

some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 100db

±0.3 mL/L from 100db to 500db

±0.1 mL/L from 500db to 1000db

±0.05 mL/L below 1000db

For details on the processing see document: 2016-06_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

A cross-reference list was produced.

The sensor history was updated.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. For casts at P1 to P26 values ranged from 101% to 107%. Surface DO saturation was about 130% in Saanich Inlet, and 80% in Haro Strait and 75% in Juan de Fuca Strait. The values suggest that the DO calibration is reasonably good.

##### Final Bottle Files

The MRGCOR2 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

PAR was

A second SBE DO channel was added for both the CTD DO and titrated DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

One line was removed from cast 2016-06-0023 that was from a misfired bottle with no sampling.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing. Note that format F7.3 was used for the Salinity:Bottle channel as recommended by the salinity analyst.

Data were exported from the CHE files to file 2016-06-bottles-final.xlsx. The entries were compared with the rosette log sheets to ensure no samples had been missed. A few problems were found:

- Oxygen sampling is indicated on the sampling log for events #19, 40, 50, 64, 85, 95 and 97 but no such DO samples were reported; they are listed in groups with other samples taken by non-DFO researchers, so are assumed to not be intended to go into the CHE files.

- There is no CHL sample for event 21 (sample 130); the duplicate samples were analyzed but not found on the final QF page. The analyst was contacted and provided a new spreadsheet. The CHE file was recreated.

- Salinity sample 292 was missing from the final QF page, but was found in the raw data and precision study. The analyst was contacted and provided the missing data. The CHE file was recreated.

- Cast #75 had 2 lines for Niskin #1 – this problem was included in the notes from the Chief Scientist and one line was removed from the bottle file. Niskin #1 had been fired twice.

- Nutrients for event #82, P26 indicate duplicates for most samples and quadruplicates for some. The analysis shows mostly single samples but dups for the cases shown as having 4 samples. So it is assumed that the missing samples were taken for the use of another group.

Standards check and a header check were run on all files. No problems were found.

The track plot looks ok.

Plots of each file were examined. A problem was found in event #93 which has a PAR channel with all values =0 and the depth is >>2000m, so the PAR sensor was obviously not mounted. The steps from REMOVE onwards were repeated for that cast.

A few errors were found in bottle files in the final review process; these were corrected and the bottle files were rebuilt.

A cross-reference listing was produced for the CHE files.

##### Thermosalinograph Data

There were 3 hex files. There were extracted chlorophyll and salinity samples taken while underway and some were taken at about the same time as the 5m bottle was closed during CTD casts. There were also dissolved oxygen samples taken in two different parts of the ship and from a few 5m rosette bottles.

a.) Checking calibrations

There were 2 configuration files which were identical. One file was renamed as 2016-06-tsg.xmlcon. No errors were found in the calibration parameters. There was no configuration file for the first TSG file, but the chief scientist noted that there were no changes to the configuration during the cruise.

b.) Conversion of Files

The 3 files were converted to CNV files using configuration file 2016-06-tsg.con.

Those CNV files were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

A time-series plot showed that the data look good. There are one-sided spikes in the salinity of size 0.5 to 1psu in parts of the record, but they disappear abruptly when the ship stopped at station P26 and were not seen again. File #3 covers the return trip and there are no noteworthy salinity spikes in that file. While the ship is stopped spikes are not seen. The spikes are all towards lower salinity suggesting bubbles.

The flow rate was mostly very steady at either 1 or 0.95 with a single brief drop-out in each of files #1 and 2.

The track plot looks fine. The plot was added to the end of this report.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db and exported to a spreadsheet which was saved as 2016-06-ctd-tsg-comp.xls. There were 47 casts which overlapped with TSG files.

The TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, lab temperature, salinity and fluorescence and the files were reduced to the times of CTD files and loop samples.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude and longitude were all ≤0.0007° and the median differences were 0.0001° and 0.0002°. This shows both the times and positions are reliable for both systems.

d.) Comparison of T, S and Fl from Loop & Rosette Samples and TSG and CTD data

- T1 vs T2 The intake temperature sensor worked throughout the cruise. The median and average differences were both ~0.178C° during stops for CTDs. There was the usual dependence on intake temperature with warming in the loop getting lower as the intake temperature increased and approached the ambient temperature of the ship.

- Flow Rate The flow rate was steady varying from 0.92 to 1.0 with a median of 0.94.

- TSG vs CTD The spreadsheet comparing CTD and TSG files was then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG.

|  | Tintake-Tctd | Tlab-Tctd | SALtsg-SALctd | FLtsg/FLctd |
| --- | --- | --- | --- | --- |
| average | 0.0046 | 0.1832 | -0.0078 | 0.209 |
| min | -0.1224 | 0.0823 | -0.0524 | 0.107 |
| max | 0.0914 | 0.2971 | 0.1824 | 0.372 |
| median | 0.0037 | 0.1838 | -0.0067 | 0.180 |

26 well-mixed casts were identified by finding the depth at which the salinity had increased by 0.005 from the 4m reading. For 2 of those casts the temperature differences were high and examination of profiles showed they were noisy between 0 and 4m so those 2 were not used for any of the variables. For another 3 the salinity differences were >0.03 but temperature seemed ok, so only the salinity values were not included in the average for those.

|  | Tintake-Tctd | Tlab-Tctd | SALtsg-SALctd | FLtsg/FLctd |
| --- | --- | --- | --- | --- |
| average | 0.0028 | 0.1774 | -0.0084 | 0.2130 |
| median | 0.0024 | 0.1806 | -0.0066 | 0.2040 |
| stdev | 0.0055 | 0.0201 | 0.0042 | 0.0697 |

1. Intake Temperature The intake temperature is higher than the CTD temperature by a median of 0.004C° using all casts with a standard deviation of 0.025C°. Using the well-mixed casts the  median difference was 0.0024C° with a standard deviation of 0.006 C°.

2. SALINITY TSG salinity data are lower than the CTD salinity by a median value of 0.0067 with a standard deviation or 0.03. From the well-mixed casts the median was 0.0066, so not much different except that the standard deviation was much lower at 0.004.

3. FLUORESCENCE

The TSG fluorescence data are uncalibrated and expressed in volts. The ratio of the TSG fluorescence to that from the CTD has a median value of 0.18 and standard deviation 0.07. Using the well-mixed casts the results are similar. To make an estimate of TSG fluorescence, multiplying by between 2.6 and 9.3 would be required with a factor of 5 or 6 being a common result offshore.

(See 2016-06-ctd-tsg-loop-comp.xls.)

- Loop vs Rosette

The loop salinity was higher than the rosette salinity by an average of 0.006psu, but when 2 outliers were excluded the loop samples were low by an average of 0.001psu, and all differences were within 0.008psu of the rosette samples. Six of the 9 samples were within 0.002psu of the rosette samples. The rosette samples for the two outliers were in good agreement with the CTD salinity in the CHE files. One of the outliers had a loop sample that was lower than the rosette sample by ~0.03 while the other was higher by 0.09.

The loop extracted chlorophyll was lower than the rosette chlorophyll by an average of 0.11ug/L, but when the samples from the 2 samples taken at P1 and P2 are excluded the loop CHL is lower by an average of 0.03ug/L. P1 and P2 are in areas of high variability so a slight mismatch in time may be significant. The P1 loop salinity sample was somewhat out of line as well and there was no salinity sample from P2.

Overall, the loops are as close to the rosette as we can expect.

- Dissolved Oxygen Loops

There were oxygen samples taken both in the lab and in the transducer compartment. They were sometimes taken at significantly different times.

The DO from the lab was higher than that from the surface rosette by an average of 0.13mL/L, but that goes down to 0.08mL/L when the casts are excluded that had a significant time difference between sampling. The DO values from the transducer compartment were slightly closer, on average, to the rosette DO than those from the lab but there was a lot more scatter. The DO values from the transducer compartment were lower than those from the lab in most cases and when outliers were excluded the transducer DO values were lower than the Lab DO values by an median of ~0.02mL/L. The casts with a significant time difference between the gathering of the 2 loop samples are associated with outliers in 3 out of 4 cases. Based on averages, the, DO appears to increase in the loop but most of the difference from the rosette samples appears to have occurred before reaching the transducer compartment. There are few bottles and an inconsistent method of drawing the DO TR samples, so these conclusions are weak.

- Loop Bottle - TSG Comparisons

TSG data was extracted for the times indicated in the log for loop samples. The differences between the latitude and longitude readings from the log and those from the TSG file are somewhat larger than noted in the comparison with CTD data. The latitudes differ by as much as 0.015º and the longitudes by as much as 0.20º while the ship is underway. While stopped the differences are within 0.003º and 0.004º. The times noted in the log are likely out by a few minutes and if the ship is moving that would lead to larger differences. So we cannot expect the underway comparison to be of the same quality as seen when the ship is stopped though it is hoped that the resulting mismatches in salinity and fluorescence may be non-systematic. A plot of the standard deviation in the TSG intake temperature versus CTD event number was made to see where variability is high; this shows that comparisons from Juan de Fuca and from the return trip will be most affected by small errors in matching time.

There were 7 significant outliers in the comparison between the TSG salinity and loop salinity, one of which was flagged by the analyst and another that came from Loop P14 for which the time is uncertain. There were 3 that came from times when the ship was underway and 2 of those were near-shore where variability was high so that a slight mismatch in temperature would be significant. The two casts from P4 both were both outliers. The standard deviation (measured over 2 minutes – 5 records) in the TSG Salinity was low for those cases. A close analysis suggests that if the 2 P4 samples are reversed, the results look good. The labels are the same, but looking at the RAW data it is seen that the sample with the lower salinity was analyzed with other samples from late in the cruise and the one with the higher salinity was analyzed with samples from around the time of event #14. So the values were reversed in the comparison. So Loop plus station name is not a good labelling method – something more is needed. For event #45 there is some doubt about whether it was taken while stopped at P14 or while underway. The salinity suggests that the loop did not get taken while stopped at P14 or that this is a bad bottle. In light of the doubts about when this loop sample was drawn, it was removed from the comparison.

The P11 loop salinity is lower than the TSG salinity by 0.018psu, lower than the rosette 5m sample by 0.028psu and lower than the CTD salinity when the rosette fired by 0.023psu. There is no note of any problem in analysis. There was no loop sample for P11 for extracted chlorophyll, and there is no indication in the log book or rosette log that this sample was to be taken, so it is likely from some other time/place.

The TSG salinity is lower than salinity samples from the loop by a median value of 0.01 including all bottles. When only those taken while stopped (excluding the dubious P14 value) the TSG salinity is low by a median of 0.008 (std dev 0.008) and while moving it is low by 0.016 (std dev 0.032). The respective averages were -0.006 and -0.009psu. Possible explanations for the differences are:

  - While underway in areas of high variability, small errors in timing could lead to significant errors. The standard deviations in salinity were largest in Juan de Fuca Strait and the differences between TSG and loop salinity samples are greatest there with varying signs.

  - It was noted earlier that there were spikes towards low salinity values in parts of the record. Large spikes (0.05-1.0psu) occurred only while the ship was underway and not at all during large parts of the cruise including the return trip. A check of times for the loop samples shows only 1 case that is likely to be affected by spiking. Event #8 was at JF4 where both high variability and spiking occurred and the ship was moving. This is the cast with the largest difference between TSG and loop (0.075psu). There may have been a little spiking at the beginning of event #9 at P1 but the spikes likely cleared before the sample was taken; the difference there is just slightly higher than the median. It is likely that the underway samples are generally affected by bubbles.

The TSG salinity is lower than loop bottles by 0.008 when stopped and 0.017 when underway. It is likely that the difference between the two values is due to bubbles as even where large spikes in salinity are not noted, there are likely some small bubbles that will not be present in the bottle samples when they are analyzed.

Extracted chlorophyll values were mostly very low and comparisons with a fluorometer will always have a lot of variability at those levels. The loop chlorophyll is higher than the fluorescence voltage by a median factor of ~4 while stopped and ~3 while underway. There are too few bottles to consider this significant, but it may be because the CHL values while stopped were mostly very low and the fluorometer tends to read higher than CHL when CHL is <1ug/L..

(See 2016-06-ctd-TSG-loop-comp.xlsx.)

- Calibration History

The temperature and conductivity sensors were recalibrated in November 2015 and were used during 2016-01 and 2016-47. For 2016-01 the salinity was spiky so comparisons with loops and CTD were not trusted. There were also variations in flow rates that complicated comparisons.

The TSG fluorometer values were about 24% of the CTD fluorometer. Chlorophyll values were low for that cruise. The fluorescence data were archived in volts. For 2016-47 salinity quality was higher but there was too much scatter in comparisons to justify recalibration. Fluorescence was left in voltage and a multiplier of 7.5 was suggested for a rough estimate in concentration units in areas where CHL>0.5ug/L.

Conclusions

1. The TSG clock worked well.

2. The TSG flow rate was steady with values ranging from 0.92 to 1.0 and just 2 short drop-outs.

3. The temperature in the loop increases by roughly 0.18Cº but as usual heating varies with the intake temperature.

4. The TSG intake temperature was higher than the CTD temperature by a median of ~0.0024Cº using 23 well-mixed casts; the standard deviation was 0.0055Cº.

5. The TSG Salinity is lower than the CTD salinity by about 0.008 with a standard deviation of 0.004. Compared to loop samples the TSG salinity is low by 0.01 if all samples are included and by 0.008 while stopped and 0.017 while moving. Large spikes were not noticeable at times when the loop samples were gathered, but when spikes were seen the ship was in motion. There may well be effects due to bubbles that do not lead to obvious spikes. While these appear to be larger when the ship is motion, that does not prove that they are not also present while stopped. So while the comparisons with both CTD and loops are consistent, they may not be due to calibration drift. No matter what the cause it looks reasonable to add 0.008 to the TSG salinity. This may not be enough to correct data in motion, but will reduce any error.

6. The fit of CTD fluorescence in ug/L against TSG fluorescence in volts has too much scatter for a reasonable fit. For a very rough estimate in the offshore region, multiply fluorescence voltage by about 3 while underway and 4 while stopped. The ratio varies with chlorophyll concentration, so this estimate is inappropriate for CHL>1ug/L.

7. The quality of the TSG data is good, with only limited sections corrupted by salinity spikes.

8. Calibration of salinity is justified given that comparisons with both CTD and loop samples give similar results. Adding 0.008psu is appropriate for when the ship is stopped, but most of the TSG record is collected in motion. While there is a lot of scatter in the “moving” comparisons adding 0.0015 looks like a reasonable correction for salinity.  TSG fluorescence will be left in voltage units, but a multiplier of 3 can be applied for a very rough estimate in chlorophyll units for casts west of station P1.

9. Loop 45 appears to have been taken when the ship was moving but we have no record of the time, so it should not be included in the loop file.

f.) Editing

The ATC files were copied to *.EDT.

Each file was opened in CTDEDIT and the salinity channels were cleaned in files #1 and #2 where there were single-point spikes not seen in the temperature trace. File #3 had no large spikes.

In the course of editing it was noted that there were areas where salinity gradually falls and then recovers. This has been seen before, but in this case the drops were smaller than seen in the past. This is assumed to be due to bubbles accumulating and then clearing. These are impossible to edit except by removing large sections of salinity data. They were left in place as it is far from clear which data are bad.

g.) Recalibration

File 2016-06-tsg-recal1.ccf was used to add 0.015 to channel Salinity:T0:C0.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Difference, Conductivity:Primary, Flag and Position:New channels.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats.

Those files were saved as TOB files.

The Standards Check and Header Check were run; no problems were found.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

##### Loop File

The CHE files were put through program DERIVE to obtain sigma-t.

Data from those files (including sigma-T and draw temperature) were exported to file 2016-06-che-surface.csv. Data from below 6m were removed.  (Note that the SBE DO channel for mass units need not be included if there are loop DO samples, since it is easy to do the calculation later for both types; but if there are no DO loop samples as was the case for these data, then include the mass units at this stage.)

Sampling number and sample method columns were added. ROS was entered for the method.

The Start Time was copied into a second column and the first was formatted for date and the second for time. Columns were rearranged to fit a model 6-line header.

Loop data were then added to the file and lined up appropriately. The sampling method column was entered as USW.

Note that the only loop samples from event 59 (P18) were DO samples so that line was dropped since DO samples are not being included.

The lines for events 31 (P11) and 45 (P14) were also excluded due to unknown time of sampling.

The data were sorted on event number, then pressure and added to the 6-line header file.

The file break column was filled with value 1 so all data will be in a single file when converted.

The file was then saved as 2016-06-surface-6linehdr.csv.

(Note: For future reference this step failed several times and no cause could be found. Closing all EXCEL files and trying again eventually worked – the program conversion program seemed to remember an input file that had sinse been fixed.)

A comment file was prepared which was essentially the same as the one used in preparing CHE files but including a description of the loop system.

CONVERT was run to produce an IOS Header file. The flags and comments were entered in the headers in the conversion process.

CLEAN was run to get start and stop times and positions.

Header Edit was used to correct channel names and formats and to add comments. The final file was renamed as 2016-06-surface.loop. A track plot turned up one error in position. That was fixed then the track looks reasonable and a plot of temperature and salinity versus longitude looks reasonable.

Particulars

PAR OFF: 34, 53, 68, 82, 93.

Out of Order firing: 19, 40, 50, 64, 85, 93.

1. PAR 4615 mounted for this cast, then removed.

2. Two drops to 10m and back

13. Possibly on bottom – no altimetry signal

23. Niskin 2 fired accidentally – remove from file. Niskin 3 used for 2m sample.

25. Bottle 1 closed at surface – misfire.

46. Started early to check PAR

50. Niskin 2 closed at wrong depth (250 instead of 24 - so Niskin 3 used for sample #265.

75. Niskin 1 closed a second time after Niskin 24 – remove line for 2nd firing.

95. Wrong station name in file. (Fixed in raw files.)

105. File saved as 105 should be 106. (Fixed in converted files.)-

TSG1 –no con file saved but same as for other files. TSG file 3 has whole return trip.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 17Dec2014 | Factory |  |  |
| Conductivity | 2280 | 19Dec2014 | Factory |  |  |
| Secondary Temp. | 2663 | 15Jan2015 | Factory |  |  |
| Secondary Cond. | 2754 | 19Dec2014 | Factory |  |  |
| Transmissometer | 1185DR | 27Apr016 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 17Jan2015 | Factory |  |  |
| PAR | 70613 | 21Mar2016 | Factory |  |  |
| PAR | 4615 | 16 Mar2011 | Factory |  |  |
| SeaPoint Fluor. | 3640 | n/a |  |  |  |
| Pressure Sensor | 0443 | 17Dec2014 | Factory |  |  |
| Altimeter | 62354 | n/a | Factory |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3411       Cruise ID#:	2016-06

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3411 | 7Nov15 | Factory |  |  |
| Conductivity | 3411 | 7Nov15 | Factory |  |  |
| WetLabs Fluorometer | Ws3s-953p |  |  |  |  |
| Temperature:Secondary | ? | 0842 |  |  |  |
| Flow meter | ? | n/a |  |  |  |