## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 27 Mar 2025 | Updated channel names & formats in TOB files. GG |
| ### 25 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| 16 Aug 2020 | Added HPLC Data. S.H. |

## PROCESSING NOTES

Cruise: 2016-07

Agency: Ocean Sciences Division

Location: Strait of Georgia and Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Strait Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: June 15, 2016 – June 20, 2016

Processed by: Germaine Gatien

Date of Processing: 28 September 2016 – 12 October 2016

Number of original HEX files: 89

Number of CTD files: 89	Number of bottle files: 27

Number of TSG files: 2

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#1119) on the secondary pump, a SeaPoint Fluorometer (#3865) on the primary pump with a 3X cable, a Biospherical QSP-400 PAR sensor (#4565), Surface PAR (#20518) a pH sensor (#0692) and an altimeter (#1253).

The data logging computer was Vector CTD Laptop.

The data acquisition program was Seasave v7.23.1.

The CTD deck unit was an SBE model 11+ s/n 508.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log book included an equipment list but lacked a personnel list. The sampling/rosette log sheets were generally in good order, but 2 sample numbers were used twice during cast #92. A leading 9 was added to the second instance of each of those sample numbers.

Salinity samples were analyzed within 15 to 20 days after collection.  Both salinity channels were within 0.002 of bottles. Overall, the primary salinity data were closer to bottles, but near the mouth of Juan de Fuca Strait, where the CTD was moving more during bottle stops than in other areas, the secondary salinity data were closer to bottles. Since this is likely where flushing of Niskin bottles was best, the secondary salinity was chosen for archiving.

Salinity data were removed from event #96 from 192.9db to 259.3db as they were clearly bad. The primary salinity looks better at those depths but worse above that. The secondary salinity in the upcast file was better at depth but worse above 190db with many unstable features and less detail in the profiles.  So the secondary downcast data were selected for archiving to ensure high quality data, but salinity is only available down to 192db.

Similarly, channel Oxygen:Dissolved:SBE was recalibrated based on results near the mouth of Juan de Fuca Strait. The downcast Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.3 mL/L from 0 to 50db

±0.15 mL/L from 50db to 150db

±0.10 mL/L below 150db

As some of these errors will be due to incomplete flushing of Niskin bottles, this is likely an underestimate of the accuracy of the dissolved oxygen sensor during downcasts.

The fit of CTD fluorescence versus extracted chlorophyll was flatter than usual, with fluorescence averaging about 64% of the extracted chlorophyll. Usually this type of fluorometer reads higher than CHL for very low extracted CHL, but there were few such values.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data were available at the time of processing. Calibration is required for each cast to get absolute values although

general trends within a cast are likely real.

There are 2 TSG files. Data acquisition did not start until 05:17 on June 17th and there was a 6-hour gap between files. There was no loop sampling, flow rate meter or intake temperature sensor. Comparisons with CTD data had too much scatter to enable an estimate of TSG temperature and salinity accuracy.  Temperature and salinity data were removed from one section where flow in the loop appears to have stopped.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The sensor history was found for the pressure, conductivity and DO sensors. This was the 6th use of the pressure sensor and the 3rd use for the conductivity and DO sensors since they were last calibrated at the factory.

The configuration file did not change through the cruise. The calibration constants were checked for all instruments and no problems were found. One file was saved as 2016-07-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2016-07-ctd.xmlcon to create CNV files.

A few casts were examined.  All expected channels are present. The primary and secondary temperature and conductivity channels are close during downcasts, but as usual the upcasts differ more due to noise in both channels. There are some spikes in both primary and secondary conductivity and temperature. Dissolved oxygen, fluorescence, PAR, pH, altimetry and transmissivity profiles all look normal. PAR maximum values were usually lower than Surface PAR, but when at the end of cast #7 the CTD was very close to the surface, PAR looks close (and sometimes higher) than SPAR. The descent rate was mostly fairly high and moderately steady, except near the mouth of Juan de Fuca Strait where it is very noisy.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2016-07-ctd.xmlcon.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and problems were seen in 2 files.

CTDEDIT was used to edit the BOT files for events #19 and #100. Comments about editing were added to the headers. The output files were then copied to BOT.

A preliminary header check was done and no problems were found. Fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records. There were 3 cases of sample numbers being repeated. The surface sample from event #77 was change from 241 to 9241 since 241 was used again on the next cast; since only particles were gathered from the first bottle with that number, this is unlikely to cause a problem. There were 3 different sorts of samples from the second instance, so it would require corrections to a variety of analyses. The other two cases had 2 numbers used twice during cast #92. The dissolved oxygen and nutrient analysts noted this and renamed the second instance with a leading 9, so the order of sample #s is 294, 295, 296, 9295, 9296, 297, etc.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine. Those files were bin averaged on bottle number.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2016-07-bot-hdr.txt which will be updated as needed during processing. The nutrient comments were amended to remove the reference to deep sampling since there were no casts deeper than 400m.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2016-07chl*.xlsx. The file included comments, flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2016-07chl.csv, event numbers were added, and the file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2016-07oxy.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2016-07oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2016-07SAL.xlsx. The analysis was done within 15-20 days of collection. The files were simplified and saved as 2016-07sal.csv.

That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2016-07nuts.xlsx.

Then the file was simplified, reordered on sample numbers and saved as 2016-07-nuts.csv.

The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only, with output files named MRGCLN1..

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number as the merge channel.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. It was found that the oxygen from event 65 had been mislabelled as from event 64. That was fixed, the merges rerun and no further problems were found.

CLEAN was run on the MRG files to add 0 flags to empty flag channels and to update header limits.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There is a lot of scatter in the comparison with bottles. When data with a standard deviation in the CTD salinity >0.001 and differences >0.01 are excluded the 50m bottle from Baynes Sound stands out as an outlier. This is not unexpected since Baynes Sound is an area where flushing of bottles is likely to be less efficient and the vertical salinity gradient is high enough that only a slight flushing problem could easily explain why the CTD salinity is lower than the bottle by -0.1psu. When that case is also excluded the primary salinity is found to be lower than bottles by an average of 0.0003 and the secondary is lower by 0.0013. The standard deviation is 0.002 in both cases. The difference between the primary and secondary is 0.001 which matches the observations of section 9.

When differences are plotted against time a different pattern emerges. The 3 bottles closest to the mouth of Juan de Fuca Strait show the primary salinity to be higher than the bottles by an average of 0.0016psu while the secondary salinity was extremely close to the bottles; the standard deviations were <0.0001. This is the area where we would anticipate good flushing of bottles because there was a lot of vertical motion during bottle stops.

There could be a very small effect on bottle samples due to evaporation and/or adsorption and it might affect the bottles collected earlier more. But that would tend to make the CTD look lower than it really is, not higher, so it does not explain the different results in Juan de Fuca.

All of the outliers are explained by either noisy CTD salinity data or because they are from levels where strong gradients and incomplete flushing are expected.

For full details for the COMPARE run see file 2016-07-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

When all casts were selected, excluding cases where the differences were <-0.4 or >+0.02 and residuals >0.03the fit found was:

CTD DO Corrected = CTD DO * 1.023 	(R2 = 0.33)

The major outliers are all from near the surface except for the sample from 100m in Saanich Inlet which is always an outlier in such comparisons due to large vertical gradients.

When only the 4 casts closest to the mouth of Juan de Fuca Strait were used, and a few outliers removed based on standard deviation in the CTD DO and residuals, the fit was:

CTD DO Corrected = CTD DO * 1.0165	 (R2 = 0.8)

When only the cast at the mouth was used and a few outliers excluded using the same method, the slope was 1.0176 and R2 = 0.86.

It makes sense that the slope is higher where descent rates are noisy enough to ensure the Niskin bottles flush well. Poor flushing means that the bottle samples will have lower DO values than the in situ values, thus making the CTD DO values look closer to ambient values than they really are.

During 2016-05 there were also problems that appeared to be due to poor flushing, so the fit used for recalibration was made using only 2 casts near the mouth of Juan de Fuca Strait. The fit was:

CTD DO Corrected = CTD DO * 1.0152	    (R2 = 0.74)

Multiplying all DO values by 1.0165 looks like the best estimate we can make.

No outliers need flags.

For more detail see document 2016-07-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No problems were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor. There was a large range of CHL values with 77 samples between ~0.5ug/L and ~10.4ug/L. The fit is:

SBE Fluorescence = 0.64 * CHL

The fit looks less dependent on CHL than usual, perhaps because there were few very low CHL values. At very low CHL the fluorescence tends to be higher than CHL, but there are few instances of that in these data.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

A large spike in conductivity in event #5 was removed by this step.

##### ALIGN DO

Tests were run on ALIGNCTD to determine the best setting to advance the DO signal. When this sensor was used in April the best results were with a setting of +2.5s. That setting worked well on these data as well.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. Tests on previous cruises using these sensors showed the default setting of (α = 0.0245, β=9.5) did the best job and it does improve the data for both conductivity channels for these data.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 2 of the deeper casts to examine differences between sensor pairs. None are very deep. The shaded entries are from the last time these sensors were used.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2016-05-0060 | 350 | -0.0002 | +0.0001 N | +0.0013 | High, Steady |
| 2016-05-0085 | 350 | -0.0001 | +0.0001 N | +0.0012 | High, Steady |
| 2016-05-0101 | 330 | -0.0002 | +0.0001 N | +0.0013 | High, F.Steady |
| 2016-07-0069 | 350 | -0.0005 | -0.0001 | -0.0010 | High, F. Steady |
| 2016-07-0108 | 350 | -0.0002 | -0.0001 | -0.0011 | High, F. Steady |

The differences are small and similar to those seen in April though the conductivity and salinity differences have changed signs.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. No errors were found.

A header check was run. No problems were found. No off-scale fluorescence values were found.

The cruise track was plotted and added to the end of this report.

There are a few sources of information about the pressure accuracy:

- Surface Check gave the average surface pressure as 2.5db which is within the normal range for the Vector.

- A deck pressure measurement of 0.7db was made during cast #82 which suggests pressure is too high.

- During cast #7 the CTD ran with pumps on near the surface while the CTD tech worked with a new winch operator on recovery of the rosette. The sensors appear to have been in water when the pressure was 0.3db and maybe right at the surface or out of water for 0.2db; the pumps were on and transmissivity suggests that the CTD might have moved in and out of water, but was mostly in the water. This suggests that pressure is within 0.2db.

- At the end of event #19 the CTD is clearly in water at 0.2db with pumps on. The pressure goes to close to 0 for a few records with no sign of being out of water, but that might not have been long enough to affect values. This suggests pressure is within 0.2db.

- During 2016-46 which followed this cruise there was some evidence that pressure was slightly too low, but probably ok to 0.2db.

The mismatch between the deck pressure of +0.7db and the fact that the CTD clearly was in water at 0.3db during cast #7, may indicate that the deck measurement was taken without sufficient time to equilibrate. Or it may be indicative of the inherent error in the pressure sensor. Both this cruise and 2016-46 suggest small errors, but of opposite sign. Most evidence suggests that the pressure is within ±0.2db which is well within the specifications.

The altimeter and bottom depth headers were exported from the SAMAVG and CLN files to spreadsheets. The bottom depths were checked against the log book. There were many discrepancies; mostly the log entries had been changed so they may be a better reflection of the depth when the CTD was at the bottom.

The following calculation was made to see if the bottom depths and altimetry readings are good:

Check=maximum depth sampled (Max press*0.99) - water depth +altimeter reading

We don’t expect a 0 reading since the altimetry reading is an average over 2db and the bottom depth is subject to changes during the cast. Where bottom depths in the headers differed from those in the log, this calculation was used to see if changing the bottom depth brought the Check value closer to 0. In most cases it did, so the water depth was changed. Where it didn’t a plot was made to ensure the altimetry looked ok. Cases where the log and header depths agreed but the Check value was >4m were plotted to ensure the altimetry header made sense. The only case where the altimetry header looked bad was for cast #1 during which the CTD did not get close to the bottom; a spike in altimetry was misinterpreted as a bottom value.

The following changes were made:

- Bottom depth changed to match the log entry: 7, 28, 40, 51, 60, 62, 65, 66, 68, 69, 86, 89, 90, 93, 106, 107, 110 and 112

- Altimetry header was removed for: 1

The only one of these changes that applied to the bottle files was a minor change to depth for event #28. This was made in the SAM, SAMAVG, MRG and MRGCLN2 files.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and as usual, there is a lot of variability with the up and downcast traces sometimes closer than temperature and sometimes further apart. This is likely due to varying vertical gradients and descent/ascent speeds. Overall the choice made earlier looks appropriate, so no further alignment will be applied.

Conductivity

For 2 recent cruises using the same equipment the best results varied from cast to cast and feature to feature, but overall -0.4 records for the primary and -0.5 records for the secondary looked best. Tests were run on 5 casts and the best results were with -0.5 records for the primary and -0.8 for the secondary conductivity.

SHIFT was run twice on all casts using -0.5 records for the primary and -0.8 for the secondary conductivity.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. The last 2 times this sensor was used a setting of +70 records had the best results and it looks ok for this cruise. As usual the best choice varies with descent rates and local gradients, but a shift of +70 records definitely improves all the data.

SHIFT was run on the pH:SBE channel using a setting of +70 records.

Plots were examined after this step and the results look ok.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

The temperature, conductivity and dissolved oxygen sensors were used during 2 cruises in April 2016. The first cast had delayed salinity bottle analysis and flushing of Niskin bottles was likely poor during both cruises. While the CTD salinity looked low compared to bottles, it is likely because the bottle salinity was too high. Dissolved oxygen comparisons also showed signs of poor flushing. Recalibration was based on results from an area where more ship motion likely improved flushing.  No recalibration was found necessary for pressure.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. As has been seen in all Strait of Georgia/ Juan de Fuca Strait cruises over the past year or two, temperature is frequently above the top of the range. For this cruise this was seen at mid-depths (~100-200m) at all casts in the northern Strait of Georgia, west of 124ºW. No other temperature data was outside the climatology. Salinity is mostly within the climatology range except for frequent excursions near the bottom – sometimes to the low side (1 cast at the mouth of Juan de Fuca and 1 in Baynes Sound) and many to the high side between casts 36 and 46, south of Boundary Bay in the Strait of Georgia. The 3-standard deviation climatology is not recommended for near-shore work. This check is usually done to check for systematic errors in calibration. The salinity excursions are not systematic and the temperature excursions have been seen from a variety of sensors over the past year.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

There is little to distinguish between the primary and secondary T and S channels. Both have spikes though the secondary looks slightly better in T-S space. The comparison with bottles does not make it clear which salinity is more accurate. The secondary was selected for editing and archiving.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes. For a few casts there were sections of bad salinity points where temperature was ok; in those cases only salinity points were removed. All files required some editing, but mostly it was removal of records from near the top and bottom of casts and light cleaning of salinity.

Cast #96 had an offset in the secondary salinity from 192.9db to 259.3db, so those data were removed. The primary salinity looks better at those depths but worse above that; attempts to edit the secondary channels proved unsuccessful. The secondary salinity in the upcast file was examined but, again, while it would provide better data at depth, it would be worse above 190db with many unstable features and less detail in the profiles.  So the secondary downcast data were selected for archiving to ensure high quality data, but salinity is only available down to 192db.

##### Initial Recalibration

Pressure does not need recalibration.

There has been a change in salinity since the last use during 2016-05 when the primary salinity was lower than the secondary. A relative change of ~0.002 has occurred. The bottle comparison for 2016-05 was not trusted due to a delay in salinity analysis and some problems with bottle flushing, so we cannot establish which sensor is drifting. The primary salinity data look best using data from the whole cruise while the secondary look best in the area considered most reliable for comparison. The sensors have been used on 3 cruises since last calibrated at the factory and this cruise provides the best comparison available. Any of the comparisons lead us to suspect that the salinity is good to ±0.002, so recalibration will not be applied.

Based on the comparison of section 5, CALIBRATE was run using file 2016-07-recal1.ccf to correct the Oxygen:Dissolved:SBE channel in the SAM and MRGCLN2 files using:

CTD DO Corrected = CTD DO * 1.0168	 (R2 = 0.77)

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly. The average of differences in the fit once outliers were removed was +0.0167mL/L and the standard deviation is 0.016mL/L. The differences are higher for high DO which generally occurs where gradients are higher, so this is what would be expected if flushing of bottles is incomplete. DO is high by about 0.004mL/L at DO=2mL/L and by 0.034mL/L at DO=7mL/L.

Looking only at the bottles from the 3 casts nearest the mouth of Juan de Fuca Strait, DO is high by an average of 0.006mL/L.

See file 2016-07-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed based on residuals, the CTD DO was higher than the bottles by an average of ~0.07mL/L and standard deviation of 0.05mL/L. The results vary with depth with the CTD reading high by about 0.09 at the surface and by 0.03mL/L at 300m. This pattern is likely due to incomplete flushing in the presence of significant DO gradients above 200m and slow response of the CTD DO in the presence of frequent reversals of DO gradient. No further recalibration of DO is justified.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

The CTD files from rosette casts were clipped to 50m; sigma-T was derived and the data were exported to a single file, 2016-07-SOG.csv.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have a few small unstable features, but those are from sites where active mixing is expected.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts except #45 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, PAR and PAR:Reference data are nominal and unedited

except that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.3 mL/L from 0 to 50db

±0.15 mL/L from 50db to 150db

±0.10 mL/L below 150db

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration

data were available at the time of processing.

Calibration is required for each cast to get absolute values, although general

trends within a cast are likely real.

For details on the processing see document: 2016-07_Processing_Report.doc.

The Standards Check routine was run and a few problems in formats were found and corrected.

The Header Check was run and no problems were found.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values were highly variable ranging from 80% to 130%. As usual values were quite low in Haro Strait, at some casts in the eastern Juan de Fuca Strait and the southern Gulf Islands region and a few of the casts in narrow channels to the north-west. In the open Strait of Georgia values ranged widely in the south and were mostly between 110% and 115% to the north. At the mouth of Juan de Fuca the most southerly cast was at 110% while the other 2 were at 90% and 95%; we normally see values close to 105% in this area. Given the wide range in surface saturation, these observations are not helpful in assessing the DO sensor calibration except to say that are neither consistently high nor low.

##### Final Bottle Files

CALIBRATE was run on all files and then the MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Plots were made of all casts to look for problems and it was found that the temperature and salinity looked odd near the surface for event #46, with salinity significantly lower at 5m than above and below that level. Temperature was also higher at 5m than at 2m. The data were left in place, but a note was added to the header to indicate that the CTD data for bottle #11 are out of line in profile. Temperature rose and salinity fell throughout the stop for the bottle. The changes continued as the CTD was raised from 5m to 2m but they reversed direction shortly before the CTD stopped for the next bottle. There is no evidence that this was due to instrumental malfunction.

Standards check and a header check were run on all files and no errors were found.

The track plot looks ok.

Data from the CHE files were exported to a spreadsheet and compared with rosette sheets and no errors were found.

Plots of each file were examined. The transmissivity looks extremely low at the bottom of casts #36, 40 and 51, so the full files were examined to see if the CTD might have hit bottom. However, the transmissivity reached those low values while the CTD was still descending and the altimetry does not suggest that a bottom touch occurred, or if it did, that it is the reason for low transmissivity. There casts are in the southern part of the Strait of Georgia. No other suspicious features were noted.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

##### Thermosalinograph Data

There were no loop samples, flow meter or intake thermistor.

The only method to check calibration is to compare with the CTD casts.

a.) Checking calibrations

There were two files. The configuration files were identical.

One configuration file was renamed as 2016-07-tsg.xmlcon. No errors were found in the calibration parameters.

b.) Conversion of Files

The hex files were converted using file 2016-07-tsg.xmlcon.

They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots were produced. Salinity has a lot of small-scale variability but that is reasonable for the region. There is a large single-point spike in temperature to 19.7ºC at about 20:30 UTC on June 18th. There is a CTD cast at about that time that has a maximum temperature of ~16ºC at 2db. There is no obvious spike in salinity – in fact, both temperature and salinity are very smooth before the spike.

The track plots mostly look fine, though there are no data from the early part of the cruise and there is a gap between the two files.

c.)  Checking Time Channel

A quick check was made to ensure that the early gap was not due to a clock problem, but the starting time and position are close to the time and position of a CTD cast at station ADCP.

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.3db of 3db because there are not sufficient data at 2db. These were exported to a spreadsheet which was saved as 2016-07-ctd-tsg-comp.xls. There were 61 casts which overlapped with TSG files. There were no TSG data from Events #1-23 and #83-89.

The 2 TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for lab temperature and salinity and the files were reduced to the times of CTD files.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. There were no differences >0.0004º and the median differences were <0.0001º.

An initial comparison showed large differences between the TSG and CTD temperature and salinity data, so a second set of data were selected from 2m, though only 21 casts had data from that shallow. The data were added to the original spreadsheet file on a separate sheet.

d.) Comparison of T and S from TSG and CTD data

- TSG vs CTD The spreadsheet comparing 3m CTD and TSG files was examined first. The differences between the salinity and temperature channels from the CTD and the TSG varied greatly with particularly large differences in the southern part of the Strait of Georgia, and smaller differences in the eastern part of Juan de Fuca Strait and around the Gulf Islands and in the Discovery Islands area. The sites with smaller differences tended to have lower vertical salinity gradients.

1. LAB TEMP The lab temperature was higher than the CTD temperature by a median value of 0.3C° (standard deviation is 0.7C°). When only the 10 casts with the lowest standard deviation in the TSG temperature were included the TSG was high by 0.37C°. If only the first 10 casts in Juan de Fuca Strait are included temperature is high by a median value of 0.11C° (std dev 0.04).

2. SALINITY The TSG salinity data are lower than the CTD salinity by a median value of 0.62psu (standard deviation 3.0) using all casts and low by 0.03psu if only the 10 casts with the lowest standard deviation in the TSG salinity. When only the first 10 events are included the salinity is high by a median of 0.001psu (std dev 0.02).

Next, the CTD 2m data were considered. There are 21 casts with data from that level. A table was prepared with the 2m and 3m data for those 21 casts and the corresponding TSG data.

Using all 21 casts the temperature difference is smaller if the 2m data are used, while the salinity difference is smaller if the 3m data are used. In any case the standard deviations in the differences and in the TSG data are so high that no weight can be put on the results. The first 3 casts are from the Juan de Fuca area that was earlier discovered to have low standard deviations in the TSG and in the differences and low near-surface vertical gradients in salinity. Using just those 3 casts, the TSG temperature is high by 0.10C° (std dev 0.06) using 3m and 0.2C° (stdev 0.50) using 2m. For salinity values are low by a median of 0.002psu using 3m (std dev 0.03) and low by 0.10psu (std dev 0.20) using 2m.

The results are too scattered to conclude much, but the 3m CTD data appears to be the better match. This comparison is complicated by the fact that there are frequent reversals in temperature and salinity near the surface.

(See 2016-07-ctd-tsg-comp.xls.)

- Calibration History

The TSG temperature and conductivity were recalibrated in December 2015 and this is the first cruise known to have used it since then.

Conclusions

1. The TSG clock worked well.

2. The TSG lab temperature was higher than the CTD temperature by between 0.1Cº and 0.2Cº depending on how data were selected. We have little experience with the Vector loop system and no knowledge of the flow rate during the cruise. From the Tully we might expect heating of <0.15Cº given similar intake temperatures, but the Vector loop is shorter so that would reduce the amount of heating. If the ship’s ambient temperature was particularly high that would increase it. The difference between TSG and CTD temperature could also be due to water in the loop coming from slightly higher than 3m and/or calibration drift.

3. The difference between the TSG Salinity and that from the CTD is highly variable. The TSG salinity is close to the CTD values for the Juan de Fuca casts that were well-mixed at the surface.

4. Overall the quality of the data was good with just one section that looks suspicious including a large spike in temperature after a section with little variability.

5. There is insufficient evidence to justify making an estimate of intake temperature.

f.) Editing

CTDEDIT was used to remove a section of data (scans 4485 to 4697) where temperature rose very slowly and steadily and salinity varied little followed by an unbelievably large spike in temperature. It is likely that the flow had stopped in the loop and the spike occurred just as the flow started again. Salinity was cleaned very lightly elsewhere in the first file. No editing was needed for the second file.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Conductivity:Primary, Flag and Position:New channels.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats. The Temperature:Primary was renamed as Temperature:Lab to make it clear that it is not the intake temperature.

Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

Particulars (mostly notes from log)

7. File ran long as they pulled the rosette out of the water.

19. Reduced oxygen sampling to allow analyst to catch up after equipment problem.

75. Fluorescence max at about 20m.

82. Deck pressure 0.7db.

92. Sample numbers 295 and 296 were used twice; the second instances were named 9295 and 9296.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 18Dec2014 | Factory |  |  |
| Conductivity | 1766 | 20Aug2014 | Factory |  |  |
| Secondary Temp. | 4054 | 20Aug2014 | Factory |  |  |
| Secondary Cond. | 3531 | 6Feb2013 | Factory |  |  |
| Transmissometer | 1396DR | 2Feb2016 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 19Dec2015 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| Surface PAR | 20518 | 21Mar2016 |  |  |  |
| pH | 0692 | 12Jan2016 | Factory |  |  |
| SeaPoint Fluor. | 3865 | n/a |  |  |  |
| Pressure Sensor | 0550 | 17Dec2014 | Factory |  |  |
| Altimeter | 1253 | n/a |  |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3363       Cruise ID#:	2016-07

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3363 | 17Dec15 | Factory |  |  |
| Conductivity | 3363 | 17Dec15 | Factory |  |  |