## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 27 Mar 2025 | Updated channel names & formats in TOB files. GG |
| ### 25 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| 16 Aug 2020 | Added HPLC Data. S.H. |

## PROCESSING NOTES

Cruise: 2016-10

Agency: Ocean Sciences Division

Location: Strait of Georgia and Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Strait Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: September 22, 2016 – September 26, 2016

Processed by: Germaine Gatien

Date of Processing: 3 January 2017 – 11 January 2017

Number of original HEX files: 85

Number of CTD files: 84	Number of bottle files: 27

Number of TSG files: 2

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#1119) on the secondary pump, a SeaPoint Fluorometer (#3865) on the primary pump with a 3X cable, a Biospherical QSP-400 PAR sensor (#4565), Surface PAR (#20518) a pH sensor (#0692) and an altimeter (#1252).

The data acquisition program was Seasave v7.26.2.13.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log book contained no equipment list and no personnel list. While the latter can be found in the cruise report, the former cannot. There were also no photos. Since the serial numbers of external sensors are not recorded in the hex files, there is no way to confirm that the calibration parameters entered for those sensors are correct. It was established that the same equipment was used for a Vector cruise in late June 2016, but external sensors could have been switched in the intervening 3 months. There is also no way to determine on which pump the dissolved oxygen sensor and fluorometer were mounted. When there is a problem with a pump this information is helpful in establishing what happened. The entries for events were clear with good notes in the log book about problems encountered.

The sampling/rosette log sheets were generally in good order except that the event numbers were entered incorrectly for events #2 and 14.

Due to the Autosal being out of service, salinity analysis was delayed so that samples sat for 57 to 62 days after collection. This delay is expected to lead to increased values due to evaporation and adsorption of samples. The descent rate of the CTD was steady for many of the casts which may mean the Niskin bottles did not flush well; this would lead to significant differences between CTD and bottle salinity. So it is no surprise that there is a lot of scatter in the comparison of CTD and bottle salinity. When outliers were removed until a fairly flat fit of differences versus pressure is achieved, it is found that both salinity channels are low by about 0.003. Since evaporation and adsorption are thought to raise bottle salinity by about 0.004psu in 2 months the CTD salinity is likely reading slightly high. No recalibration was applied to salinity.

Channel Oxygen:Dissolved:SBE was recalibrated based on results near the mouth of Juan de Fuca Strait. The downcast Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.3 mL/L from 0 to 50db

±0.15 mL/L from 50db to 150db

±0.06 mL/L below 150db

As some of these errors will be due to incomplete flushing of Niskin bottles, this is likely an underestimate of the accuracy of the dissolved oxygen sensor during downcasts.

The fit of CTD fluorescence versus extracted chlorophyll showed the usual pattern with the ratio of CTD Fluorescence to Extracted CHL samples starting at about 1.2 for the lowest CHL and gradually falling to about 0.4 or 0.5 as CHL approached the maximum of 9.2ug/L.

Channel pH:SBE was removed because the sensor malfunctioned throughout the cruise.

There are 2 TSG files. Data acquisition did not start until 17:53 UTC on September 23rd. The Vector TSG set-up has no loop sampling, flow rate meter or intake temperature sensor. Comparisons with CTD data had too much scatter to enable an estimate of TSG temperature and salinity accuracy.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The sensor history was found for the pressure, conductivity and DO sensors. This was the 7th use of the pressure sensor and the 5th use for the conductivity and DO sensors since they were last calibrated at the factory.

The configuration file did not change through the cruise. The calibration constants were checked for all instruments and no problems were found. One file was saved as 2016-10-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2016-10-ctd.xmlcon to create CNV files.

A few casts were examined.  All expected channels are present. The primary and secondary temperature and conductivity channels are close during downcasts, but as usual the upcasts differ more due to noise in both channels. There are some spikes in both primary and secondary conductivity and temperature. Dissolved oxygen, fluorescence, PAR, SPAR, altimetry and transmissivity profiles all look normal. PAR maximum values were usually lower than Surface PAR, but the PAR data is not from right at the surface. The descent rate was mostly fairly high and moderately steady, except near the mouth of Juan de Fuca Strait where it is noisy.

The pH signal looks odd for at least some casts with sudden jumps in values at depth. Problems were noted with this sensor during 2016-45. This will be investigated later.

##### BOTTLE FILE PREPARATION

ROS files were created using file 2016-10-ctd.xmlcon.

The hex file for event #42 was named 2016-10-0042hex.ros. The converted file names were changed to standard format.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and problems were seen in a few files.

CTDEDIT was used to edit the BOT files for events #14, 82, 95 and 100. Comments about editing were added to the headers. The output files were then copied to BOT.

Because the surface bottle was accidently missed during event #38, a second cast was run to fire that bottle and it was saved as event 39. The BOT files for the two events files were renamed as 2016-10-0038.BOTx and 2016-10-0038.BOTy. The bottle number was changed to 38 for file BOTy and then the two segments were joined as 2016-10-0038.BOT.

A preliminary header check was done. The only problem found is that some data in event #2 came from a time when the pumps were turned off.  However, there was no sampling from that bottle, so losing data is not a problem. All the data for bottle #10 was removed from the BOT file using a text editor.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records. A number of complications were encountered:

- Event 2 was called #1 on the rosette log sheet and most analysts used that ID.

- Event 14 was called #12 on the rosette log sheet and most analysts used that ID.

- Event #42 had no sampling, so will not be processed further.

- During event #89 Niskin #9 was fired 3 times though the rosette log sheet notes that it “did not fire”. No sampling was done, so those lines were removed from the addsamp file.

- As noted in the log there was a problem with Niskin #11 from event #82 to the end. A sample number had already been assigned to that bottle for casts #85 to #100. The sample numbers were left as planned, so the sample #s intended for Niskin #11 were removed from the addsamp file.

The analysts were informed about the event number errors.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine. Those files were bin averaged on bottle number.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2016-10-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2016-10chl*.xlsx. The file included comments, flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2016-10chl.csv, event numbers were added, and the file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2016-10oxy.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2016-10oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2016-10SAL.xlsx. The analysis was done within 57-62 days of collection. The files were simplified and saved as 2016-10sal.csv.

That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2016-10nuts.xlsx.

Then the file was simplified, reordered on sample numbers and saved as 2016-10-nuts.csv.

The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only, with output files named MRGCLN1.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number as the merge channel.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. A few problems were found and fixed.

CLEAN was run on the MRG files to add 0 flags to empty flag channels and to update header limits.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There is a lot of scatter in the comparison with bottles. The near-surface data look particularly unreliable with the CTD being lower than bottles by up to 0.07psu. So data above 150db were excluded. Then bottles were removed with large differences until a flat fit was achieved. The average of the 8 bottles shows the primary to be low by 0.0027 and the secondary low by 0.0032psu. Since there was a long delay in analysis we expect bottles to be reading high by roughly 0.004, so it is likely that the CTD is reading slightly high. Evaporation of sample and flushing errors are random which will account for much of the scatter. From 2 casts near the mouth of Juan de Fuca Strait where flushing is expected to be good, the CTD data looks low by about 0.002psu and 0.0026psu. Those were fairly shallow casts. One of the Juan de Fuca casts was way out of line.

The most significant outlier is associated with a very high vertical salinity gradient. The CTD salinity was steady around the time of bottle firing, but values close to that of the bottle were seen within 1m of the stopping depth. Many outliers are associated with noisy CTD data. Evaporation and adsorption of samples likely account for most of the difference when large outliers are removed. Poor flushing in the presence of large vertical salinity gradients likely explains the major outliers. No quality flags are recommended as none of the problems are clearly related to analysis or collection. No recalibration is justified as the CTD salinity is likely within ±0.002psu.

There were no duplicates. In the absence of a precision study and given delayed analysis and a lot of scatter in the comparison, the bottle salinity flagged 3 and comments added.

For full details for the COMPARE run see file 2016-10-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The cast from Saanich Inlet was excluded from the comparison (as usual) because most bottles from that site do not fit the rest of the cruise because of the extreme gradients. Some other bottles were excluded due to high standard deviations in CTD DO and some as outliers based on residuals. There were no CTD DO values <1.4mL/L except in Saanich Inlet, so the offset was forced to 0. The fit found was:

CTD DO Corrected = CTD DO * 1.0161

The correction is fairly small. This may be because of poor flushing of Niskin bottles, but the vertical gradients were quite low for many casts, so flushing errors should be fairly small. However, flushing is best near the mouth of Juan de Fuca Strait, so a fit was done of the 2 casts nearest the mouth. There were a number of outliers confusing the comparison but the fit found was:

CTD DO Corrected = CTD DO * 1.0205

Including 4 casts from Juan de Fuca produced a similar result.

CTD DO Corrected = CTD DO * 1.0199

The deep casts in the central Strait of Georgia had very steady descent rates where we might expect poor flushing. A comparison showed that near-surface samples were poorly flushed with the CTD DO reading higher than bottles. When only data below 25m was included the fit was:

CTD DO Corrected = CTD DO * 1.0176

Overall the Juan de Fuca fit seems most appropriate for recalibration, so CTD DO values will be multiplied by 1.02.

Most of the significant outliers are explained by high standard deviation in the CTD dissolved oxygen channel which is generally due to high local gradients so that there is a mismatch between what the bottles contain and what the CTD is measuring. A few others involve bottle samples that had been flagged by the analyst. There remain a few others that are mostly near the surface. For several the difference could be explained by the bottles not flushing completely after a shed wake passed by the CTD during the stop. Cases that were not readily explained were all from the top 20m:

- Event #17; ~20m; sample looks more like 10m than 20m but there is a lot of variability and this is not a major outlier.

- Event #51; ~2m; probably just poorly flushed bottle, difference not far out of line.

- Event #89; ~2m; high enough local gradient that distance between CTD and bottle explains some of the difference so it would not be far out of line.

- Event #92; ~10m; could be a mis-sample as it looks more like a 5m CTD value (after recalibration). The 20m and 5m samples are closer to the 10m and 2m CTD values as well, but the differences are not so large that these stand out as outliers. This was a cast where one bottle was not sampled (Niskin #11) so it seems possible that instead of skipping 1 Niskin, maybe 2 were skipped. The DO analyst decided to assign flag 3 to samples 301 to 303.

| Sample # | Pressure | CTD DO-recalibrated | Bottle DO |
| --- | --- | --- | --- |
| 301 | 20.2 | 3.334 | 3.549 |
| 302 | 10.0 | 3.535 | 3.938 |
| 303 | 5.2 | 3.909 | 4.010 |
| 304 | 2.2 | 4.011 | -99.000 |
| 305 | 2.6 | 4.016 | -99.000 |

The DO analyst asked for an opinion on samples 45 and 46 as there was some question that they might have been mixed up, but the comparison shows a good correspondence with CTD DO for both.

For more detail see document 2016-10-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further problems were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor. There were 80 samples with a range of ~0.2ug/L to ~9.2ug/L. The fit is:

SBE Fluorescence = 0.75 * CHL

As usual, for low CHL the fluorescence tends to be higher than CHL, and for CHL>5ug/L the fluorometer reads low.  There is a large scatter that may be partly due to incomplete flushing of Niskin bottles. For details see file 2016-10-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

A large spike in conductivity in event #5 was removed by this step.

##### ALIGN DO

Tests were run on ALIGNCTD to determine the best setting to advance the DO signal. When this sensor was used in June the best results were +2.5s for one cruise and +1.5s for the other. Tests were run using a variety of settings and the best results were with 2.0s.

ALIGNCTD was run on all casts using +2.0s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. The default setting of was selected and it does improve the data for both conductivity channels for these data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 3 of the deeper casts to calculate differences between sensor pairs. None of the casts are very deep. The shaded entries are from 2 other recent cruises during which these sensors were used.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2016-05-0060 | 350 | -0.0002 | +0.0001 N | +0.0013 | High, Steady |
| 2016-05-0085 | 350 | -0.0001 | +0.0001 N | +0.0012 | High, Steady |
| 2016-05-0101 | 330 | -0.0002 | +0.0001 N | +0.0013 | High, F.Steady |
| 2016-07-0069 | 350 | -0.0005 | -0.0001 | -0.0010 | High, F. Steady |
| 2016-07-0108 | 350 | -0.0002 | -0.0001 | -0.0011 | High, F. Steady |
| 2016-10-0022 | 351 | -0.0003 | -0.00005 | -0.0003 | High F.Steady |
| 2016-10-0044 | 343 | -0.0002 | -0.00007 | -0.0004 | High, steady |
| 2016-10-0052 | 275 | -0.0004 | -0.0007 | -0.0003 | High, steady |

The differences are small and similar to those seen in June. Salinity differences are close to the average of -0.0005psu deduced from the comparison with bottles.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. The only error found was in the station name for one cast. This was fixed in the CLN file.

A header check was run. No problems were found. No off-scale fluorescence values were found.

The cruise track was plotted and added to the end of this report.

The Surface Check gave the average surface pressure as 2.3db which is within the normal range for the Vector. The lowest pressure in the CTD files was about 0.8db and pumps were on then. The CTD data show that the CTD was definitely in water at that level.

The altimeter and bottom depth headers were exported from the SAMAVG and CLN files to spreadsheets. The bottom depths were checked against the log book. There were many discrepancies; mostly the log entries had been changed so they may be a better reflection of the depth when the CTD was at the bottom.

The following calculation was made to see if the bottom depths and altimetry readings are good:

Check=maximum depth sampled (Max press*0.99) - water depth +altimeter reading

We don’t expect a 0 reading since the altimetry reading is an average over 2db and the bottom depth is subject to changes during the cast and the calculation of maximum depth is rough. Where bottom depths in the headers differed from those in the log, this calculation was used to see if changing the bottom depth brought the Check value closer to 0. In most cases it did, so the water depth was changed. Cases where the log and header depths agreed but the Check value was >4m were plotted to ensure the altimetry header made sense. There were 2 cases where the altimetry header looked a little off.

The following changes were made to the CLN files:

- Bottom depth changed to match the log entry: 38,46,60,67,68,72,73,93,96.

- The altimetry header was adjusted for casts: 36,104

- Altimetry header was removed for: 2

The only one of these changes that applied to the bottle files was a change to depth for event #38. This was made in the SAM, SAMAVG and MRG files.

Cast #39 will not be processed further as it only went to 2m to fire a bottle.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and as usual, there is a lot of variability with the up and downcast traces sometimes closer than temperature and sometimes further apart. This is likely due to varying vertical gradients and descent/ascent speeds. Overall the choice made earlier looks appropriate, so no further alignment will be applied.

Conductivity

For 2 recent cruises using the same equipment the best results varied from cast to cast and feature to feature, but overall -0.4 records for the primary and -0.5 records for the secondary looked best for 2016-07 and -0.5 records for the primary and -0.8 records for the secondary for 2016-46. Tests were run on 5 casts and the best results were with -0.4 records for the primary and -0.6 for the secondary conductivity.

SHIFT was run twice on all casts using -0.4 records for the primary and -0.6 for the secondary conductivity.

pH

The data from the pH sensor look bad, with some sudden shifts and other sections with little to no variation. Problems were noted at the end of cruise 2016-46 when it was last used. A comparison was made between profiles from near the mouth of Juan de Fuca from this cruise and from the same area and season in 2015. The 2016-10 data shows lower values for most casts, but much higher for one and there is less variability with depth. These data look unsuitable for archiving. Alignment will not help.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were warnings for 3 casts but they pertained to upcast parts of the casts so are of no concern.

##### Other Comparisons

Previous experience with these sensors –

The temperature, conductivity and dissolved oxygen sensors were used during 2 cruises in April 2016 and 2 during June 2016. The first cruise had delayed salinity bottle analysis and flushing of Niskin bottles was likely poor during all the cruises. While the CTD salinity looked low compared to bottles, it is likely because the bottle salinity was too high. During 2016-07 near the mouth of Juan de Fuca Strait where flushing was likely good, the primary salinity was found to be low by 0.0003psu and the secondary salinity low by 0.0013psu, so both calibrations appear to be ok. Dissolved oxygen comparisons also showed signs of poor flushing. Recalibration was based on results from an area where more ship motion likely improved flushing.  No recalibration was found necessary for pressure.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. As has been seen in all Strait of Georgia/ Juan de Fuca Strait cruises over the past year or two, temperature is frequently above the top of the range. For this cruise this was seen at mid-depths (~40-100m) at some casts north of Texada Island in the Strait of Georgia, and for the cast in Baynes Sound. No other temperature data was outside the climatology. Salinity was within the climatology range except for the cast at station 56 which had high salinity below 100db with the “South Georgia Strait” climatology. The deep salinity falls within the ”Gulf Islands” climatology; the border between those two regions is close to station 56.  None of these excursions suggests a problem with sensor calibration.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

There is little to distinguish between the primary and secondary T and S channels. Both have some spikes though the secondary looks slightly better. The comparison with bottles does not make it clear which salinity is more accurate. The secondary was further from bottles but if we assume the bottles are high by about 0.004psu then the secondary calibration would be better. The secondary has been archived for the 4 earlier 2016 cruises that used these sensors. The secondary data were selected for editing and archiving.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes. All but 1 file (#98) required some editing, but mostly it was removal of records from near the top and bottom of casts and light cleaning of salinity.

##### Initial Recalibration

- Pressure does not appear to need recalibration.

- The salinity calibration evidence is weak but suggests that there has not been much drift in the sensor calibration. No recalibration will be applied.

- Based on the comparisons discussed in section 5 , CALIBRATE was run using file 2016-10-recal1.ccf to correct the Oxygen:Dissolved:SBE channel in the SAM and MRGCLN2 files using:

CTD DO Corrected = CTD DO * 1.02

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly. Using just the 4 casts nearest the mouth of Juan de Fuca Strait the CTD DO is found to be higher than bottles by an average of 0.0017mL/L; before recalibration is was lower by an average of 0.0456mL/L.

See file 2016-10-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed based on residuals, the CTD DO was higher than the bottles by an average of ~0.06mL/L and standard deviation of 0.03mL/L. Looking only at the 4 casts near the mouth of Juan de Fuca Strait and excluding the same outliers as in the original fit, the average is also 0.06mL/L, but the standard deviation is higher at 0.06mL/L.  If 2 further outliers are removed, the average is 0.05mL/L and standard deviation 0.04mL/L. The results vary with depth with the CTD reading high by about 0.1mL/L near the surface and by 0.035mL/L at 300m. This pattern is likely due to incomplete flushing in the presence of significant DO gradients above 200m and slow response of the CTD DO. DO reversals were common which complicates the comparison. No further recalibration of DO is justified.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

The CTD files from rosette casts were clipped to 50m; sigma-T was derived and the data were exported to a single file, 2016-10-SOG.csv.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have a few small unstable features, but those are from sites where active mixing is expected.

The pH data look bad. Plots of all casts from this cruise were compared with a plot of data collected in the same region and same time of year during 2015-18, and while the 2015 data fell into a fairly tight group, the 2016-10 data look random and almost all profiles fall above or below the 2015-18 group. Similar problems were noted during the last 5 casts of cruise 2016-46 when the sensor was last used. The CTD technician stated that the sensor is not worth having serviced and will not be used in future.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, PAR and PAR:Reference data are nominal and unedited

except that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.3 mL/L from 0 to 50db

±0.15 mL/L from 50db to 150db

±0.06 mL/L below 150db

Channel pH:SBE was removed because the sensor malfunctioned throughout the cruise.

For details on the processing see document: 2016-10_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values were highly variable ranging from 55% to 120%. As usual values were quite low in Haro Strait, at some casts in Juan de Fuca Strait and the southern Gulf Islands region and one cast in a narrow channel to the north-west. In the open Strait of Georgia values ranged widely but were mostly between 90% and 110%. Given the wide range in surface saturation, these observations are not helpful in assessing the DO sensor calibration.

##### Final Bottle Files

CALIBRATE was run on all files and then the MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

A note was added to the header of file #38 to explain that it is a combination of events 38 and 39.

Plots were made of all casts to look for problems and none were found.

Standards check and a header check were run on all files and no errors were found.

The track plot looks ok.

Data from the CHE files were exported to a spreadsheet and compared with rosette sheets. The only problem noted is that there are records for 3 firings of Niskin #9 during cast #89. There had been no sampling and no sample numbers, so these lines were removed from the MRGREO file and Head Edit was rerun.

Plots of each file were examined and no problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

##### Thermosalinograph Data

There were no loop samples, flow meter or intake thermistor.

The only method to check calibration is to compare with the CTD casts.

a.) Checking calibrations

There were two files. The configuration files were identical.

One configuration file was renamed as 2016-10-tsg.xmlcon. No errors were found in the calibration parameters.

b.) Conversion of Files

The hex files were converted using file 2016-10-tsg.xmlcon.

They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots were produced. Temperature and salinity traces look ok until the end of the 2nd file – latitude and longitude don’t change and temperature rises smoothly, so it is assumed the TSG was left running after the ship returned to port and flow was turned off. Those data should be removed later.

The track plot looks fine. The first file started when the ship was in the northern Strait of Georgia after event #10.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.3db of 3db. These were exported to a spreadsheet which was saved as 2016-10-ctd-tsg-comp.xls. There were 73 casts which overlapped with TSG files. The data from 2db were also extracted though there were only 37 casts that overlapped with TSG data and contain data as shallow as 2db. They were added to the 3db spreadsheet.

The 2 TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for lab temperature and salinity and the files were reduced to the times of CTD files.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. There were no differences >0.0007º and the median differences were 0.0001º.

d.) Comparison of T and S from TSG and CTD data

In the table below T3 and T2 refer to CTD data from 3db and 2db, respectively. Tmed and SALmed refer to median TSG temperature and salinity, with the median taken over 5 records.

|  | Tmed-T3 | Tmed-T2 | SALmed-S3 | SALmed-S2 |
| --- | --- | --- | --- | --- |
| Maximum | 1.1416 | 0.9628 | -0.0060 | -0.8171 |
| Minimum | -0.1369 | 0.0082 | -7.3622 | -6.8473 |
| Average | 0.2376 | 0.2159 | -1.3509 | -1.3948 |
| Median | 0.1292 | 0.1267 | -1.0723 | -1.0679 |
| Std Dev | 0.2374 | 0.1953 | 1.0907 | 1.1391 |
| Median of casts with low vert. grad. | 0.1271 | 0.1258 | -1.0664 | -1.0632 |
| Median of casts with high vert. grad. | 0.3332 | 0.2543 | -1.0982 | -1.5348 |
| 20 cases with lowest TSG std dev | 0.1194 | 0.1214 | -0.9373 | -0.9188 |

- TSG vs CTD The spreadsheet comparing 3db CTD and TSG files was examined first. The differences between the salinity and temperature channels from the CTD and the TSG varied greatly with particularly large differences in the southern part of the Strait of Georgia. The sites with smaller differences tended to have lower vertical salinity gradients. The differences were much higher than noted during 2016-07 in June, but this may also reflect variations in near-surface vertical gradients. Based on a quick assessment of vertical gradients between 2m and 10m, some casts were identified as having high vertical gradients. The CTD data from 2db are in closer agreement with the TSG than those from 3db and the lower gradient casts also have smaller differences.

1. LAB TEMP The lab temperature was higher than the CTD temperature by a median of 0.127C° (standard deviation is 0.24C°). When only the 20 casts with the lowest standard deviation in the TSG temperature were included the TSG was high by 0.12C°.  Using 2db data produces a median difference of 0.0126C°.

2. SALINITY The TSG salinity data are lower than the CTD salinity by a median value of 1.07psu (standard deviation 1.09) using all casts and low by 0.93psu if only the 20 casts with the lowest standard deviation in the TSG salinity. The differences are slightly lower when the 2db data are used.

The temperature difference could be due chiefly to heating in the loop, but that would not explain why the TSG salinity is low. There are two possible explanations. The TSG may be drawing water from above 2db. There could be a layer of fresh water that we are not seeing in the CTD records; there was fairly heavy rain during the week before the cruise. The largest differences were in the southern Strait of Georgia where the influence of the Fraser River often results in very high near-surface gradients. Bubbles in the loop would lower the TSG salinity, but there is no evidence of large bubbles in the traces. There could be small ones but we would not expect such a large effect from those.

(See 2016-10-ctd-tsg-comp.xls.)

- Calibration History

The TSG temperature and conductivity were recalibrated in December 2015 and has only been used on one other cruise that has been processed: 2016-07 in June 2016 in the same region. At that time the lab temperature was higher than the CTD temperature by between 0.1C° and 0.2C°.  Salinity was found to be low by 0.62 using all casts, but the standard deviation was very high at 3.0psu. The median difference was 0.03psu when only the 10 casts with the lowest standard deviation in the TSG salinity were included.

Conclusions

1. The TSG clock worked well.

2. The TSG lab temperature was higher than the CTD temperature by about 0.12Cº for cases with low standard deviations in the TSG data. We have little experience with the Vector loop system and no knowledge of the flow rate during the cruise. From the Tully we might expect heating of ~0.15Cº given similar intake temperatures, but the Vector loop is shorter so that would reduce the amount of heating. If the ship’s ambient temperature was particularly high that would increase it. We don’t know the flow rate which would also affect the amount of heating in the loop. The difference between TSG and CTD temperature could also be due to water in the loop coming from slightly higher than 3m and/or calibration drift.

3. The difference between the TSG Salinity and that from the CTD is highly variable and higher than found during the June 2016 cruise in this area. The TSG salinity is especially low compared to CTD when vertical salinity gradients are higher. This suggests that the intake is coming from higher in the water column. There could be some effect of bubbles in loop water, but there is no sign of large bubbles in the traces.

4. Overall the quality of the data was good with just one section at the end that was recorded while the ship was not moving and there was apparently no flow in the loop.

5. There is insufficient evidence to justify making an estimate of intake temperature or to recalibrate salinity.

f.) Editing

CTDEDIT was used to remove records from both files:

2016-10-0001 – Temperature and salinity data were removed from the first 9 records before flow was well established.

2016-10-0001 – Records from the end of the file (records 5403-6737) were removed since the ship did not move and a steady rise in temperature showed that flow in the loop had stopped.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Conductivity:Primary, Flag and Position:New channels.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats. The Temperature:Primary was renamed as Temperature:Lab to make it clear that it is not the intake temperature.

Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series and all look fine.

The cruise plot was added to the end of this report.

Particulars (mostly notes from log)

1. Test station to bottom – 10.

2. Test station to 100m

20. Depth report from bridge in error – should be 296, not 269. Fixed in files.

22. Bottle 12 fired but not tripped.

31. Bottle 12 did not trip – the mechanism was changed.

38. Bottle 3 was missed.

39. Event to trip bottle 3 that should have been tripped during event 38. No CTD file needed. Combine 38 and 39 bottle files.

42. Vents left open – bottle data discarded. No bottle file needed.

78. Bottle 11 has chip in bottom seal. Taken out of service.

89. Bottle 9 was fired 3 times though log says it did not fire. The 3 lines were removed from the bottle file as there was no sampling.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 18Dec2014 | Factory |  |  |
| Conductivity | 1766 | 20Aug2014 | Factory |  |  |
| Secondary Temp. | 4054 | 20Aug2014 | Factory |  |  |
| Secondary Cond. | 3531 | 6Feb2013 | Factory |  |  |
| Transmissometer | 1396DR | 2Feb2016 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 19Dec2015 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| Surface PAR | 20518 | 21Mar2016 |  |  |  |
| pH | 0692 | 12Jan2016 | Factory |  |  |
| SeaPoint Fluor. | 3865 | n/a |  |  |  |
| Pressure Sensor | 0550 | 17Dec2014 | Factory |  |  |
| Altimeter | 1253 | n/a |  |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3363       Cruise ID#:	2016-10

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3363 | 17Dec15 | Factory |  |  |
| Conductivity | 3363 | 17Dec15 | Factory |  |  |