## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2016-12

Agency: OSD

Location: North-East Pacific

Project: WCVI Moorings / WCTSS

Party Chief: Spear D.

Platform: John P. Tully

Date: 4 July 2016 – 19 July 2016

Processed by: Germaine Gatien

Date of Processing: 24 October 2016 – 9 November 2016

Number of original HEX files: 92	Number of CTD files: 90 (2 shallow test casts not processed)

Number of bottle files: 44		Number of bottle casts processed: 42

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), a SBE 43 DO sensor (#997) on the secondary pump, a SeaPoint Fluorometer (#3640) on the primary pump, QSP-400 PAR sensor (#70613) and an altimeter (#62354).

Seasave version V7.23.2 was used for acquisition.

The data logging computer was the Tully CTD Laptop (Acer).

The deck unit was a Seabird model 11+, serial number 0425.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

The oxygen kit was Scripps kit #2.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log, rosette log sheets and analysis logs were generally in good order with detailed comments. However, there were a number of problems with sample numbers and rosette records not being in agreement. There was also a case of 2 sample numbers being repeated but that was noted in the log so adjustments were made to labels before analysis began. In another case a hand-written sample number on an oxygen sample did not match the rosette record and there is no note of explanation.

Thanks to the science crew for making a number of deck pressure checks during the cruise. The results were consistent and confirmed indications from a previous cruise that the CTD pressure offset needed a significant adjustment. That offset was used in processing these data.

The comparison of salinity samples with the CTD salinity suggested that the primary salinity was low by about between 0.007 and 0.009 and the secondary was within 0.001 of bottles.

Having a lot of salinity bottle data from areas with good depth coverage and gathered in different regions enabled some useful studies.

- Separating cases where bottles were fired at the bottom of a downcast made it clear that incomplete flushing of bottles near the bottom has the opposite effect to that found for bottles fired on the upcast. This seems obvious in theory, but it was useful to be able to demonstrate it clearly. Even in conditions that would normally lead to good flushing, firing bottles near the ocean bottom is likely to capture water from well above the bottom of the cast. Waiting a while and then coming up a few metres before firing may produce better results.

- Separating cases with steady CTD descent rates from those in open waters where those rates were fairly noisy showed that flushing of bottles is poorer in the quiet areas. It seems obvious that flushing would be less efficient in areas where calm conditions keep the ship from bouncing around, but it is useful to demonstrate this. Another factor is that near-surface vertical gradients are often larger in inlets as well, so the effect of poor flushing is more significant, but below 50m the inshore gradients are similar and often lower than offshore. So there must be differences in flushing efficiency as well. There may be flushing errors under all conditions but sampling from open waters and depths with low vertical gradients will produce the most reliable calibration information.

As noted during previous uses of this equipment, there is an unusually large difference between the two temperature channels, with the secondary lower than the primary by ~0.002Cº. There is also a large difference between salinity channels that is only partly explained by the temperature difference. It is not known which temperature sensor is more out of line, but correcting the primary temperature by subtracting the difference and recalculating salinity brings that channel closer to the bottle salinity. Correcting the secondary temperature by adding the difference would decrease the secondary salinity and hence increase its difference from bottles. The secondary channels were selected for archiving for all casts, and no recalibration of temperature or salinity was applied. If a post-cruise calibration shows this choice to have been inappropriate, a recalibration can be applied later.

For this cruise the CTD was lowered to about 2m with pumps turned off, the CTD was stopped while the pumps were turned on and there was a brief wait (roughly 40s to 70s). There is sometimes noise in the salinity channels in the top 10m that may come from bubbles in the CTD system; the two channels come into better correspondence before the CTD reaches 10m. Some of these casts are in protected areas where doing a 10m soak may leave the near-surface waters disturbed for a long time. The research requires a profile of transmissivity near the surface in undisturbed water. Monitoring the 2 salinity channels after the pumps come on, and waiting until they are fairly close together before lowering the CTD might improve the near-surface salinity data.

The CTD fluorescence shows the usual pattern, reading higher than the extracted CHL for CHL<1ug/L and then falling relative to CHL until it is about 70% of CHL for the maximum CHL value of 10.7ug/L.

The SBE DO channel was recalibrated based on comparisons with bottles fired during offshore casts. The comparisons from inshore suffer from similar problems to those found for salinity, with poor flushing in the presence of large gradients producing results quite different from the offshore group.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples, using offshore casts to minimize the effect of incomplete flushing of Niskin bottles. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

The Oxygen:Dissolved:SBE data in the CTD files (downcast) are considered, very roughly, to be:

±0.4 mL/L from 0 to 100db

±0.2 mL/L from 100db to 200db

±0.1 mL/L from 200db to 500db

±0.04 mL/L below 500db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked. All were used during 7 previous cruises since they were last serviced.

All configuration files were the same. File 2016-12-ctd.xmlcon was prepared by making 2 changes to the file used at sea.

- The PAR sensor is a new one and the factory calibration includes an offset that was not entered in the con file, so that was added.

- An adjustment of -1.0db to the pressure offset was determined to be necessary during processing of cruise 2016-06, partly based on the many deck measurements made during this cruise.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2016-12-ctd.xmlcon to create CNV files. The hysteresis and Tau functions were selected since there are some deep casts. There were 2 PAR sensors in the configuration file throughout the cruise but according to the log book sensor #4615 was not mounted, so channel PAR2 was not converted after a few tests conversions confirmed that it was not in use.

A few casts were examined and all expected channels are present.

Temperature and conductivity are generally close during downcasts but, as usual, they are further apart during upcasts. Primary temperature and conductivity channels have some odd excursions and spikes. For example, in cast #105 during the surface soak with pumps turned on, the primary and secondary temperature and conductivity are quite different. This may be related to the deployment method used for this cruise. The CTD was lowered to 2m with pumps off. The pumps were then turned on and after 40s to 70s wait the full cast was run. The primary and secondary traces track well below 10m, so the problems above that level may be related to bubbles. These can be removed by doing a 10m soak before the full cast, but the researchers wanted to observe undisturbed surface waters.

Fluorescence, dissolved oxygen, PAR and transmissivity traces look normal. The PAR sensor was removed before cast #65.

The altimetry usually looks fine but for some casts there are spikes near the bottom so the header entries will need to be checked carefully after conversion to IOS Header format.

Two casts contain only surface data, 2016-12-9999 and 2016-12-0046. The first was a test and the second was run just to fire a single bottle at the surface. File 2016-12-0047 is at the same site. Those 2 casts will not be processed.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2016-12-ctd.xmlcon. The naming format was changed for 3 casts.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. For casts #105 and 106 a few spiky points were removed from channel Salinity:T0:C0 using CTDEDIT.

The output files were copied to *.bot.

Event #1 was just a test to ensure bottles closed – there was no sampling. There was particle sampling only from event #2. Those 2 files will not be processed further.

A preliminary header check and no problems were found. CTD fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs. Note that samples 332 and 333 were used for 2 casts, so the second occurrence was changed to 9332 and 9333. Some sample numbers were not used.

The bottle file for event #100 is problematic. Some spigots were noted to be open when the CTD got to the surface so the CTD was lowered again and more bottles fired to replace the missing levels. The problem is that there could be leakage into bottles fired near the surface when they move to deeper waters since the pressure is higher outside than inside the bottle. The shallowest bottle was fired at 2db and during the 2nd drop was lowered to about 125db. Comparing DO samples at 20db from nearby casts does not suggest that there is a significant effect. Only dissolved oxygen samples were drawn. It is difficult to judge if data are out of line because surface variability is very high.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2016-12-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2016-12chl*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2016-12chl.csv. The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2016-12oxy.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and was then saved as 2016-12oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in file 2016-12SAL.xlsx which included a precision study. The analysis was done between 9 and 28 days of collection. The files were simplified and saved as 2016-12sal.csv. File 2016-12sal.csv was then converted to individual SAL files. There were no duplicates.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2016-12_nutrients*.xlsx. This includes a precision study. The file was simplified, reordered on sample numbers and saved as 2016-12-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps. This process took many reruns because of some missing sample numbers that were not noticed in the first building of the ADDSAMP file. File #65 was the biggest problem. Also the SAL file for event #68 had to be adjusted to sample number order by moving sample 9333 to the end. There was no other sampling from the 2 bottles with the odd numbering. There were also discrepancies between the rosette log sheets and the sample numbers reported by the analysts for event #29.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions and a few problems were found and fixed, but others were found later.

##### Compare

COMPARE was run early before final bottle files were ready. It was later found that there were errors in the sample numbers in bottle files for events #29 to 46 but tests runs showed they had no effect on the COMPARE results.

Salinity

Compare was run with pressure as reference channel.

All severe outliers in the comparison (difference >0.05) are cases where the standard deviations in the CTD salinity are >0.001, so they are likely due to high local variability.

Outliers with low standard deviation and differences >0.02 were investigated and in all cases there were local gradients high enough that poor flushing could easily explain the differences, with values similar to those of the bottles being found within a few metres of the firing depth. There were a few cases of bottom bottles and differences out of line in the opposite direction, and that can also be explained by flushing problems since the Niskin bottle would have carried water from above.

The bottles flagged by the analyst did not show up as major outliers but some were in places where outliers would not be obvious because of high variability.

When all casts are included after excluding bottles above 150db, those with standard deviations in salinity >0.001 and those with differences more than 0.01 from the initial estimate of average difference, the primary salinity was found to be low by 0.0067 (std dev 0.0043) and the primary low by 0.0001 (std dev 0.0043). Excluding data from above 150db was made to avoid the largest errors due to poor flushing in the presence of high vertical gradients.

Next offshore casts were examined separately since flushing is likely better so that surface data are more likely useful. When outliers were identified in the same way as for the full set except that there was no restriction on pressure, the primary salinity was found to be low by 0.0087 (std dev 0.0027) and the secondary was found to be low by 0.0010 (std dev 0.0035).  The difference between these two results is close to the test results reported in section 9. Excluding values above 150m made a small difference with the primary low by 0.0082 and the secondary low by 0.0008; that difference might be due to removing bottles with poor flushing, but if flushing is an issue it may be the result of having a bigger influence from bottles fired at the bottom of downcasts. There may also be some small effect of evaporation/adsorption so that bottle values may be slightly high, but this effect should not be >0.001 given the quick analysis. So considering primary CTD salinity to be low by ~0.008psu looks like a reasonable estimate and secondary salinity is likely within 0.001psu.

This cruise presented good salinity sampling so 2 studies were done by examining the secondary salinity data with differences <0.01.

- Separating cases where bottles were fired at the end of the downcast made it clear that incomplete flushing of bottles near the bottom has the opposite effect to that found for bottles fired on the upcast. This seems obvious in theory, but it was useful to be able to demonstrate it clearly. Bottles fired at the end of the downcast have CTD salinity higher than bottles with only 2 out of 14 bottles showing the CTD salinity lower. For bottles NOT fired at the bottom there is a lot of scatter near the surface, but if we look at bottles below 150m, only 2 out of 11 show the CTD reading higher than bottles. Bottles fired below 600m are all bottom bottles, so a comparison was done of bottles between 100 and 510db where there are both types. The bottles fired at the bottom were higher than CTD salinity, on average, by ~0.002 while those fired on the way up were lower by an average of ~0.0015. These results would differ according to conditions. These data mostly came from the offshore region.

| # bottles | 9 | 14 |
| --- | --- | --- |
|  | Bottom of cast | not bottom |
| pressure | 100-510db | 100-510db |
| average | 0.001953 | -0.00154 |
| Std dev | 0.004844 | 0.004095 |

- The CTD-bottle comparison includes plots of differences against file pair number, and the results show an obvious difference between those from waters with steady CTD descent rates and those in open waters where those rates were noisy. Again, it seems obvious that flushing would be poor in areas where calm conditions keep the ship from bouncing around, but it is good to demonstrate this to be true. The study of offshore casts also led to different results from those using all casts. This is hard to quantify since some of the differences seen inshore are due to high vertical gradients exaggerating the effect of poor flushing. We could just compare deeper data, but then the difference between bottles from the bottom compared to the others confuses the issue. It is also likely that vertical gradients vary between the two groups, so the effect may be somewhat exaggerated. In any case it is clear that calibration sampling is best done in open waters or in stormy conditions. Where that is not possible, the best we can do is sample in low gradients and keep in mind that flushing is likely inefficient when interpreting the results.

For full details for the COMPARE run see file 2016-12-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

- All significant outliers come from the top 20m, almost all of those from around 10m where variations are high and there are local reversals. They are almost all found in Effingham Inlet and Douglas Channel where descent rates were very steady so that flushing of Niskin bottles is likely incomplete, so that bottle values are often significantly lower than ambient conditions. For example, during event #100,  sample 494 (4.79mL/L) looks much too low for 10m but during the bottle stop the CTD DO varied from 4.4mL/L to 6.7mL/L, so this is likely a case of slight inefficiency in flushing combined with a very high DO gradient.

There were some problems on a first run that were fixed by correcting 2 errors:

- Event #2 – This problem arose from the oxygen samples being mislabeled as from event #2; they really came from event #3.

- Event #65 – The sample labels did not agree with the rosette log sheet – 2 sample #s were skipped on the rosette sheet, but not by the label maker. When the analysts changed their samples to match the log, many outliers disappeared.

Fits are highly dependent on how outliers are excluded. When all casts are used excluding some outliers based on residuals, the fit was:

CTD DO Corrected = CTD DO * 1.0249 + 0.0251   (1)

Another fit was done using only the offshore casts with descent rates that suggest flushing of bottles should be reasonably good.

CTD DO Corrected = CTD DO * 1.0256 + 0.0324   (2)

The slopes are almost the same, but the offsets vary. Removing outliers based on residuals removed a lot of the inshore data from the first comparison. All sets contain many bottles with DO <2mL/L.

To get the best recalibration of the CTD sensor it is best to exclude cases where flushing is more likely to be affecting the comparison.

This sensor was last used during a Line P cruise, 2016-06, when we expect a good fit from deep casts where flushing is not considered to be a problem.

Most bottles flagged by the analyst come from areas where flushing is likely poor so it is hard to determine if the values are significantly out of line. The 46 flag for the bottom sample of cast #59 was examined and it was found that the higher value (1.968) looked way out of line. The lower value (1.723) fits into the comparison very well as the CTD value is 1.67, so after recalibration will be ~1.72. The analyst decided to reject the higher value.

For more details see document 2016-12-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor.

The plot of FL/CHL vs CHL below shows the usual pattern with CTD fluorescence reading a little higher than the extracted CHL for CHL<1ug/L and then falling relative to CHL until it is about 70% of CHL at the highest CHL value of 10.7ug/L.

A fit forced through the origin has a slope close to 1 when the one high CHL sample is excluded.

For full details of the comparison see file 2016-12-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on a few casts. The results are always hard to judge because the upcast data are noisy, but a setting of +3s looked reasonable and was better than either +2s or +4s. A+3s setting was the choice for the past 3 cruises that had this equipment. ALIGNCTD was run on all casts using +3s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. One cast was checked using the default setting and it improved the data for both channel pairs.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. The shaded values are from 3 previous cruises that used these sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2015-21-0034 | 1000 | -0.0013 | -0.00014 | -0.0002 | High, Mod |
|  | 1900 | -0.0014 | -0.00006 | +0.0008 | Med, Noisy |
| 2015-21-0099 | 1000 | -0.0014 | -0.00008 | +0.0004 | High, XNoisy |
|  | 1900 | -0.0013 | -0.00002 | +0.0012 | “ |
| 2016-47-0039 | 1000 | -0.0019 | +0.0001 | +0.0032 | High, V Noisy |
|  | 1900 | -0.0018 | +0.0002 | +0.0040 | “ |
| 2016-47-0105 | 1000 | -0.0017 | +0.00015 | +0.0038 | High, Noisy |
|  | 1900 | -0.0018 | +0.0002 | +0.0045 | “ |
| 2016-06-0029 | 1000 | -0.0021 | +0.0003XN | +0.0045 VN | High, Moderate |
|  | 1900 | -0.0017 | +0.0003 | +0.0050 | “ |
| 2016-06-0068 | 1000 | -0.0016 | +0.0002 | +0.0045 | High, Moderate |
|  | 1900 | -0.0016 | +0.0003 | +0.0052 | “ |
|  | 3800 | -0.0020 | +0.0003 | +0.0059 | “ |
| 2016-06-0079 | 1000 | -0.0016 | +0.0003 | +0.0048 | High, Moderate |
|  | 1900 | -0.0017 | +0.0003 | +0.0055 | “ |
| 2016-12-0067 | 1000 | -0.0017 | +0.0004 | +0.0078 | VHigh, Noisy |
|  | 1900 | -0.0018 | +0.0005 | +0.0083 | “ |
| 2016-12-0068 | 1000 | -0.0022VN | +0.0005 | +0.0080 | High, VNoisy |
|  | 1900 | -0.0019VN | +0.00055 | +0.0085 | “ |

The temperature differences are similar to those from 2016-47 and 2016-12. The conductivity and salinity differences are slightly higher.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. Errors were found in the station names for events #2, 3, 20, 32 and 46. These were fixed in the headers of the CLN and SAM files. The SAM files were averaged again.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 0.45db which is low but expected from this cruise as there was a need for data from the surface. The associated salinity values are very low. For the offshore casts the pumps were not turned on until the CTD was at 2 to 4db. During the previous cruise a decision was made to change the pressure offset to -1.0db. During this cruise deck pressures were taken frequently and were quite stable at 1.0 to 1.4db, so the -1.0db offset looks appropriate since the pressure offset used in acquisition was +0.05db. The -1.0db offset was used in the conversion of the files during processing. Examination of a few files showed that when pressure was close to 0db the CTD appeared to be very close to the surface. In one case when pressure is +0.2db, the transmissivity suggests the CTD was in air. Salinity looks to be in water but the pumps were off. In another case when pressure was about 0db and definitely in water judging by the transmissivity. So the pressure appears to be good to ±0.2db, and perhaps reading slightly too high.

The altimeter and water depth readings from the headers of the CLN and SAM files were exported to a spreadsheet.

A rough calculation was made to find faulty water depth or altimetry header entries:

Check = (Pressure*0.99 -Depth +Altimeter Reading)

Where the Check value was >2.5db the water depths were compared with log entries. There were many cases where there were differences and using the log value produced a better result, so the water depth entry was changed in both the CLN and SAM files. In cases where the water depth was correct or fixing that still produced a high “check” value, a plot of altimetry near the bottom was examined and most values looked reasonable, but 2 header entries were removed because the entry was based on faulty interpretation of spikes.

The SAM files were bin-averaged again.

The final merge step and CLEAN run were repeated for the bottle files.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 5 casts to see what shift in conductivity does the best job of removing noise in the salinity channels Nothing improved the primary which seemed unusually smooth to start with, while -0.6 records worked best for the secondary.

SHIFT was not run for the primary conductivity.

SHIFT was run on all casts using -0.6 records for the secondary conductivity. Salinity was recalculated.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

After an initial run, casts #1 and 26 had initial records removed and DELETE was rerun.

##### Other Comparisons

Previous experience with these sensors –

The sensors were used during 4 cruises in 2015 and 3 just before this cruise in 2016. The pressure offset was updated in stages on the previous 3 cruises. Salinity calibrations were mostly not trusted due to flushing problems, delayed analysis and/or noisy comparisons for most of the cruises. The secondary salinity has been closer to bottles than the primary for all cruises. The temperature sensors were found to be further apart than usual during the previous 2 cruises. There were good comparisons for 2016-47 and 2016-06 for which the primary salinity was found to be low by ~0.0074 and ~0.0065, respectively, but after a correction to temperature the salinity channel was low by 0.004 and 0.0045. The secondary salinity was low by 0.002 and 0.0015 for the last 2 cruises. The dissolved oxygen comparisons have been reasonably consistent.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Where local climatology was available most temperature values were within the ranges except for a few instances of temperature being above the maximum between 100 and 200m at station SOGN and a few places in Hecate Strait; one deep cast from Queen Charlotte Sound had slightly high temperature at about 1500m. High temperatures have been widely noted over the past few years, and in particular high values have been seen during other cruises in summer 2016 in the northern part of the Strait of Georgia below 100m. Salinity values were within the climatology except for one cast for which salinity was a little low at the bottom. However, the climatology for that cast goes only to about 200m while the profile was more than 400m deep, so the climatology must not include the area sampled. None of the excursions looks indicative of calibration problems.

Repeat Casts – There was a repeat cast in the Strait of Georgia, but local variations are too high to expect good repeatability in this area. Differences along lines of constant σt were about 0.02 for temperature and 0.005 for salinity at σt= 23.6 (near 280m) for both T-S pairs. That is a reasonable agreement for this area.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

An initial examination to decide which channels to archive showed that the primary was noisier with more unstable features. It was also much further from bottle salinity. The secondary T and S channels were selected for archiving for other recent uses of this equipment, and were also chosen for this cruise.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of most casts.

After editing T-S plots were examined for all casts. The only unstable features in the T-S plots were small ones near the surface which may be real features.

##### Recalibration

There is no evidence to suggest that the pressure channel needs further adjustment.

The secondary salinity was found to be within 0.001 based on offshore bottle sampling, so no recalibration will be applied to that channel.

There is no plan to archive the primary channels, but if there is a later decision to revisit that decision, recalibration would be required by subtract 0.002 from the primary temperature, followed by recalculation of the primary salinity. That step would raise the primary salinity by about 0.002, so that it would then be reading low by ~0.006psu. So, a second recalibration would then be needed to add 0.006psu to the primary salinity, unless post-cruise calibration reports suggest otherwise.

Dissolved oxygen recalibration was discussed in section 5.

File 2016-12-recal1.ccf was prepared to apply the following correction to channel Oxygen:Dissolved:SBE:

CTD DO Corrected = CTD DO * 1.0256 + 0.0324

These corrections were first applied to the SAM and MRGCLN2 files. COMPARE was rerun and the results confirm that the calibration was done correctly. When all casts are included and outliers removed in the same way as in the original comparison, the average of differences in the DO fit was +0.0011mL/L and the standard deviation was 0.021mL/L. But if the casts are divided between onshore and offshore, the results are:

Inshore:	Average = +0.0288mL/L; Std.Dev = 0.026

Offshore:	Average = -0.0018mL/L; Std Dev = 0.019

The recalibration used was based on the offshore data, so we expect a good result and we find it. Had the fit been based on inshore data the correction applied would have been smaller. This fits the assumption that flushing is poor in protected areas, at least when conditions are calm. Some of the effect may be due to lower near-surface gradients offshore, but even below 100m where offshore gradients tend to be higher, the differences between CTD and bottles is similar or lower than for inshore.

(See file 2016-12-DO-comp2.xlsx for details.)

CALIBRATE was run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from roughly the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When outliers were removed based on residuals the offshore casts show the CTD DO to be higher than the bottles by ~0.05mL/L while the inshore casts show it to be higher by an average of 0.13mL/L. The slope of the fit of differences against DO is quite flat for the offshore casts, with small differences for high and low DO values and larger differences between those. The comparison is very good in areas of low DO gradient and for the offshore the top 20m were usually quite well mixed. For the inshore the slope of the fit is higher with large differences for high DO. This is to be expected since the near-surface gradients were higher for the inshore casts, so flushing problems should be more significant there. The scatter is large with standard deviations of 0.05mL/L and 0.08mL/L for the offshore and inshore comparisons. The results suggest that even for the offshore casts the flushing of bottles may be somewhat inefficient. The results in the well-mixed regions suggest the recalibration was appropriate and that no further recalibration is justified.

See 2016-12-dox-comp3.xlsx for details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

PAR was removed for casts #65-108.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and unedited except that

some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

The CTD deployment method was to lower the CTD to about 2m with the pumps off.

The pumps were turned on, there was a wait of 40 to 70s and then the

full cast was run. For some casts the two salinity channels differ markedly

near the surface after the pumps came on; this may be due to bubbles from

the CTD system. Below 10m the two salinity channels track well.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 100db

±0.2 mL/L from 100db to 200db

±0.1 mL/L from 200db to 500db

±0.04 mL/L below 500db

For details on the processing see document: 2016-12_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

A cross-reference list was produced.

The sensor history was updated.

The track plot looks fine.

##### Special files for Dr. Johannessen

All DEL files were clipped to 10db and then bin-averaged to 0.5db bins. Recalibration was applied. The data were extracted to a spreadsheet, including the pump status channel. The file was ordered on pump status and for 0 values the pumped channels were removed. Also removed were the primary temperature and salinity, both conductivity channels and dissolved oxygen, scan number, altimeter, descent rate and flag. The first bin may contain some data collected while the CTD was out of water since pressure at the surface varies slightly from cast to cast.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Surface DO saturation varied from 80% to 135%. The lowest values were in areas where active mixing is likely to mix DO down. The highest values were in inlets including Effingham Inlet. The deep offshore casts had values from 104% to 108% which is in the about what is expected offshore. The values suggest that the DO calibration is reasonably good.

##### Final Bottle Files

The MRGCOR2 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

PAR was removed for casts #65-108.

A second SBE DO channel was added for both the CTD DO and titrated DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2016-12-bottles-final.xlsx. The entries were compared with the rosette log sheets to ensure no samples had been missed. A few problems were found and corrected in the bottle files and the spreadsheet was then recreated.

- An oxygen sample for cast #24 was entered as #973 but according to the rosette sheet should be #74. That was corrected in the 2016-12oxy.csv file and the merge process repeated.

- Cast #97 had a line for which there was no sample number and no sampling, so that was removed.

Standards check and a header check were run on all files. No problems were found.

The track plot looks ok.

A cross-reference listing was produced for the CHE files.

Particulars - Notes from Log Book

1. Test cast – bottles fired but no samples taken

3. Deck pressure 1.1

4. Deck pressure 1.2

7. Deck pressure 1

19. Altimetry odd at bottom

23. Fluorometry spiky

25. Deck pressure 1.3

26. Started descent with pump off, returned to surface and ran again.

28/29/31. Deck pressure 1.1 for all 3

40. Missed tripping bottles so back to b-5 to redo

47. Station name wrong. Should be SC61, not WC43. Fixed after conversion.

52. Bottle 19 Si tubing broke – replaced with normal line and spring.

53. Deck pressure 1.1

54/55. Deck pressure 1.1

62. Bottles 2 and 3 both tripped at 200m.

63. Deck pressure 1.2. Went back to trip second 50m bottle.

64. Altimeter and sounder in serious disagreement. Were either 5mab(sounder) or 50mab (altimeter).

65. PAR removed.

66. Termination re-taped prior to cast.

68. Samples 332 and 333 are repeats from previous cast – relabelled 9332 and 9333.

69. Deck pressure 1.3.

70. Deck pressure 1.1.

71. Deck pressure 1.3

72. Deck pressure 1.2

80. Pause at 168m due to large aft lead

93. Serious groove in plastic guide on LARS.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 17Dec2014 | Factory |  |  |
| Conductivity | 2280 | 19Dec2014 | Factory |  |  |
| Secondary Temp. | 2663 | 15Jan2015 | Factory |  |  |
| Secondary Cond. | 2754 | 19Dec2014 | Factory |  |  |
| Transmissometer | 1185DR | 27Apr016 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 17Jan2015 | Factory |  |  |
| PAR | 70613 | 21Mar2016 | Factory |  |  |
| SeaPoint Fluor. | 3640 | n/a |  |  |  |
| Pressure Sensor | 0443 | 17Dec2014 | Factory |  |  |
| Altimeter | 62354 | n/a | Factory |  |  |