## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2016-29

Agency: PBS, Salmon and Freshwater Ecosystems Division, Nanaimo, B.C.

Party Chief: Neville C.					Location: Strait of Georgia

Project: SoG Juvenile Salmon Trawl Survey		Platform: Nordic Pearl

Date: 4 July 2016 – 14 July 2016

Processed by: Germaine Gatien

Date of Processing: 7 November 2016 – 9 November 2016

Number of original HEX files:	26		Number of CTD files:  26

# INSTRUMENT SUMMARY

Two CTDs were used for this cruise:

- SeaBird Model SBE-25+ CTD (s/n 1123) was used with temperature sensor #4888, conductivity sensor #4513, Wetlabs ECO Fluorometer #2216, oxygen sensor #3234 and pressure sensor 1123.

- SeaBird Model SBE-25 CTD (s/n 0404) was used with temperature sensor #2095, conductivity sensor #1764, Wetlabs ECO Fluorometer #4185, pH sensor #0852, oxygen sensor #1176 and pressure sensor 0482.

# SUMMARY OF QUALITY AND CONCERNS

The log book had equipment lists and there was a good record of which CTD was in use for each cast as well as problems encountered. The log times were in UTC. A spreadsheet was provided with positions, station names, water depths and CTD serial numbers; this was very helpful since position information is not in the headers. There were a few errors in file names that were adjusted to match log event numbers.

The CTD used for the first 4 casts was a new one. The raw files had extensions XML rather than the usual HEX and the headers did not contain the usual instrument information. After the 4th cast the CTD was changed.

The CTD was soaked at the surface for about 1 minute, lowered to 10m and brought back to the surface. After a wait of about 1 minute the full cast was run, though the wait was shorter for the first cast.

A pH sensor was mounted on CTD #0404 but not on CTD #1123.

CTD #0404 had poorer pressure resolution than expected from a SeaBird 25. This was also noted during 2016-58 and the CTD technician was informed that there may be a problem with the settings for the instrument. The fluorometer on CTD #0404 produced some slightly negative values at the bottom of casts which were replaced with pad values.

CTD #01123 performed as expected.

There was no calibration sampling.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea.

2. Preliminary Steps

CTD #0404 used for parts of cruises 2016-31 and 2016-58 with the same sensors. There is no history available for CTD 1123.

The Daily Log was obtained.

The cruise summary sheet was completed.

A single configuration was file used for each CTD. The calibration parameters were checked and no errors were found.

## 3. Conversion of Raw Data

All casts were converted using the appropriate configuration files.

The pressure at the beginning and end of casts for both CTDs was about 0.2db so there is no evidence of a problem in pressure calibration. The pressure from CTD #0404 is full of steps of >1db. This was noted on other 2016 cruises using this CTD. CTD #1123 looks very smooth with no obvious steps but the sampling rate is 16Hz while it is 8Hz for CTD #0404.

Plots show that the channels all have reasonable profiles.

The times in the log are in UTC.

The times in the header for CTD #1123 are in PDT so 7 hours will need to be added later.

As was found for other recent cruises using CTD #0404, the times in the headers are in PDT but a day ahead, so actually 17 hours ahead. So 17 hours will need to be subtracted later.

4. WILDEDIT

This step was skipped as it does not usually work well on SBE25 data and no obvious spikes were noted.

5. FILTER

Tests were run on a few casts to choose a filter for these data.

- SBE25+ 1123– The pressure looks fine and it can be filtered in the DELETE process, but temperature and conductivity were investigated to see if the salinity improved with filtering and it did. Salinity was quite spiky when the filter was not used. So WFILTER was run on casts 1-18 using a width of 5 for pressure, temperature and conductivity

- SBE25 0404 -The pressure looks steppy for all widths tested using WFILTER but the filter worked well on the temperature and conductivity and the resulting salinity looks improved. The pressure readings from this sensor are odd with the steps being ~1.3db, much grosser resolution than expected from an SBE25. This same problem was noted in 2015. WFILTER did improve the pressure but there were still many reversals until the setting was increased to width 9 for pressure. WFILTER was run on casts #32-139 a width of 9 for pressure, and a width of 5 for temperature and conductivity.

6. ALIGNCTD

For both CTDs a setting of +2.5s does improve the alignment of dissolved oxygen with temperature, though there only 2 files with full down and upcast data from CTD 1123.

ALIGNCTD was run on all files using +2.5s.

7. CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25. CELLTM was run on all casts using that setting.

8. DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers. This was done separately for each CTD to get the correct information in the headers.

CLEAN was run to replace pad values in the pressure channel with interpolated values and to add event numbers to the headers.

Time was adjusted using program ADD TIME CHANNEL as follows:

- For CTD #1123 the times were adjusted by adding 7 hours.

- For CTD #0404 the times were adjusted by subtracting 17 hours.

10. Checking Headers

Some header information was missing: latitude, longitude, station name and water depth.

A spreadsheet was provided and that was used to prepare file 2016-29-header_mrg.csv. IOS SHELL program “MERGE:CSV file to headers” was used to add information from the csv file to the IOS files with output MRH.

A track plot produced looked ok.

A cross-reference listing was produced and checked against the log entries.  There are consistent differences in time but as they are ~2 to 3 minutes and it is not known which is more reliable, no adjustment was made. There were some more serious problems found in event numbers:

- Event #19 was saved with name 2016-29-0018.hex.

- Event #37 was saved with name 2016-29-0038.hex.

- Event #140 was saved with name 2016-29-0139.hex and was also saved with that name in file 2016-29-CTD-Bongo Data.xlsx.

The converted file names were adjusted to match the log book, and the steps described in this section were repeated. The cross-reference list was produced again and no further problems were found.

A header check was also run and the only problems noted are that there are some negative values in fluorescence and dissolved oxygen. These may disappear when the initial soak and upcast data are removed: if not, CLEAN can be run to replace them with pad values.

The surface check was run and the average surface value was 0.86db. For CTD #1123 the average was 0.06db, while the #0404 values were all 0.17db or 1.40db. These initial values all had extremely low salinity and were clearly gathered just as the CTDs entered the water.

##### 11. SHIFT

pH

- CTD #1123 did not have a pH sensor.

- For CTD #0404 a shift of +30 records bring the upcast and downcast pH trace into better alignment with temperature traces, though it is hard to judge because the pH traces cannot capture sharp changes. SHIFT was run on casts #32-140 to advance pH:SBE by 30 records.

Conductivity

For CTD #1123 tests showed some improvement in stability in T-S plots with a shift of +1.2 records followed by a recalculation of salinity. There were only 4 casts and few unstable features from which to judge. SHIFT was run on all 4 casts using an advance of +1.2records.

As was found the last time CTD #0404 was used, all shifts applied to conductivity produced unstable features in the T-S plots. No adjustment was applied.

Fluorescence

The fluorometers on the SBE25s were not pumped, so a shift in alignment is expected to be small or unnecessary. Profile plots of temperature and fluorescence were examined but the steps in the fluorescence are hard to relate to variations in temperature which was filtered so the steps are gone. No alignment adjustment will be made.

Oxygen

The dissolved oxygen channel was aligned earlier in processing, but was checked again. As usual, the results vary from feature to feature as local vertical gradients and descent rates affect the alignment. No further shift will be applied.

##### 12. DELETE, CLIP and CTDEDIT

Plots were made to find out how many records needed to be removed to eliminate the initial soak.

For CTD 1123 the number varied, but for all the CTD 0404 casts removing 1000 records worked well.

CLIP was run on the SHFC0 files for CTD #1123 and on the SHFpH files for CTD #0404.

DELETE was run on all casts using the following parameters:

Remove Surface Records in the top 10m.

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 9 points for CTD #1123 and over 5 points for CTD #0404) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

The DEL files were copied to *.EDT. All files were examined in CTDEDIT and some editing was applied to 23 casts. This mostly involved removing some near-surface records and a few records from the bottom of the cast.  There was also some cleaning of salinity.

14. Inter-comparisons

Sensor History – There is no history available for the sensors on CTD #1123. CTD #0404 has been used on 2 other 2016 cruises but there was no calibration sampling.

Comparison of repeat casts –There were no repeat casts.

Historic Ranges – The local climatology is a 3-standard deviation climatology covering 3 months and a fairly large region, so it is considered a severe criterion for casts near shore. Nonetheless, profile plots were examined with the climatology superimposed.

- Temperatures were above the climatology maximum below 100db for all the casts from the northern part of the cruise (events 32-84). Similar observations have been made during other cruises in the same area in 2016. Excursions from climatology in this area are likely real and not due to sensor calibration problems.

- The only salinity values that fell outside the climatology were some low values between 20 and 30db at 1 cast in the Broughton Islands area.

Post-cruise calibrations – None available.

15. CALIBRATE

There was no calibration sampling and no useful history to enable estimates. Pressure appears to be ok with the offset applied in conversion. No recalibration was applied.

16. Fluorescence Filter

This step will be skipped as the fluorescence is not very spiky.

17. Bin Average and REMOVE

The files were bin averaged using 1db bins.

T-S plots were examined and no further editing was found necessary.

REMOVE was run to remove Scan_Number, Conductivity:Primary, Oxygen:Voltage, Descent Rate and Flag channels from all casts.

Dissolved oxygen% surface saturation was calculated and plotted. Near-surface values ranged from 60% to 140% with values in the open parts of the Strait of Georgia tending to be between 95% and 115%. The lowest values came from casts that were well-mixed in the top 15m so oxygen had probably been mixed downwards and the highest values were seen in casts with high near-surface vertical DO gradients. . There is insufficient information to comment on DO sensor accuracy.

18. HEADER EDIT and final checks.

CLEAN was run to replace a few slightly negative fluorescence values with pad values.

Header Edit was used to fix headers, fix formats and channel names and to add the following notes to the headers:

For CTD #01123

Data Processing Notes:

Fluorescence and Oxygen:Dissolved:SBE channels are nominal and unedited

except that some records were removed in editing temperature and salinity.

While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly

for casts far from shore.

There was no salinity or dissolved oxygen calibration sampling and no other

information available upon which to base a recalibration.

For details on the processing see the report: 2016-29-proc.doc.

For CTD #0443

Data Processing Notes:

----------------------

Fluorescence, pH and Oxygen:Dissolved:SBE channels are nominal and unedited

except that some records were removed in editing temperature and salinity.

While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore.

There was no salinity, pH or dissolved oxygen calibration sampling and no other

information available upon which to base a recalibration.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration

data are available. Calibration is required for each cast to get absolute

values, although general trends within a cast are likely real.

For details on the processing see the report: 2016-29-proc.doc.

---------------------------------------------------------------------------------

A cross-reference listing was produced.

A header check was run on the CTD files. A problem was found in the header identification of the CTD used for the first 4 casts. This was fixed using a text editor.

The problem derives from the original converted files which do not contain this information. The format of the raw files was XML, not HEX.

The sensor history was updated.

Profile and T-S plots were examined and no problems were found, though it is noted that the Puget Sound pH values were higher than those from the Strait of Georgia.

Particulars

3. Winch speed issues. Part of upcast missing.

16. Upload failed part way through – have full downcast but only 10m of upcast.

32. Change of CTD to #0404

127. CTD not switched on – no data logged.

## CRUISE SUMMARY

| Cruise ID#: 2016-29 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Dates: Start: 4 July 2016 End: 14 July 2016 |  |  |  |  |
| Location: Strait of Georgia |  |  |  |  |
| Party Chief: Neville C. |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? |
| 1 | SEABIRD | 25+ | 1123 | No |
| 2 | SEABIRD | 25 | 0404 | No |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25plus/1123	Cruise ID#:		2016-29

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4888 | 7Nov2015 | Factory |  |  |
| Conductivity | 4513 | 28Oct2015 | Factory |  |  |
| ECO Fluorometer | 2216 | ? | ? |  |  |
| SBE43 Oxygen | 3234 | 3Nov2015 | Factory |  |  |
| Pressure | 1123 | 11Nov2015 | Factory |  |  |

Make/Model/Serial#:	SEABIRD/SBE25/0404	Cruise ID#:		2016-29

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2095 | 1Apr2011 | Factory |  |  |
| Conductivity | 1764 | 3Jan2014 | Factory |  |  |
| pH | 0852 | 18Jan2016 |  |  |  |
| ECO Fluorometer | 4185 | 2Nov 2015 | ? |  |  |
| SBE43 Oxygen | 1176 | 19Dec2015 | Factory |  |  |
| Pressure | 482 | 4Jan2016 | Factory |  |  |