## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2016-30

Agency: PBS, Ecosystem Science Division, Nanaimo, B.C.

Party Chief: Neville C.					Location: Strait of Georgia/Puget Sound

Project: SoG Juvenile Salmon Trawl Survey		Platform: Frosti

Date: 17 October 2016 – 26 October 2016

Processed by: Germaine Gatien

Date of Processing: 3 February 2017 – 6 March 2017

Number of original HEX files:	30		Number of CTD files:  30

# INSTRUMENT SUMMARY

SeaBird Model SBE-25 CTD (s/n 0404) was used with temperature sensor #2095, conductivity sensor #1764, Wetlabs ECO Fluorometer #4185, pH sensor #0852, oxygen sensor #1176 and pressure sensor 0482.

# SUMMARY OF QUALITY AND CONCERNS

Copies of pages from the log book were available, but there was no equipment list or list of personnel. The log times were in UTC. A spreadsheet was provided with positions, station names, water depths and CTD serial numbers; this was very helpful since position information is not in the headers. The file names did not correspond to event numbers in the log, but they were assigned sequentially. The names were changed so that the final 2 digits correspond to the event numbers in the log.

There was an initial soak at 10m followed by a return to the surface before the full cast was run. The data from the initial soak were included in the raw data, so had to be removed so that they would not be included in the final data.

The CTD had poorer pressure resolution than expected from a SeaBird 25. This was also noted during 2016-58 and 2016-40 and the CTD technician was informed that there may be a problem with the settings for the instrument. It has been sent to the factory for service.

The raw files for casts 116-137 were missing but were later found on the ship computer.

The fluorometer produced some slightly negative values at the bottom of casts. Based on the results of cruise 2016-40 a correction was made to the fluorescence offset.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data are available. Calibration is required for each cast to get absolute values, although general trends within a cast are likely real.

There was no calibration sampling.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea.

2. Preliminary Steps

CTD #0404 used for parts of cruises 2016-31, 2016-58 and 2016-29 with the same sensors and for all of 2016-40.

The Daily Log was obtained.

The raw files for casts 116-137 were missing but were later found on the ship computer.

The cruise summary sheet was completed.

A single configuration was used. The calibration parameters were checked and no errors were found, though no confirmation was found for the pressure calibration.

## 3. Conversion of Raw Data

All casts were converted using the appropriate configuration files.

The file names are in sequential order of CTD casts, but do not correspond to the event numbers in the log. The names were changed after conversion so that the last 2 digits correspond to the log events.

The pressure at the beginning and end of casts was about 0.2db so there is no evidence of a problem in pressure calibration. The pressure is full of steps of ~1.2db. This was noted on other 2016 cruises using this CTD. The CTD technician was alerted about this problem.

Plots show that the channels all have reasonable profiles except for descent rate which is extremely noisy, presumably because of the odd pressure steps.

The data from 10m soak are included, so will need to be removed before running DELETE.

The times in the log are in UTC.

As was found for other recent cruises using CTD #0404, the times in the headers are in PDT but a day ahead, so actually 17 hours ahead. So 17 hours will need to be subtracted later.

4. WILDEDIT

This step was skipped as it does not usually work well on SBE25 data and no obvious spikes were noted.

5. FILTER

For other recent uses of this CTD, the pressure looked steppy for all widths tested using WFILTER but the filter worked well on the temperature and conductivity and the resulting salinity looks improved. The pressure readings from this sensor are odd with the steps being ~1.2db, much grosser resolution than expected from an SBE25. This same problem was noted in 2015. WFILTER did improve the pressure but there were still many reversals until the setting was increased to width 9 for pressure.

WFILTER was run with a width of 9 for pressure, and a width of 5 for temperature and conductivity.

6. ALIGNCTD

As for other recent uses of this equipment a setting of +2.5s improves the alignment of dissolved oxygen with temperature. ALIGNCTD was run on all files using +2.5s.

7. CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25. CELLTM was run on all casts using that setting and it did improve correspondence of upcast and downcast data in T-S space.

8. DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

CLEAN was run to replace pad values in the pressure channel with interpolated values and to add event numbers to the headers.

Time was adjusted using program ADD TIME CHANNEL to subtract 17 hours.

10. Checking Headers

Some header information was missing: latitude, longitude, station name and water depth.

A spreadsheet was provided and that was used to prepare file 2016-30-header_mrg.csv. IOS SHELL program “MERGE:CSV file to headers” was used to add information from the csv file to the IOS files with output MRH.

A track plot was produced and 1 position was obviously wrong. The spreadsheet differs from the log and when the log value was substituted in the file, the plot looks fine. The spreadsheet was corrected.

A cross-reference listing was produced and checked against the log entries.  The only error was the station name for event #89 which was Set 52 in the file but set 53 in the log. There was no CTD cast during Set 52, so the file was changed to match the log book entry.

A header check was also run and the only problems noted are that there are some negative values in fluorescence (minimum -0.275). There were a few casts with pressure -1.075db that were single-point spikes at the surface. Other than those values the lowest pressure was -0.3db which was seen after the pump was turned off and conductivity was close to zero.

The surface check was run and the average surface value was 0.76db; most of the surface values are associated with near-zero conductivity.

CLIP was used to remove data collected during the initial soak; the number of records removed varied.

##### 11. SHIFT

pH

Tests were run using shifts of +10 to +30 records. A shift of +20 records did the best job overall of bringing the upcast and downcast pH traces into alignment with temperature traces, though it is hard to judge because the pH traces cannot capture sharp changes.

SHIFT was run on all casts to advance pH:SBE by +20 records.

Conductivity

As was found the last time CTD #0404 was used, all shifts applied to conductivity produced unstable features in the T-S plots. No adjustment was applied.

Fluorescence

The fluorometers on the SBE25s were not pumped, so a shift in alignment is expected to be small or unnecessary. Profile plots of temperature and fluorescence were examined but the steps in the fluorescence are hard to relate to variations in temperature which was filtered so the steps are gone. No alignment adjustment will be made.

Oxygen

The dissolved oxygen channel was aligned earlier in processing, but was checked again. As usual, the results vary from feature to feature as local vertical gradients and descent rates affect the alignment. No further shift will be applied.

##### 12. DELETE and CTDEDIT

DELETE was run on all casts using the following parameters:

Remove Surface Records in the top 10m.

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

The DEL files were copied to *.EDT. All files were examined in CTDEDIT and some editing was applied to 27 casts. This mostly involved removing records that appear to be corrupted by shed wakes.  There was also some cleaning of salinity.

14. Inter-comparisons

Sensor History –CTD #0404 has been used on 3 previous 2016 cruises but there was no calibration sampling. It was also used during a later cruise with salinity and extracted chlorophyll sampling. The CTD salinity appeared to read slightly higher than bottles but there was too much uncertainty about the firing depth of the Niskin bottles to justify recalibration.

Comparison of repeat casts –There were no repeat casts and nearby casts are shallow and/or well separated in time.

Historic Ranges – The local climatology is a 3-standard deviation climatology covering 3 months and a fairly large region, so it is considered a severe criterion for casts near shore. Nonetheless, profile plots were examined with the climatology superimposed.

- Temperatures were mostly within the climatology except in the northern part of the Strait of Georgia where there were some high temperatures between about 40 and 70db.

- Salinity was occasionally high around Texada Island and low near the surface.

Excursions from climatology in this area are likely real and the excursions were not systematic, so they are not likely to be due to sensor calibration problems.

Post-cruise calibrations – None available at the time of processing.

15. CALIBRATE

During the cruise that followed (2016-40) there was extracted CHL sampling and there were many negative values found in CTD fluorescence. For that cruise the CTD fluorescence was corrected by adding 0.275ug/L to the fluorescence. This looks appropriate for this cruise too.

CALIBRATE was run to add 0.275ug/L to all fluorescence values in the EDT files.

16. Fluorescence Filter

This step will be skipped as the fluorescence is not very spiky.

17. Bin Average and REMOVE

The files were bin averaged using 1db bins.

T-S plots were examined and no further editing was found necessary.

REMOVE was run to remove Scan_Number, Conductivity:Primary, Oxygen:Voltage, Descent Rate and Flag channels from all casts.

Dissolved oxygen% surface saturation was calculated and plotted. Near-surface values were quite low, ranging from 58% to 98%. The lowest values were in Haro Strait and a few casts in Puget Sound where waters were very well mixed and low saturation is expected. All other values were between about 85% and 98%.  In many cases there were no data above 3m. This is significant if the surface waters are not well-mixed. The highest value was in Howe Sound and had a fairly high near-surface gradient. While it is 97.5% at 1.5m it is only 93% at 4m, so it may well be closer to 100% at the surface.

This type of sensor tends to drift towards lower values, so there is reason to expect the dissolved oxygen to read a little low. It is tempting to recalibrate by multiplying by a factor ~1.02 or 1.03. However, the sensor was also used during cruise 2016-40 which immediately followed this cruise. There were some offshore casts during 216-40 that showed surface saturation rates of ~103% which is a typical value for those conditions. So no correction will be made.

18. HEADER EDIT and final checks.

Header Edit was used to fix headers, fix formats and channel names and to add the following notes to the headers:

Data Processing Notes:

----------------------

Fluorescence, pH and Oxygen:Dissolved:SBE channels are nominal and unedited

except that some records were removed in editing temperature and salinity.

While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore.

There was no salinity, pH or dissolved oxygen calibration sampling and no other

information available upon which to base a recalibration.

During the cruise that followed and included extracted chlorophyll sampling it

was determined that a small correction to the offset for the fluorometer was

appropriate to remove many small negative values.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration

data are available. Calibration is required for each cast to get absolute

values, although general trends within a cast are likely real.

For details on the processing see the report: 2016-30-proc.doc.

---------------------------------------------------------------------------------

A cross-reference listing was produced.

A header check was run on the CTD files. No problems were found.

The sensor history was updated.

Profile and T-S plots were examined and no problems were found.

Particulars

132. bottom depth should be 169m, not 179

## CRUISE SUMMARY

| Cruise ID#: 2016-30 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Dates: Start: 17 October 2016 End: 26 October 2016 |  |  |  |  |
| Location: Strait of Georgia / Puget Sound |  |  |  |  |
| Party Chief: Neville C. |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? |
| 1 | SEABIRD | 25 | 0404 | No |

CTD CALIBRATION INFORMATION

ake/Model/Serial#:	SEABIRD/SBE25/0404	Cruise ID#:		2016-30

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2095 | 1Apr2011 | Factory |  |  |
| Conductivity | 1764 | 3Jan2014 | Factory |  |  |
| pH | 0852 | 18Jan2016 |  |  |  |
| ECO Fluorometer | 4185 | 2Nov 2015 | ? |  |  |
| SBE43 Oxygen | 1176 | 19Dec2015 | Factory |  |  |
| Pressure | 482 | 4Jan2016 | Factory |  |  |