## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2016-41

Agency: Ocean Sciences Division

Location: Douglas Channel

Project: WC – Douglas Channel

Party Chief: Johannessen S.

Platform: John P. Tully

Date: May 16, 2016 – May 23, 2016

Processed by: Germaine Gatien

Date of Processing: 28 July 2016 – 17 August 2016

Number of original HEX files: 25

Number of CTD files: 25	Number of bottle files:	22

Number of TSG files: 0

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0443) was used for this cruise. It was mounted in a rosette and attached were a WetLabs CSTAR transmissometer (#1185DR), a SBE 43 DO sensor (#0997) on the secondary pump, a SeaPoint Fluorometer (#3640) on the primary pump, a pH sensor (#0692) and an altimeter (62354).

The data logging computer was the Tully CTD Laptop (Acer).

The data acquisition program was Seasave.

The CTD deck unit was an SBE model 11+, #425.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log book was generally in good order. There were many notes about spooling problems on the early casts.

Comparison of CTD salinity to bottle samples suggests that the CTD salinity is low, but it is likely that the bottle values are high for 2 reasons. The samples were analyzed about 7 to 8 weeks after collection so there is likely some evaporation and/or adsorption of samples, which would both raise salinity. Incomplete flushing of Niskin bottles is also likely, due to very steady descent rates; this would lead to the Niskin bottles containing water from lower in the water column where salinity is higher than the ambient salinity at the bottle stop level. Recalibration of salinity was based on the results of cruise 2016-47 which immediately followed this cruise, had the same sensors and many bottle samples.

The comparison of titrated dissolved oxygen samples with data from the SBE Dissolved Oxygen sensor indicate that the sensor data are closer to bottle samples than expected based on the history of this sensor. This is likely the result of poor flushing since many samples will have come from lower in the water column, thus will have DO values that are lower than ambient waters at the bottle stop level. The errors will be especially high where the vertical DO gradient is large. While the sensor readings are fairly close to the samples, they are likely lower than samples from the correct depth. Recalibration was based on the comparison with bottles during 2016-47 when the same sensor was used.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5 mL/L from 0 to 50db

±0.2 mL/L from 50db to 200db

±0.1 mL/L below 200db

The fluorescence gain setting entered in the configuration file used at sea was 30X. The log book indicates it was 10X and during the cruise that followed it was found that the configuration file was incorrect and it was changed from 30X to 10X. The data were processed assuming a 10X cable was in use and the values look reasonable. There was no extracted chlorophyll sampling.

The pH channel was removed from events #12, 13 and 14 because downcast values started low and gradually increased until they looked believable by about 150m.  This is likely because the soaker bottle was not placed on the sensor between casts. This is the second 2016 cruise during which these odd pH profiles have been noted. It is important that the sensor be hydrated unless there is only a very brief stop between casts. Judging by these casts, a 1.5hour period out of water is long enough to cause serious problems with this type of sensor.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data were available at the time of processing.  Calibration is required for each cast to get absolute values, although   general trends within a cast are likely real.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. There were many comments about issues with spooling, but an adjustment during event #14 appears to have solved the problem.

Nutrients, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked.

The XMLCON files did not change through the cruise. The calibration constants were checked for all instruments and there were some errors:

- The gain setting for the fluorometer was entered as 30X whereas the log indicates it was 10X. During the cruise that followed it was discovered that the configuration file was wrong. The 10X cable was in use.

- The calibrations entered for the pH sensor were out of date. There was a January 2016 factory calibration available.

- While not an error, the pressure sensor was found to be reading too high when used during 2015-54 and an offset of -0.6 was used in recalibration. That cruise frequently had the CTD right at the surface, so there was good evidence available. So for this cruise an offset of -0.5476db (factory offset +0.05243db – adjustment 0.6db) should be applied. There was also evidence from 2016-47 which followed this cruise that the pressure adjustment of -0.6db was reasonable.

The corrected file was saved as 2016-41-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2016-41-ctd.xmlcon to create CNV files.

A few casts were examined.  All expected channels are present. The primary and secondary temperature and conductivity channels are reasonably close during downcasts, but as usual the upcasts differ more due to noise in both channels. There are some spikes in both conductivity channels – these are reduced by running WILDEDIT. There were obvious problems during bottle stops with pressure increasing significantly during a stop. This stops after event #14 with only minor changes in pressure seen thereafter. Dissolved oxygen, altimetry and transmissivity profiles all look normal. The descent rate was generally fairly high and varied from very steady to a little noisy.

The downcast pH traces for some casts are very different from upcasts. This will be investigated later.

It is impossible to determine if the fluorescence values are realistic because we have no chlorophyll samples, but given it was May and the CTD technician believes that the gain cable was not changed between this cruise and the next one and the log indicates the gain was 10X, we have to assume that the gain is 10X. The values found seem reasonable, but those obtained with the assumption of 30X are not clearly bad. There are some off-scale fluorescence values.

The surface pressure was checked and acquisition generally starts at about 0.7db – if the pressure offset were not applied, this would be about 1.3db. The salinity looks reasonable for either depth. There are no notes in the log about the deployment method. There are lots of salinity samples including some for event #41 where the descent rate indicates some ship movement to shake up the Niskin enough to help flush the contents so that the comparison with CTD salinity may help establish whether the pressure offset was appropriate. However, this is also one of the few casts with a well-mixed surface layer which means a small pressure difference is likely not significant.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2016-41-ctd.xmlcon.

The ROS files were converted to IOS format.  They were put through CLEAN to create BOT files; off-scale fluorescence values were replaced with pad values.

Temperature and salinity were plotted for all BOT files to check for outliers and no problems were seen that required editing.

A preliminary header check was done and no problems were found. Fluorescence had gone off-scale, but the CLEAN routine fixed that.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv.

Sample numbers were added to the file based on the rosette log records.

For event #19 the log indicates there were 5 bottles fired but there were actually 6 bottles fired. A note in the log says that Niskin #4 was accidentally fired at 30m instead of 20m, so Niskin #6 was fired at 20m. The samples should be fine. There are no samples from Niskin #4, although the note on the log sheet suggests the 20m sample may be labelled that way. There may be some trouble merging these files if particle data are to be added to it.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine. Those files were bin averaged on bottle number.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2016-41-bot-hdr.txt which will be updated as needed during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2016-41oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2016-41oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2016-41SAL.xls. The analysis was done within 51 to 56 days of collection. The files were simplified and saved as 2016-41sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2016-41nuts.xls.

Then the file was simplified, reordered on sample numbers and saved as 2016-41-nuts.csv. The file was converted to individual NUT files.

All files for event #41 were reordered on increasing sample number.

The SAL, OXY and NUT files were merged with CST files in 3 steps.

After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. A number of problems were found:

- The DO for casts #6 and 7 were misidentified as #3 and 4.

- The nutrients for cast #20 were misidentified as #19.

- 3 CHL samples are indicated on the rosette log but could not be found.

CLEAN was run on the MRG files to add 0 flags to empty flag channels and to update header limits. (MRGCLN2)

##### Compare

Salinity

Compare was run with pressure as reference channel.

There is a lot of scatter in the comparison. All of the major outliers are from the top 10m where vertical salinity gradients were extremely high.

Both CTD salinity channels read lower than the bottles on average. When bottles above 200m and those fired close to the bottom are excluded, the primary salinity is low by an average of 0.0084 and the secondary is low by 0.0065, with standard deviations being 0.003 and 0.004 respectively. There are only 12 bottles from 3 casts in that comparison. Looking at bottom bottles below 200m excluding 2 outliers, the primary salinity is low by 0.0014 and the secondary high by 0.0011. These are cases where poor flushing would have the opposite effect as on upcast bottles. Near the surface these differences are larger which is likely caused by poor flushing in the presence of high vertical salinity gradients.

Another explanation for the CTD salinity looking low would be evaporation or adsorption of samples due to the 7 to 8 week wait for analysis. This could lead to bottle salinity being high by roughly 0.002 to 0.005.

There were no duplicates, but there were 6 casts during which 2 bottles were fired at the same depth and salinity samples taken from both. These came from casts 13, 20, 22, 29, 41 and 42. The first 3 had a very steady descent rate indicating calm waters, the 4th and 5th had just a little noise in the descent rate while the 6th had a little more noise so that we might expect better flushing of bottles. However, the 6th case also had higher local gradients so flushing errors would be exaggerated.

The differences between bottles at the same depth were somewhat lower for cast #42 and the CTD salinity is closer to bottle values for that cast. This could be evidence that flushing was better when there was more ship motion, especially if we allow for larger differences due to a relatively high vertical salinity gradient for that cast. But the evidence is weak.

For full details for the COMPARE run see file 2016-41-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There are no noteworthy outliers, though at around 10m there are very high vertical DO gradients and large changes during bottle stops. No further quality flags are suggested.

There are no bottle samples with DO<1.9mL/L, so getting a fit with a reliable offset is a problem.

- With the offset free

CTD DO Corrected = CTD DO * 1.0066 +0.0652

That correction looks lower than expected and may reflect flushing problems.

- With the offset set to zero

CTD DO Corrected = CTD DO * 1.0202

That slope is slightly lower than those found in other recent cruises.

- When only the casts with noisier descent rate (31-42) are used the fit is:

CTD DO Corrected = CTD DO * 1.0167 +0.0346

Or CTD DO Corrected = CTD DO * 1.0230

But there are not many bottles in those fits.

The result from 2015-21 in Sept. 2015 when there was good sampling with a wide range of DO values was:

CTD DO Corrected = CTD DO * 1.0225 + 0.0306	(2015-21)

A break was taken from this processing job to examine the cruise that followed and had many DO samples including many of low value, and where incomplete flushing was likely to have only a minor effect.  The fit from 2016-47 in May 2016 was:

CTD DO Corrected = CTD DO * 1.0208 +0.0377

The correction derived from the 2016-47 fit is similar to 2015-21 and the slope is very close to that for this cruise when the offset is forced to zero. It is also very close to the results from 2015-54 in October 2015 so it looks like a good choice for this cruise.

For more detail see document 2016-41-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. The only outliers are from the high gradient zones where they are inevitable.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on 2 casts with no bottles fired but it was very hard to judge which setting was best for aligning DO with temperature since both profiles were complex. A setting of between 2.5s and 3s looked best overall. On all other cruises using that DO sensor since it was last serviced a setting of +3.5s was used.

ALIGNCTD was run on all casts using +3s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. Tests on previous cruises using these sensors showed the default setting of (α = 0.0245, β=9.5) did the best job and it does improve the data for both conductivity channels for these data.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. While none are very deep, it does allow us to check for sudden changes. Differences from earlier cruises using the same equipment are also shown with dark shading.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2015-10-0099 | 500 | -0.0011 | -0.00016 | -0.0009 | High, Very Noisy |
|  | 1000 | -0.0010 | -0.00006 | +0.0004 | “ |
|  | 1900 | -0.0013 | -0.00003 | +0.0010 | “ |
| 2015-19-0097 | 350 | -0.0004 | -0.00003 | +0.0008 | High, V.Steady |
|  | 400 | -0.0003XN | -0.00006VN | +0.0015VN | “ |
| 2015-54-0057 | 450 | -0.0011 | +0.00003 | +0.0013 | High, Moderate |
| 2015-54-0373 | 450 | -0.0008 | +0.00025 | +0.0033 | High, Steady |
| 2016-41-0001 | 410 | -0.0015 | +0.00010 | +0.0024 | High, Steady |
| 2016-41-0013 | 410 | -0.0016 | +0.00016 | +0.0026 | F High, V Steady |
| 2016-41-0033 | 410 | -0.0013 | +0.00012 | +0.0026 | High, Noisy |
| 2016-47-0039 | 1000 | -0.0019 | +0.0001 | +0.0032 | High, V Noisy |
|  | 1900 | -0.0018 | +0.0002 | +0.0040 | “ |
| 2016-47-0069 | 400 | -0.0022VN | -0.00004 | +0.0022VN | High, V Noisy |
|  | 1000 | -0.0019 | +0.0001 | +0.0036 |  |
|  | 1900 | -0.0017 | +0.0001 | +0.0039 | “ |

The differences are small, though slightly larger than for 2015-10 and 2015-19. They are intermediate between the 2 casts included from 2015-54 and smaller than 2016-47 except for conductivity.  However, differences at 400m tend to be noisy, and were especially so for 2016-47.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number and to replace off-scale fluorescence values with pad values. The off-scale criteria was chosen as FL>14.816ug/L based on examination of plots.

##### Checking Headers

A cross-reference list was checked against the log book. The only error found was in the station name for event #34 which was changed from SCC74 to CMP1.

A header check was run. There are obviously some bad pH values, so the profiles were examined. There were spikes in pH in 4 casts (21, 28, 29, 33), but they were all in upcast sections except for cast #28 which had a bad section around 105.5db. 15 values were replaced with pad values.

It was also noted that casts 12, 13 and 14 have pH values that are too low at the beginning of the downcasts, but gradually increase until they look believable by 150m.  It appears that the soaker bottle was sometimes not placed on the sensor between casts. This is the second 2016 cruise during which these odd pH profiles have been noted. It is important that the sensor be hydrated unless there is only a very brief stop between casts. Judging by these casts, a 1.5 hour period out of water is long enough to cause serious problems with this type of sensor.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 1.34db which is reasonable for this area. There were very low salinity values at 1db for casts near the head of Gardiner Canal, but that looks reasonable; the CTD was clearly in the water based on other variables.

The altimeter and depth readings from the headers of the CLN and SAMAVG files were exported to spreadsheets to check that they are reasonable. A rough check was made by subtracting the altimeter reading from the depth entered in the header to find the maximum depth sampled. That was then compared with the maximum pressure recorded in the files using a rough estimate that depth is ~99% of pressure. When those two estimates differed by more than 5m, a plot was made to see if the altimetry reading was reasonable. Despite many spikes near the bottom, the header altimetry entries look reasonable except for cast #28 for which the header entry of 4.6m looks like the result of a spike; 8.9m is a better estimate, so the header entry was changed and a note of explanation added to the header. There were also a few cases (1, 2, 7, 18 & 30) where the log book entry for depth looks likely to be better than the one entered in the headers, at least for the time when the CTD was at the bottom of the cast. Those depths were changed in the CLN files. The same changes were made to the bottle files (SAMAVG) except that there was no altimetry header for cast #28 since there were only spikes at the time the bottom bottle was closed.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and as usual, there is a lot of variability with the up and downcast traces sometimes closer than temperature and sometimes further apart. This is likely due to varying vertical gradients. Overall the choice made earlier looks appropriate, so no further alignment will be applied.

Conductivity

Tests were run on 4 casts using a variety of shifts and the results varied from cast to cast and feature to feature with only small differences among them. The choice for the last 2 cruises when these sensors were used was -0.2 records for the primary and -0.8 records for the secondary. Those settings look reasonable for this cruise as well.

SHIFT was run twice on all casts using those settings.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. The last time the sensor was used the shift applied was +50 records but that looks a little high for these data. A setting of +30 records was found to work best.

SHIFT was run on all casts using that setting.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

The DEL files were copied to *.EDT.

##### Other Comparisons

Previous experience with these sensors –

SALINITY: The T/C sensors were used during 2015-10, 2015-21, 2015-19 and 2015-54. The primary and secondary salinity were both found to be low by~ 0.002 for 2015-10. They were low by 0.0045 and 0.0032, respectively, for 2015-21, but delayed salinity analysis reduced the reliability of those results. For 2015-19 salinity was low by 0.0051 and 0.0076 but analysis was delayed and salt crystals were found on some bottles and incomplete flushing of Niskin bottles was likely a problem. During 2015-54 salinity was found to be low by 0.008 and 0.0055, but it was likely that Niskin bottles hadn’t flushed well.

PRESSURE: During 2015-54 there was sampling close to the surface and pressure was found to be high by 0.6db.

DISSOLVED OXYGEN: The last good comparison of DO with calibration samples was from cruise 2015-21 in Sept 2015.

Experience with these sensors after this cruise

There was a cruise after this one that used the same sensors and showed that the primary salinity was low by about 0.006 and the secondary by about 0.002. The pressure correction of -0.6db was found to be appropriate. The dissolved oxygen fit was similar to 2015-61 and 2015-54 and was based on many bottles include many with low DO values.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Local climatology was only available for 5 casts. All salinity data were within the local climatology except for some low salinity values in the top 10m of station HAK1. Temperatures were mostly within the climatology except for the 2 HEC1 casts where there were small excursions towards high temperature at 50m and the deeper of the 2 casts had slightly low temperature at the bottom. These look like limitations in the climatology which likely has few or no data from these sites.

These excursions are not considered evidence of calibration problems.

Repeat Casts – There were some repeat casts but they were too shallow to expect good repeatability.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The primary channels were selected for archiving during 2015-10 and 2015-21 and the secondary were chosen for 2015-19 and 2015-54.  The secondary channels have slightly less spiking in these files, so they were selected for editing and archiving.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. Only 1 file (#6) required no editing.

##### Initial Recalibration

The pressure offset used in processing looks appropriate.

The salinity comparison is not trusted due to delayed analysis and many casts had a very steady descent rate that may have reduced flushing of Niskin bottles.  The history of comparisons for these sensors on other cruises is also hard to interpret because there were different depths sampled, different waits for sample analysis and different areas where flushing efficiency would vary. Overall the differences between the two salinity and temperature channels appear to be increasing.

We could use the results from the cruise that followed for which both flushing problems and the wait for analysis were minimal. The differences between the 2 salinity channels were larger for 2016-47 (~0.0038) than for 2016-41 (~0.0026). That may be partly because the 2016-47 sampling was much deeper and there is often some pressure-dependence in differences.

For 2016-47 the temperature difference was felt to be part of the problem, and it is almost as large for this cruise. Correcting the primary temperature by subtracting 0.0015 and recalculating primary salinity would bring the primary salinity closer to bottles by ~0.0015. The salinity for 2016-47 was further corrected by adding 0.004 to the primary and 0.002 to the secondary. This looks like as good a recalibration scheme as we can find for this cruise.

CALIBRATE was run using file 2016-41-recal1.ccf to subtract 0.0015 from the primary temperature, to recalculate primary salinity and to correct the Oxygen:Dissolved:SBE channel in the SAM and MRGCLN2 files using:

CTD DO Corrected = CTD DO * 1.0208 +0.0377

Then a second run of CALIBRATE was made using file 2016-47-recal2.ccf to add 0.004 to the primary salinity and 0.002 to the secondary salinity.

COMPARE was rerun for salinity and dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly. The two salinity channels are lower than bottles by ~0.003 which is likely due to the bottle samples having been affected by evaporation and/or adsorption and/or incomplete flushing of Niskin bottles. The average of differences in the fit of (DO:BOT-DO_CTD) vs DO:CTD once cases were removed with large standard deviation in CTD DO, was +0.04mL/L. When only bottles with DO<3mL/L are included the average is high by +0.008mL/L. This is in line with expectations that flushing of Niskin bottles was incomplete and in areas of high DO gradient the result is that the CTD DO will look too high. This effect is much smaller at depth where DO gradients are lower.

See file 2016-41-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. As expected the CTD DO is higher than bottles in most cases. When differences >0.02 are removed the CTD is higher by an average of 0.08mL/L. When only bottles below 300m that were not fired at the bottom are included, the average is +0.049mL/L. When only bottles below 300m that were fired at the bottom are included, it is -0.001mL/L. The bottom bottles probably have DO values that are a little too high due to incomplete flushing though the vertical DO gradients should be pretty low there. Since the CTD DO is very close to those bottles, it is likely reading a bit too high as well. This is consistent with the slow response of the DO sensor. For the bottles fired during the upcast, the CTD is reading much higher, which is likely due to poor flushing as bottles would contain deeper water with lower DO. The recalibrated values look reasonable.

No further recalibration of DO is justified.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots look fine. Profiles look ok except for the bad pH in casts 12, 13 and 14 that was noted earlier, and there are a few small spikes at depth in channels for which that is unlikely. It will be left for researchers to determine what to do with those.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

Also, pH was removed from events #12-14.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and pH are nominal and unedited except

that some records were removed in editing temperature and salinity.

The deployment method used for the CTD did not include a soak at 10m.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5 mL/L from 0 to 50db

±0.2 mL/L from 50db to 200db

±0.1 mL/L below 200db

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration

data were available at the time of processing.

Calibration is required for each cast to get absolute values, although general

trends within a cast are likely real.

For details on the processing see the report: 2016-41-proc.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and problems were found.

A file list was produced.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values were highly variable ranging from 95% to 140%, with the highest values in the upper section of Douglas Channel and lowest values in the southern Douglas Channel. Hecate Strait had values between 110% and 120%. These values seem reasonable for this region and season, but the evidence is too weak to suggest how accurate the DO data are.

##### Final Bottle Files

CALIBRATE was run twice using the same scheme as for the full data files. Then the MRGCOR2 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

The pH channel was removed from events #12-14 only.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

Data from the MRGREO files were exported to a spreadsheet and compared with rosette sheets. The only problem found was that the dissolved oxygen and salinity for sample #296 had been flagged 2 because the Niskin bottle was fired without the usual 30s wait, but the nutrient sample was not flagged. The same flag was added to the nutrient sample and to the nutrient spreadsheet for consistency.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check and a header check were run on all files and no errors were found.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in. No problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars

1. Some pauses due to wire spooling issues

2. Descending slowly to reset the wire on the drum

6. Pause at ~220m - spooling issues

10. Return to surface slowly - spooling issues

11. Stopped at 89m upcast – spooling issues

12. Particles only cast

14. Paused at 54m and at 30m paid out 25m to correct the spooling issue

42. Bottles #13 and 14 fired without 30s wait.

CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 17Dec2014 | Factory |  |  |
| Conductivity | 2280 | 19Dec2014 | Factory |  |  |
| Secondary Temp. | 2663 | 15Jan2015 | Factory |  |  |
| Secondary Cond. | 2754 | 19Dec2014 | Factory |  |  |
| Transmissometer | 1185DR | 5Feb2014 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 17Jan2015 | Factory |  |  |
| pH | 0692 | 12Jan 2016 | IOS |  |  |
| SeaPoint Fluor. | 3640 | n/a |  |  |  |
| Pressure Sensor | 0443 | 17Dec2014 | Factory |  |  |
| Altimeter | 62354 | n/a |  |  |  |