## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2016-45

Agency: WVL, Marine Environment and Habitat Science, Vancouver, B.C.

Location: Strait of Georgia

Project: Coastal Ecosystem Research Program

Party Chief: Sutherland T.

Platform: Vector

Date: April 27, 2016 – May 2, 2016

Processed by: Germaine Gatien

Date of Processing: 4 July 2016 – 28 July 2016

Number of original HEX files: 38

Number of CTD files: 38	Number of bottle files: 21

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#1119) on the secondary pump, a SeaPoint Fluorometer (#3865) on the primary pump with a 3X cable, a Biospherical QSP-400 PAR sensor (#4565), Surface PAR (#20518) a pH sensor (#0692) and an altimeter (#1253).

The data logging computer was Vector CTD Laptop.

The data acquisition program was Seasave v7.23.2.

The CTD deck unit was an SBE model 11+.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

There was no list of personnel in the log. Otherwise the Daily Science Log book and sampling/rosette log sheets were in good order.

The comparison of CTD and bottle salinity has a lot of scatter. Several bottles came from the bottom of casts where comparisons are often poor. Flushing of Niskin bottles was likely poor due to extremely calm waters as evidenced by a very steady descent rate. When outliers were removed from the comparison, both salinity channels were found to be low by an average of 0.02. The deepest sample showed the two salinity channels low by ~0.005. There was a delay of ~7.5 weeks in analysis of salinity samples which can lead to evaporation and adsorption that could easily increase salinity by 0.003 to 0.005. So the CTD is likely more accurate than it appears. The sensors have only been used for one other cruise which also had problems with flushing and delayed analysis. No recalibration is justified.

There were some spikes in the conductivity and salinity channels that were not seen in temperature. Sometimes they appear in both channels but often in only one. Where they involve only a few records salinity was cleaned using interpolation; they were removed if more extensive. These spikes are mostly fairly small and do not look like the results of poor alignment. There were also many near-surface unstable features seen in both temperature and salinity. Some of these could be real, or they could be due to ship effects, surface debris affecting plumbing or the wait at the surface may have been too short. Some data were removed using a graphical editor as they appeared to be bad; others were left as they were not clearly bad.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.7 mL/L from 0 to 25db

±0.25 mL/L from 25db to 200db

±0.06 mL/L below 200db

No TSG data was acquired due to a problem with the loop system.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data were available at the time of processing. Calibration is required for each cast to get absolute values although

general trends within a cast are likely real.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Nutrients, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The conductivity and DO sensors had been used for one other cruise, 2016-05, since they were last recalibrated; the DO sensor was used for 3 other cruises.

The XMLCON files did not vary through the cruise, and all parameters were correct. One file was saved as 2016-45-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2016-45-ctd.xmlcon to create CNV files.

A few casts were examined.  All expected channels are present. The primary and secondary temperature and conductivity channels are mostly reasonably close during downcasts, but as usual the upcasts differ more due to noise in both channels. There are some spikes in the conductivity channels that are not seen in the temperature or dissolved oxygen. Dissolved oxygen, fluorescence, PAR, pH, altimetry and transmissivity profiles all have the usual shapes. PAR maximum values were usually lower than Surface PAR, but the CTD was at about 2m when those values are found. The descent rate was mostly steady but fairly low.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2016-45-ctd.xmlcon.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and problems were seen in the primary salinity in 1cast and in the secondary salinity in 3 casts.

CTDEDIT was used to edit the BOT files for events 47, 58 and 88. Comments about editing were added to the headers.

The output files were then copied to BOT.

A preliminary header check was done and no problems were found. Fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

There were 7 bottles planned for event #17 but 8 were fired – 2 were fired at the surface, but no sample number was assigned to the 2nd bottle.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine. Those files were bin averaged on bottle number.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2016-45-bot-hdr.txt which will be updated as needed during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2016-45oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2016-45oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2016-45SAL.xls. The analysis was done within 64 to 69 days of collection. The files were simplified and saved as 2016-45sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2016-45nuts.xls.

Then the file was simplified, reordered on sample numbers and saved as 2016-45-nuts.csv. The file was converted to individual NUT files.

The SAL, OXY and NUT files were merged with CST files in 4 steps.

After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. Niskin #8 was fired during event #17 but no sample # was assigned and no samples taken, so that line was removed from the MRG file.

CLEAN was run on the MRG files to add 0 flags to empty flag channels and to update header limits. (MRGCLN2)

##### Compare

Salinity

Compare was run with pressure as reference channel.

There is a lot of scatter in the comparison. When bottles are removed for which the standard deviation in the CTD salinity is >0.0008 plus all bottles fired at the bottom, both salinity channels are low by an average of 0.02. However, the plot shows evidence of poor flushing. Some bottles were fired at the bottom and they all have bottle values lower than the CTD; this is to be expected as flushing errors at the bottom are of reverse sign as the water in the Niskins will come from higher in the water column. At other levels most bottles have higher salinity than the CTD. As well as flushing, such differences may arise due to delayed analysis of salinity samples. The delay was not long enough to lead to very large effects, but could certainly account for readings that are too high by 0.003 to 0.005

The deepest sample that did not come from the bottom was from 332db and for that one the primary and secondary salinity values were lower than the bottle by 0.0054 and 0.0049, respectively.

During 2016-05 the two CTD salinity channels were also close to each other; they were lower than bottles by an average of 0.007. That cruise had less evidence of problems with poor flushing but an even longer wait for sample analysis.

The only extreme outliers in the comparison can be explained by flushing problems; there is no evidence that any samples require quality flags.

For full details for the COMPARE run see file 2016-45-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The fit of (CTD DO – Titrated DO) vs CTD DO is:

CTD DO Corrected = CTD DO * 1.0065 +0.0867

or	CTD DO Corrected = CTD DO * 1.0086

The result is unusual as the slopes of corrections are usually larger and offsets smaller. However, during 2016-05 which used the same sensor a week earlier, similar results were found. For 2016-05 it was felt that the odd comparison was due to a combination of low vertical gradients so that response of the sensor was better than usual and poor flushing of Niskin bottles. The usual comparison leads to a linear correction with near-zero offset and the CTD DO reading lower than bottles by 2 to 4%.  But poor flushing of Niskin bottles leads to lower DO values in the titrated samples with the effect most dramatic in the higher gradients which are mostly from 5m to 25m. Given that most casts are shallow, this does not leave much to compare.

For 2016-05 the descent rate was generally quiet but there was some variety with casts in some areas more likely to have better flushing than in the casts for the current cruise. The result found during 2016-05 for casts with the steadiest descent rates was:

CTD DO Corrected = CTD DO * 1.0012 +0.0487

or	CTD DO Corrected = CTD DO * 1.0111

And casts with noisier descent rates (none were very noisy) had fits:

CTD DO Corrected = CTD DO * 1.0101 +0.0168

or	CTD DO Corrected = CTD DO * 1.0136

For the current cruise there are no casts with a noisy descent rate, but those with a little variability led to the fit:

CTD DO Corrected = CTD DO * 1.0138

For 2016-05 the fit

CTD DO Corrected = CTD DO * 1.015

was used to recalibrate the CTD DO and it looks like the best choice for this cruise as well though it may be an underestimate.

The most significant outlier had noisy CTD DO data and no other outliers were significant enough to require flagging of samples.

For more detail see document 2016-45-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. In these plots the outliers are typical of what is expected due to poor flushing of bottles. Bottom samples frequently look like the water came from higher in the water column. Above bottom the significant outliers suggest the Niskin contained water from lower in the water column.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

A few casts were checked before and after this step to see if spikes in conductivity were reduced and they were, though some spikes remain. So, the problems noted in the bottle files may be less serious in the profile files.

##### ALIGN DO

Tests were run on ALIGNCTD to determine the best setting to advance the DO signal. The results were difficult to judge because the casts are shallow with complex temperature and DO profiles. During the previous cruise a setting of 2.5s was found appropriate, but it looks too high for these data, likely because the descent rate was lower. Settings from 1.5s to 3s were tested and +2s looked best overall.

ALIGNCTD was run on all casts using +2s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. Tests on previous cruises using these sensors showed the default setting of (α = 0.0245, β=9.5) did the best job and it does improve the data for both conductivity channels for these data.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on the only cast that sampled as deep as 300m  to examine differences between sensor pairs. All the differences were very small; ∆Temp ~ -0.0003, ∆Cond ~0, ∆Sal ~+0.0001.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. No errors were found.

A header check was run. No problems were found. No off-scale fluorescence values were found.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.4db which is typical for the Vector. The shallowest data came from 1.8db.

The altimeter and bottom depth headers were exported from the SAMAVG and CLN files to spreadsheets. The bottom depths were checked against the log book. There were 4 discrepancies and 1 case with no log entry for depth.

As a further check the following calculation was made:

Check=maximum depth sampled (Max press*0.99) - water depth +altimeter reading

We don’t expect a 0 reading since the altimetry reading is an average over 2db, altimetry was often noisy and the bottom depth is subject to change during the cast. Changing the bottom depth to match the log entry improved the results for 3 of the 4 casts mentioned above, so for those 3 casts the header entry was adjusted in the SAM, SAMAVG and CLN files. In the other cases there is insufficient evidence to change entries.

Bottom depth was changed to match the log entry in the CLN files for casts 36, 45 and 52 and in the SAM and SAMAVG files for cast #52.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and as usual, there is a lot of variability with the up and downcast traces sometimes closer than temperature and sometimes further apart. This is likely due to varying vertical gradients. Overall the choice made earlier looks appropriate, so no further alignment will be applied.

Conductivity

Tests were run on 2 casts using a variety of shifts and the best results were found using the same settings as for 2016-05, namely -0.4 records for the primary and -0.5 records for the secondary.

SHIFT was run twice on all casts using those settings.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. The last time this sensor was used a setting of +70 records had the best results. Tests on a few casts suggest +70 is best for this cruise too though +90 also looked good for some features.

SHIFT was run on the pH:SBE channel using a setting of +70 records.

Plots were examined after this step and the results look ok.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There was 1 warning, but as it pertains to the upcast it is of no significance.

##### Other Comparisons

Previous experience with these sensors –

The only cruise that used the temperature, conductivity and dissolved oxygen sensors since the last factory calibration was 2016-05. The dissolved oxygen correction factor was 1.15 but there was a lot of scatter in the fit, and there was likely incomplete flushing for some casts. The 2 salinity channels were very close and lower than bottles by ~0.007 but because of delayed analysis and incomplete flushing it is believed that the CTD salinity is much more accurate than it appears from the comparison. The pressure sensor was used for 3 Ricker cruises in 2015 as well as 2016-05; no recalibration was found necessary during those cruises.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. As has been seen in all Strait of Georgia/ Juan de Fuca Strait cruises over the past year or two, temperature is frequently above the top of the range. Inside Baynes Sound this was limited to minor near-surface excursions. Outside Baynes Sound the top 20m were generally above the climatology maximum as were most temperatures between 40 and 75db. Below 80db all data were within the climatology. Salinity is mostly within the climatology range except for a few cases of low salinity in the top 10db and one case of low salinity below 55db at the deepest cast in Baynes Sound (event #45). The 3-standard deviation climatology is not recommended for near-shore work. While the results might suggest a systematic calibration error, similar results have been seen in this region during a number of cruises and using a variety of sensors. The excursions from the climatology are considered real.

Repeat Casts – There were no repeat casts and nearby casts are too shallow to provide useful comparisons.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

There is little to distinguish between the primary and secondary T and S channels. There were some spikes found in the conductivity channels and thus in salinity but not in temperature. Sometimes they appear in both channels but often in only one. Where they involve only a few records they were cleaned using interpolation; they were removed if more extensive. These spikes are mostly fairly small and do not look like the results of poor alignment. The secondary channels have slightly more spiking in the salinity that is not seen in temperature, so the primary channels were selected. There are many unstable T-S features near the surface which may be real or may be due to a fairly short wait at the surface or ship effects or surface debris that affect the plumbing. Where the unstable features are small and not obviously wrong they were left in place.

CTDEDIT was used to remove records with large spikes in both T and S, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. All files required some editing.

##### Initial Recalibration

Pressure does not appear to need recalibration.

Poor flushing of bottles and possible evaporation/adsorption of samples due to some delay in analysis make the comparison between bottles and CTD salinity unreliable. The history of the sensors was not helpful except to offer the hope that since they had been recalibrated recently and compared well with each other, calibration drift in these sensors may be slight.

Poor flushing also affected the comparison of dissolved oxygen samples with that from the CTD.

The recalibration used for cruise 2016-05 looks like the best choice for this cruise; the earlier cruise also had some flushing issues, but there were more data from a variety of conditions.

CALIBRATE was run using file 2016-45-recal1.ccf to correct the Oxygen:Dissolved:SBE channel in the SAM and MRGCLN2 files using:

CTD DO Corrected = CTD DO * 1.015

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly. The average of differences in the fit once outliers were removed was +0.015mL/L, but the standard deviation is 0.057mL/L. See file 2016-45-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed, the CTD DO was higher than the bottles by an average of ~0.07mL/L and standard deviation of 0.15mL/L; the results vary with depth with the CTD reading low right at the surface, high between 10m and 25m and close to bottles below 25m, on average. This pattern is likely due to incomplete flushing in the presence of significant DO gradients between 10 and 25m and frequent reversals of DO gradient near the surface and below 25m. Slow response time of the CTD sensor could partly account for this pattern as well. No further recalibration of DO is justified.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some small unstable features, but that is normal for this region in near-shore and well-mixed casts.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, pH and PAR data are nominal and unedited

except that some records were removed in editing temperature and salinity.

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available, but there are no such data for this cruise.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.7 mL/L from 0 to 25db

±0.25 mL/L from 25db to 200db

±0.06 mL/L below 200db

WARNING: The pH:SBE:Nominal data should be used with caution; no field

calibration data were available at the time of processing.

Calibration is required for each cast to get absolute values, although

general trends within a cast are likely real.

For details on the processing see the report: 2016-45-proc.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values ranged from about 105% to 145% with the lower values in Baynes Sound and higher values in the more open waters east of Denman Island. Given the wide range in DO gradients and surface saturation, these observations are not helpful in assessing the DO sensor calibration.

##### Final Bottle Files

CALIBRATE was run on the MRGCLN2 files and then the MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check and a header check were run on all files. A problem was found with 3 draw temperatures which had been entered incorrectly in the analysis spreadsheet. Those were corrected and the mass units derivation was rerun. Another run of Header Check turned up no further errors.

The track plot looks ok.

Data from the CHE files were exported to a spreadsheet and compared with rosette sheets and no errors were found.

Plots of each file were examined to ensure no problems had crept in.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Particulars (mostly notes from log)

TSG manifold failed. No TSG data were acquired.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 18Dec2014 | Factory |  |  |
| Conductivity | 1766 | 20Aug2014 | Factory |  |  |
| Secondary Temp. | 4054 | 20Aug2014 | Factory |  |  |
| Secondary Cond. | 3531 | 6Feb2013 | Factory |  |  |
| Transmissometer | 1396DR | 2Feb2016 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 19Dec2015 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| Surface PAR | 20518 | 21Mar2016 |  |  |  |
| pH | 0692 | 12Jan2016 | Factory |  |  |
| SeaPoint Fluor. | 3865 | n/a |  |  |  |
| Pressure Sensor | 0550 | 17Dec2014 | Factory |  |  |
| Altimeter | 1253 | n/a |  |  |  |