## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2016-46

Agency: Ocean Sciences Division

Location: Baynes Sound

Project: Baynes Sound Ecosystem Study

Party Chief: Sutherland T.

Platform: Vector

Date: June 22, 2016 – June 29, 2016

Processed by: Germaine Gatien

Date of Processing: 6 October 2016 – 18 October 2016

Number of original HEX files: 41

Number of CTD files: 38	Number of bottle files: 22

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#1119) on the secondary pump, a SeaPoint Fluorometer (#3865) on the primary pump with a 3X cable, a Biospherical QSP-400 PAR sensor (#4565), Surface PAR (#20518) a pH sensor (#0692) and an altimeter (#1252).

The data logging computer was Vector CTD Laptop.

The data acquisition program was Seasave v7.23.1.

The CTD deck unit was an SBE model 11+ s/n 508.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log and the Sampling/rosette log sheets were in good order.

The Niskin bottles are believed to have flushed poorly during most casts from this cruise. This is typical of rosette casts in protected areas especially when seas are calm. It is a particular problem when most sampling is from shallow casts where vertical gradients tend to be high. Cruise 2016-07 used the same equipment a few weeks earlier in the Strait of Georgia and Juan de Fuca Strait and included some casts in deeper and/or more exposed sites. So conclusions about sensor calibrations from that cruise were applied to 2016-46. CTD Dissolved Oxygen was recalibrated, but no recalibration was applied to salinity.

Salinity samples were analyzed within 13 to 18 days after collection.

There were 3 occasions when a near-surface bottle was fired but did not close. Attempts were made to replace the missing samples by making a quick return to the missing level after resetting the bottle. In theory, this could have worked, but the wait was too short for the CTD data to equilibrate and for the contents of the Niskin bottles to flush. The CTD DO data definitely looked out of line in profile.

While lowering the rosette to 5m is likely safe, it should be kept in mind that moving it lower than the depth at which other bottles were closed exposes those bottles to pressures that may cause leakage into the bottles.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.3 mL/L from 0db to 100db

±0.1 mL/L from 100db to 250db

±0.02 mL/L below 250db

The pH:SBE data were bad from casts 95 to 104, so the pH:SBE channel was removed from those files. This does not appear to be due to either a cap left on or insufficient hydration of the sensor. The sensor likely needs repair or replacement.

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration data were available at the time of processing. Calibration is required for each cast to get absolute values although

general trends within a cast are likely real.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX. There was confusion about what the cruise number was with the result that the raw data files have the wrong number.

The raw data file names were changed from 2016-47-**** to 2016-46-****.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Nutrients, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The sensor history was found for the pressure, conductivity and DO sensors. Since last calibrated at the factory this was the 7th use of the pressure sensor and the 4rd use for the conductivity and DO sensors.

The configuration file did not change through the cruise. The calibration constants were checked for all instruments and no problems were found. One file was saved as 2016-46-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2016-46-ctd.xmlcon to create CNV files.

A few casts were examined.  All expected channels are present.

- The primary and secondary temperature and conductivity channels are generally close during downcasts, but as usual the upcasts differ more due to noise in both channels. There are some spikes in both primary and secondary conductivity and temperature.

- Dissolved oxygen, fluorescence, PAR, pH and transmissivity profiles all look normal.

- The altimetry is very noisy near the bottom; while it is easy to estimate the bottom value by eye, the values that go into the headers will need to be checked carefully.

- PAR maximum values were usually lower than Surface PAR, but since most casts start when the PAR is 2m below the surface, that may be reasonable. During cast #101 when the CTD was kept running while out of water, the PAR readings are higher than the SPAR.

- The descent rate was kept as high as is reasonable for shallow sampling and quite high for the deepest cast. The descent speeds were mostly quite steady.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2016-46-ctd.xmlcon.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and no problems were found.

A preliminary header check was done and no problems were found. Fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

There were 3 cases where a bottle didn’t fire. In 2 cases a second file was started and the CTD was fired at the missing level. In 1 case the CTD came out of the water, the Niskin was reset and the CTD returned to fire at the missing level.

An initial (and unsuccessful) attempt to form 3 files with data at the right levels was done as follows:

- Casts #63/64 – The single bottle from cast #64 can be pasted into cast #64 replacing the Niskin #7 that didn’t close. First the bottle number in the SAM file for cast #64 was changed from 1 to 7 using Ultraedit. Then those data were copied and pasted into the Niskin #7 spot in file #63.

- Casts #95/96 – The single bottle from cast #96 can be pasted into cast #95 replacing the Niskin #9 that didn’t close. First the bottle number in the SAM file for cast #96 was changed from 1 to 9. Then those data were copied and pasted into the Niskin #9 spot in file #63.

- Cast #101 – The data from the second firing of Niskin #9 needs to be moved to the position of the first firing. First the bottle number for the 2nd firing was changed from 11 to 9 to simplify later steps. Then the data from the 2nd firing were copied, removed from their position at the end of the file and used to replace the first set.

The addsamp.csv file was edited to reflect the changes made and then was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine. Those files were bin averaged on bottle number.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2016-46-bot-hdr.txt which will be updated as needed during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2016-46oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2016-46oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2016-46SAL.xls. The analysis was done within 15-20 days of collection. The files were simplified and saved as 2016-46sal.csv.

That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2016-46nuts.xls.

Then the file was simplified, reordered on sample numbers and saved as 2016-46-nuts.csv.

The file was converted to individual NUT files.

The SAL, OXY and NUT files were merged with CST files in 3 steps.

After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

CLEAN was run on the MRG files to add 0 flags to empty flag channels and to update header limits.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There is a lot of scatter in the comparison with bottles. This is not surprising as all data are shallow and likely to be subject to poor flushing which will be fairly random as it is gradient-dependent. The 2 deepest samples were from 125db and 250db. They showed the primary to be low by 0.0216 and 0.0067 and the secondary low by 0.0226 and 0.0074.

Comparing the two sets of comparisons showed that, overall, the secondary salinity was lower than the primary salinity by ~0.001 (std dev 0.0005). This difference is similar to what was observed during 2016-07, so there is no evidence of the further drift beyond what was noted between 2016-45 and 2016-07.

2 out of 3 large outliers (difference>0.05 are associated with noisy CTD salinity data. The other one has a sample salinity that looks like it could be found ~3m below the CTD reading, and there was a very steady descent rate so incomplete flushing could very easily account for the difference.

For full details for the COMPARE run see file 2016-46-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

Using all casts and then excluding 2 major outliers and cases where residuals >0.03, the fit found was:

CTD DO Corrected = CTD DO * 1.0075 	(R2 = 0.38)

There were many shallow casts in protected areas where flushing is likely to be poor. When only 4 casts were used that were at least 150m deep, and excluding data by the same method, the fit was:

CTD DO Corrected = CTD DO * 1.0107 		(R2 = 0.61)

When only the deepest cast, #67, was used the result was:

CTD DO Corrected = CTD DO * 1.0149 	(R2 = 0.53)

This is reasonably close to the result of cruise 2016-07 when only the 4 casts closest to the mouth of Juan de Fuca Strait were used:

CTD DO Corrected = CTD DO * 1.0168	 (R2 = 0.8)

And when only the cast at the mouth of Juan de Fuca was used and a few outliers excluded using the same method, the slope was 1.0176 and R2 = 0.86.

During 2016-05 there were also problems that appeared to be due to poor flushing, so the fit used for recalibration was made using only 2 casts near the mouth of Juan de Fuca Strait. The fit was:

CTD DO Corrected = CTD DO * 1.0152	    (R2 = 0.74)

Using the 2016-07 correction based on 4 casts looks like the best estimate we can make.

There are 2 major outliers and they both come from files that were patched together, 63 and 101. The CTD data look bad for both.  The fit for the other patch was also poor. This problem will be studied next, but there is no need to redo the comparison as the questionable data were not included in the fit.

For more detail see document 2016-46-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further problems were found.

##### Study of casts 63/64, 95/96 and 101

The attempt described in section 4 to use data from a second firing of some Niskin bottles proved unsuccessful and poor results were obvious in the dissolved oxygen comparison described in the previous section.

The CTD data in the 2nd casts (64 and 96) do not match those in the original casts. Temperature values are a little off, but not too bad, but dissolved oxygen is low by ~2mL/L and salinity by from ~0.2 to 0.5psu. The differences were even greater for event #101 where there was a return to the water after resetting a bottle and without starting a new file. For all 3 cases the CTD data are not reliable.

One possible answer to this problem is to replace all CTD data except the pressure with pad values and enter the bottle results with flag 2 to draw attention to a comment. But especially for casts 95 and 101 when the CTD was lowered to 5m and the Niskin may contain water from higher than that as there was also no time for flushing. For casts 63/64 the repeated bottle was very shallow so the lack of Niskin flushing would not be so significant but it is hard to justify since the stop before firing was so short.

There is no reasonable case to make for merging the missing samples, so it seems best to pad the bottle values, with a flag 9. By using the flag we can let users know that bottle samples were taken on a second lowering but they are not reported due to poor CTD data during the 2nd firing and uncertainty about the depths from which the water in the Niskin bottles originated. The planned sample number will be used. This step was applied to the MRG files and CLEAN was rerun.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on ALIGNCTD to determine the best setting to advance the DO signal. When this sensor was used in April the best results were with a setting of +2.5s, but that appears to be too high for these data, especially for the many shallow casts. The deeper casts had stops for bottles which make the tests harder to interpret. A setting of +1.5s looks best overall. This setting can be fine-tuned later, if necessary.

ALIGNCTD was run on all casts using +1.5s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. Tests on previous cruises using these sensors showed the default setting of (α = 0.0245, β=9.5) did the best job and it does improve the data for both conductivity channels for these data.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on the only cast deeper than 200m to examine differences between sensor pairs. The shaded entries are from the last time these sensors were used.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2016-05-0060 | 350 | -0.0002 | +0.0001 N | +0.0013 | High, Steady |
| 2016-05-0085 | 350 | -0.0001 | +0.0001 N | +0.0012 | High, Steady |
| 2016-05-0101 | 330 | -0.0002 | +0.0001 N | +0.0013 | High, F.Steady |
| 2016-07-0069 | 350 | -0.0005 | -0.0001 | -0.0010 | High, F. Steady |
| 2016-07-0108 | 350 | -0.0002 | -0.0001 | -0.0011 | High, F. Steady |
| 2016-46-0067 | 320 | -0.0005 | -0.00015 | -0.0011 | High, Steady |

The differences are small and similar to those seen in 2016-07. There was a change in sign in the conductivity and salinity differences between 2016-05 and 2016-07, but there is no sign of further change in this cruise which followed 2016-07 immediately.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. No errors were found.

A header check was run. No problems were found. No off-scale fluorescence values were found.

The cruise track was plotted and added to the end of this report.

There are a few sources of information about the pressure accuracy:

- Header Check gave the average surface pressure as 2.2db which is in the normal range for the Vector.

- During event #101 of this cruise the CTD was left running with the pumps on while it was out of water to reset a Niskin bottle and then put back into the water. Tracking conductivity and transmissivity suggests that the surface was close to -0.2db and the pressure in air varied slightly but was close to -0.25db. The air pressure appears to be increasing during the time the CTD was out of water with a steady drift of ~0.025db in a 30s period. The CTD may require more time than this to establish the pressure in air or the drift may be due to real movement of the instrument. There were 2 readings of -0.7db during the time out of water but pressure was varying greatly at that time and the CTD had probably just left the water.

- During cast #7 of cruise 2016-07 using the same equipment, the CTD ran with pumps on near the surface while the CTD tech worked with a new winch operator on recovery of the rosette. The sensors appear to have been in water when the pressure was 0.3db and maybe right at the surface or out of water for 0.2db.

- At the end of event #19 during 2016-07 the CTD is clearly in water at +0.2db with pumps on. The pressure goes to slightly less than 0 for a few records and the primary conductivity spikes, so that was likely right at the surface. The CTD then returned to positive values and

- There was a deck pressure measurement of 0.7db during 2016-07. The result was doubted based on other observations. The pressure may not have had time to equilibrate before the reading was taken.

The mismatch between different types of observations and the 2 consecutive cruises illustrate that it is difficult, if not impossible, to establish surface pressure to better than ±0.2db and the initial accuracy is given as ±0.9db, though that includes measurements at depth so we can probably expect better results at the surface. No recalibration is justified.

The altimeter and bottom depth headers were exported from the SAMAVG and CLN files to spreadsheets. Few bottom depths were entered in the headers and some that were did not agree with the log entries (some had been altered in the log). Bottom depths were entered in the altimetry spreadsheet and a check for reasonableness was made by calculating

Check = Max Press*0.99 – Bottom depth + altimetry header

Those values were all within 2.1m except for cast #64 which was a shallow cast and the altimetry header is clearly wrong. It was removed from the header though it is not expected that the cast will be archived.

Spreadsheet 2016-45-header_merge.csv was prepared with bottom depths to be added to the headers; this was used with routine “Merge CSV file to Headers” to add the bottom depths to the CLN and MRGCLN2 files.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and as usual, there is a lot of variability with the up and downcast traces sometimes closer than temperature and sometimes further apart. This is likely due to varying vertical gradients. Overall the choice made earlier looks appropriate, so no further alignment will be applied.

Conductivity

For recent cruises using the same equipment the best results varied with -0.4/-0.5 records looking best for the primary/secondary conductivity during 2016-05 and 2016-45 while -0.5/-0.8 records was best for 2016-07. Tests were run on 4 casts and the best results were with -0.4 records for the primary and -0.5 for the secondary conductivity.

SHIFT was run twice on all casts using those settings.

pH

The pH sensor clearly needs alignment as it lags the temperature and the offset between downcast and upcast pH is much larger than that of temperature. The last 3 times this sensor was used a setting of +70 records had the best results and it looks ok for this cruise. As usual the best choice varies with descent rates and local gradients, but a shift of +70 records definitely improves all the data.

SHIFT was run on the pH:SBE channel using a setting of +70 records.

Plots were examined after this step and the shift results look ok, but there are some problems in the pH traces.

For casts 89, 92, 95, 97, 10, 101 and 104 there are sudden offsets in pH values that do not look natural, especially given the slow response of this type of sensor. For the first 2 of the affected casts, the problem was during the upcast and the downcast profiles look ok. For the others the obvious problems occurred during the downcasts and always below 50m, though the values above 50m look unusually low for the last 3 casts. These casts were not in Baynes Sound, so there is no reason to expect such variability.

##### DELETE

Cast #2 was dropped from the list of files to be processed since the pumps were not turned on. Cast #3 was run as a replacement but no data were archived.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were warnings for events #64 and #96 because the pressure range was <10db. These casts were run only to capture a surface rosette. There are other full casts at the same sites. Those casts will not be processed further.

##### Other Comparisons

Previous experience with these sensors –

The temperature, conductivity and dissolved oxygen sensors were used during 2 cruises in April 2016 and 1 in June. Cruise 2016-05 had delayed salinity bottle analysis and flushing of Niskin bottles was likely poor during both cruises. While the CTD salinity looked low compared to bottles, it is likely because the bottle salinity was too high. Dissolved oxygen comparisons also showed signs of poor flushing. Recalibration was based on results from an area where more ship motion likely improved flushing.  For cruise 2016-45 there were problems with flushing and delayed analysis. In June during 2016-07 both salinity channels were found to be within 0.002 of bottles. Overall, the primary salinity data were closer to bottles, but near the mouth of Juan de Fuca where flushing was likely better, the secondary salinity data were closer to bottles. Dissolved oxygen analysis was based on 4 casts near the mouth of Juan de Fuca. No recalibration was found necessary for pressure for any of these cruises; that sensor has been used 7 times since the last factory calibration.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. This climatology is not really suitable for shallow casts near shore. Many casts had temperatures above the climatology maximum near the bottom. Similar results were seen in this area in June during 2016-07, and high temperatures have been observed widely over the past 2 years with the depth of the excursions varying with time. Salinity data were all within the local climatology except for 4 casts to the south of Baynes Sound that had low salinity at the bottom; a similar result was found during 2016-07 in Baynes Sound. These excursions are not systematic and are not evidence of sensor calibration problems.

Repeat Casts – There was no repeat casts and while there were casts that were close together, they were too shallow for a comparison to be informative.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

There is little to distinguish between the primary and secondary T and S channels. Both have spikes and the comparison with bottles does not make it clear which salinity data are more accurate. The secondary data were selected for archiving for 2016-07, so the secondary temperature and salinity were selected for editing for this cruise as well.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes, mostly near the bottom of casts. All files required some editing, but mostly it was removal of records from near the top and bottom of casts and light cleaning of salinity.

##### Initial Recalibration

Pressure does not appear to need recalibration.

There is a lot of scatter in the comparison with bottles. This is not surprising as all data are shallow and likely to be subject to poor flushing which will be fairly random as it is gradient-dependent. The results from cruise 2016-07 suggested that both salinity channels were good to ±0.002, so no recalibration will be applied to those channels.

Poor flushing of Niskin bottles would also affect the comparison of CTD DO with titrated DO samples, but in this case bottle values will read lower than expected for the firing depth. The comparison based on all casts (excluding some outliers) had a slope of ~1.008, the 4 deepest casts had a slope of ~1.011 and for the single cast deeper than 300m the slope was ~1.015. The last result is reasonably close to the 2016-07 slope of 1.0168. As usually seen in areas where flushing is poor the slope of the fit was lower than might be expected based on the history of the sensor and especially low in well-protected shallow casts.

The results of the 2016-07 comparison of 4 casts in an area where flushing was expected to be good were applied to these data. CALIBRATE was run using file 2016-46-recal1.ccf to apply the correction:

CTD DO Corrected = CTD DO * 1.0168

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly. The average of differences in the fit once outliers were removed was +0.033mL/L and the standard deviation is 0.030mL/L. The differences are higher for high DO which generally occurs where gradients are higher, so this is what would be expected if flushing of bottles is incomplete. DO is low by ~0.015 at DO=2mL/L, very close to bottles at DO=3mL/L, high by ~0.025 at DO=4mL/L and high by 0.055 at DO=5mL/L. For the deepest cast the CTD DO is high by an average of 0.005mL/L. The overcorrection at low DO is likely due to the low or reversed vertical DO gradients near the bottom.

See file 2016-46-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed, the CTD DO was higher than the bottles by an average of ~0.08mL/L and standard deviation of 0.05mL/L; the results vary with depth with the CTD reading very close to bottles at the bottom. This pattern is likely due to incomplete flushing in the presence of significant DO gradients everywhere except at the bottom. Near the bottom the gradients are low so even if flushing is poor the error resulting would be small. No further recalibration of DO is justified.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have only one very small unstable feature but in this area some instability is expected though most disappear in averaging.

The only problem noted was in the pH signal for casts 95 to 104. As noted earlier these data look bad and should be removed.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

Channel pH:SBE was also removed for casts 95-104.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, PAR and PAR:Reference data are nominal and unedited

except that some records were removed in editing temperature and salinity.

The pH channel was removed from casts #95-104 due to sensor malfunction.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.3 mL/L from 0db to 100db

±0.1 mL/L from 100db to 250db

±0.02 mL/L below 250db

WARNING: The pH:SBE:Nominal data should be used with caution; no field calibration

data were available at the time of processing.

Calibration is required for each cast to get absolute values, although general

trends within a cast are likely real.

For details on the processing see document: 2016-46_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % surface saturation was calculated and plotted. Values ranged from 100% to 120%. The lowest values were close to shore to the south-east in Baynes Sound and the highest were outside Baynes Sound. These results are similar to what was found in this area during 2016-07. The results show no evidence of a DO sensor calibration problem but there are too many variables that affect DO surface saturation in this area to conclude that the calibration is good.

##### Final Bottle Files

CALIBRATE was run on all MRGMRH files and then the MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

Channel pH:SBE was also removed for casts 95 to 104.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Standards check and Cross Reference Listing routines were run and no errors were found.

Data from the CHE files were exported to a spreadsheet and compared with rosette sheets and an error was found in the nutrients for casts #101 and 104. This error arose because sample numbers were wrong in the nutrient spreadsheet. The two casts were not run in the order originally planned. The spreadsheet was edited because the analyst was unavailable – he was informed by email.

A header check was run on all files. The listing shows data from the SPAR sensor for all casts, but the sensor was not listed in the INSTRUMENT section for most files. This did not happen with the CTD files. No problem was found in the data. This must be due to some changes that were being made to IOS SHELL while processing was ongoing. The first set of files converted used the old version, but 3 files that were revisited later were converted using a new version that did pick up the SPAR sensor, likely due to an increased number of instruments allowed in the new version.  The files could be reconverted, but that would involve many steps; such reprocessing often leads to error. Since this inconsistency is of no great significance, the data were not reprocessed to pick up the missing information. One other error found was in the project name which was wrong in 3 files. That was corrected, Header Check rerun, and no further errors were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars (mostly notes from log)

1. Test cast; no data archived. Sound not working – depths from bridge.

2. First attempt no 10m soak and no pumps on, so rerun.

3. No data archived –see event # 40 for data from station BS15.

5. Rosette probably touched bottom. Raised 2m and waited 30s before firing Niskin #1.

8. Event # wrong in CTD files. Archived as 6. File names corrected before conversion.

29. Archived as event #28; should be #29. File names corrected before conversion.

40. Revisit to BS15 to replace event #3.

47. Archived as event #46; should be #47. File names corrected before conversion.

63. Niskin 7 did not fire.

64. 2nd cast to get surface sample from Niskin 7.

92. Sample #s out of order due to change in order of casts.

94. Brake was on slightly on way down so winch was a bit jerky.

96. Rosette returned to 5m to fire bottle 9 that did not close during event #95. Started new event number.

101. Rosette brought on deck to reset Niskin 9 and returned to 5m to refire without stopping archiving.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 18Dec2014 | Factory |  |  |
| Conductivity | 1766 | 20Aug2014 | Factory |  |  |
| Secondary Temp. | 4054 | 20Aug2014 | Factory |  |  |
| Secondary Cond. | 3531 | 6Feb2013 | Factory |  |  |
| Transmissometer | 1396DR | 2Feb2016 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 19Dec2015 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| Surface PAR | 20518 | 21Mar2016 |  |  |  |
| pH | 0692 | 12Jan2016 | Factory |  |  |
| SeaPoint Fluor. | 3865 | n/a |  |  |  |
| Pressure Sensor | 0550 | 17Dec2014 | Factory |  |  |
| Altimeter | 1252 | n/a |  |  |  |