## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 20 March 2025 | Updated channel names & formats in TOB files. G.G. |
| ### 25 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| ### 16 Aug 2020 | ### Added HPLC Data. S.H. |
| 13 Sept 2016 | Added warning to header comments that while CTD fluorescence is given in concentration units, it does not always compare well with CHL, especially for casts far from shore. |

## PROCESSING NOTES

Cruise: 2016-47

Agency: OSD

Location: WCVI

Project: La Perouse

Party Chief: Yelland D.

Platform: John P. Tully

Date: 24 May 2016 – 5 June 2016

Processed by: Germaine Gatien

Date of Processing: 4 August 2016 – 25 August 2016

Number of original HEX files: 102 	Number of CTD files: 102

Number of bottle files: 64		Number of bottle casts processed: 59

Number of original TSG files: 8		Number of processed TSG files:	 8

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), a SBE 43 DO sensor (#997) on the secondary pump, a SeaPoint Fluorometer (#3640) on the primary pump, Two Biospherical QSP-400 PAR sensors (#70613 and #4615), and an altimeter (#62354).

A thermosalinograph (Seacat 21 S/N 3411) was mounted with a Wet Labs WETstar fluorometer (S/N ws-3s 953p), remote temperature sensor and a flow meter.

Seasave version V7 22.4 was used for acquisition.

The data logging computer was the Tully CTD Laptop (Acer).

The deck unit was a Seabird model 11+, serial number 0425.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

The oxygen kit was a Scripps kit.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log, rosette log sheets and analysis logs were generally in good order with detailed comments on problems. There was one error on a rosette sheet: a sample number was used twice during one cast which led to labels not matching the rosette sheet and a repeated sample number used on the following cast. There are also cases of samples indicated on the sheets but no such samples in the analysis spreadsheets. If samples were skipped on purpose it would help to have a note about the change, or just a line drawn through the relevant box, to avoid time wasted looking for non-existent samples.

The secondary temperature and salinity were chosen for archiving for most casts since there was slightly less noise in the T-S plots. However, there were severe problems with the sensors mounted on the secondary pump during event 30, probably due to a blockage in the intake tube. The problems largely disappeared when the CTD neared the bottom, but there continued to be more noise in the secondary salinity traces than the primary for events 32 and 33, so primary temperature and salinity were selected for archiving for events 30-33. The Oxygen:Dissolved:SBE data from the downcast of event 30 were obviously bad. In the upcast of 30 and the full casts of 32 and 33 the oxygen data do not look obviously bad and the sensor data did not stand out in the comparison with bottles, so the Oxygen:Dissolved:SBE data were left in the 3 bottle files and in the CTD files for #32 and #33.

The comparison of salinity samples with the CTD salinity suggested that the primary salinity was low by about 0.0074 and the secondary by about 0.0033, but allowing for a small amount of evaporation or adsorption of bottle samples, a reasonable estimate is that they are low by ~0.006 and 0.002, respectively. There was an unusually large difference between the 2 temperature sensors, with the secondary lower than the primary by ~0.002Cº. It is not known which temperature sensor is more out of line, but correcting the primary temperature by subtracting the difference and recalculating salinity decreases the difference from bottles. Correcting the secondary temperature by adding the difference would decrease the secondary salinity and hence increase the difference from bottles. Having the primary salinity low by 0.006 seems less likely than having it low by 0.004, so the primary temperature was recalibrated. After the temperature correction was made a further recalibration was applied by adding 0.004 and 0.002 to the primary and secondary salinity channels. If a post-cruise calibration shows this choice to have been inappropriate, salinity can be recalibrated later.

As usual the CTD fluorescence values are higher than the extracted chlorophyll for CHL<1ug/L and then falling relative to CHL until they are about 50% of CHL at high CHL values.

There were many dissolved oxygen samples and some deep casts providing a good comparison with the DO sensor on the CTD. The results were close to those from some other cruises that used this sensor. The hysteresis correction was used in converting these data and there was no sign of hysteresis in the comparison.

Downcast Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 100db

±0.2 mL/L from 100db to 300db

±0.1 mL/L from 300db to 800db

±0.02 mL/L below 800db

The thermosalinograph worked throughout the cruise and a change to the set-up appears to have helped reduce bad salinity data. However, there remain problems that are likely due to bubbles. File #1 had sections where salinity gradually dropped, then suddenly recovered. There were often a series of such drops; some are seen during station stops but the worst occurred while the ship was steaming between stations LB16 and LC12. There may be smaller drops in other places that are masked by local variability. A graphical editor was used to clean single-point spikes in salinity in the first 4 casts, and a few records were removed for a time when the flow was turned off briefly in file #3. The larger problems in salinity were not amenable to editing so those data were left unchanged.

The comparison between TSG and CTD data suggests that the TSG may sometimes draw water from higher in the water column than 4m and likely from a depth for which we have no CTD data. While the comparison with loop samples also has a lot of variation, it does show a better correspondence, as expected given the two measurements are from water sampled at the same place and at almost the same time.

TSG Fluorescence is reported in volts. A comparison with the fluorometer mounted on the CTD and with loop chlorophyll samples provides a very rough estimate of TSG fluorescence as 7.5 times the fluorescence in volts. This should be considered a rough guide only as the ratio varies.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Nutrients, extracted chlorophyll, dissolved oxygen, NH4 and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked. All were used during 5 previous cruises.

There were 3 different versions of XMLCON files used during the cruise. The only difference was in the gain entered for the fluorometer. For events #1-10 the gain was entered as 30X but the cable was in fact a 10X. That was corrected in the con file for event #12 but the fluorescence went off-scale for that cast, so the cable was then changed to 3X and the con file adjusted accordingly. The 3X cable was used for the rest of the cruise. Configuration files were prepared for the two equipment configurations as follows:

- 2016-47-ctd1.xmlcon – for events #1-12

- 2016-47-ctd2.xmlcon – for events #13-195

The calibration constants were checked for all instruments and the only corrections needed were to the offsets for pressure and PAR sensor 70613. That PAR sensor is a new one and the factory calibration includes an offset that was not entered in the con file. A -0.6db adjustment to the pressure offset was determined to be necessary during cruise 2015-54 when there was sampling very close to the surface.

Tests were done first to see if the pressure offset from 2016-54 looks suitable for this cruise. There was a deck pressure measurement done during 1 cast with value 1.4db. If we subtract 0.6db, the deck value would be 0.8db which is reasonable. There was another cast during which the rosette accidentally came out of the water very briefly with pumps turned on. The minimum CTD pressure for that cast was 1.1db. If the rosette was out of water, the pressure sensor must be fairly close to the surface, so subtracting 0.6db for a reading of 0.5db looks reasonable. The conductivity, transmissivity and fluorescence suggest that the CTD was in water though if it left very briefly it might not be obvious as some water would remain in the system. In any case, a minimum pressure of 0.5db looks more likely than 1.1db. Neither of these tests are conclusive, but taken with the evidence from 2016-54 it is wise to use an offset of -0.5476 which is the factory setting from the last servicing (0.05243db) minus 0.6db.

The new pressure offset was entered in the 2 configuration files.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2016-47-ctd.xmlcon to create CNV files. The hysteresis and Tau functions were selected. There are some deep casts.

One deep cast was also converted using the original con file used at sea to ensure that the PAR offset entered was appropriate and it was. With the offset the deep values from the new PAR sensor had a minimum value of 1.0e-12.

Cast #184 has some serious problems in both conductivity channels, mostly in the upcast.

A few casts were examined and all expected channels are present.

The temperature channels are track reasonably well during downcasts though there are some spikes in the primary temperature. During upcasts there are larger differences than usual with the primary temperature looking spiky early in the cruise. Later casts have some spikes but may be better. Temperature differences look large even in deep water.

Both conductivity channels have some spikes but the primary is worst especially during upcasts. Casts #30 and 184 look very odd.

The altimeter appears to have worked well. Fluorescence and transmissivity look ok, though there are some small spikes at depth and fluorescence did go off-scale for at least one cast (#12).

The 2 PAR channels are reasonably close and the traces look normal.

Dissolved oxygen looks normal.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2016-47-ctd1.xmlcon (events #1-12) and 2016-47-ctd2.xmlcon (events #13-195).

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Bottle files were created in testing Niskin bottles, but there was no sampling, so the files will not be processed further for events #1, 71, 73, 116 & 119

Temperature and salinity were plotted for all BOT files to check for outliers. For casts 60, 63, 64, 77, 80, 86, 113 and 131 a few spiky points were removed from both salinity channels using CTDEDIT.

The output files were copied to *.bot.

A preliminary header check and no problems were found. CTD fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

There was some confusion in sample numbers:

- Casts #69 and 74 both used sample #263. There was no OSD sampling from event #69 so there is likely no chance that there will be any confusion so no change was made.

- Cast #110 – There was an error in sample numbers that originated from a mistake in the sampling log – sample #391 was entered for both Niskins #5 and #6 for event #110, but the labels were printed correctly. Samples from Niskin 6 through 20 are incorrect on the sheet. They should be #392 to 406 but are given as #391 to 405. But sample #406 was also used for the first bottle fired for event #113 so it is not a simple matter of using the label sample numbers. The simplest fix was to use 392-405 & 9406 for event 110 and leave event 113 as it is. The sample numbers will all be correct except that we need to rename the first instance of #406 to #9405. This will apply to DO, SAL, CHL, NH4 and NUTS samples.

Sort was used to ensure that the ADDSAMP file was in sample number order.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

The SAM files were bin-averaged on bottle # and called SAMAVG.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2016-47-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2016-47chl*.xls. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2016-47chl.csv. The event numbers in the file were incorrect – they were repaired in the simplified file and the analyst was informed of the errors. There are also inconsistencies in the CHL labels for event #2 - the depth recorded is not in agreement with the sample # and Niskin #. (This was investigated after the files were merged.) The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2016-47oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. Filtered samples that were run for comparison were removed.  The spreadsheet page with the final data was simplified and was then saved as 2016-47oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2016-47SAL.xls. The analysis was done between 31 and 41 days of collection. The files were simplified and saved as 2016-47sal.csv. File 2016-47sal.csv was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2016-47_nutrients*.xlsx. This includes a precision study. The file was simplified, reordered on sample numbers and saved as 2016-47-nuts.csv. The file was converted to individual NUT files.

NH4

NH4 data were obtained in file 2016-47_NH4.xls which included a report on precision. The file was simplified (removing the filtered samples) and saved as 2016-47NH4.csv and converted to individual NH4 files.

The SAL, CHL, OXY, NUT and NH4 files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. There were a number of cases where sampling was indicated on the sampling log sheets but no sample value was found in the merged files. Checks with the spreadsheets from the analysts show that no samples were missed in the merge process, so it looks like these samples were never taken.

The only significant problem was the one noted above with the CHL sampling for event #2. The log sheet indicates that samples were to be taken from Niskin bottles 8, 9 and 11 at 20, 10 and 0m. But the sample labels indicate that they came from Niskins 7, 8 and 10. The depths on the labels were given as 20, 10 and 0m, but those Niskin bottles were actually fired at 30, 20 and 5m. The sample numbers were added after the labels were printed. The fluorescence maximum was at about 25m on the way down and about 20m on the way up, so there may have been a conscious decision to sample a little deeper than originally planned. An initial comparison of fluorescence and CHL values does suggest that the samples came from 30, 20 and 5m. The analyst confirms that Niskins 7, 8 and 10 are the right choice.

There were no sample numbers assigned to 1 bottle each from casts #38 and 69, so those lines were removed from the MRG file and CLEAN was rerun on those 2 files.

There were many cases of sampling indicated on rosette sheets but no such samples found by the analysts. It is assumed that the plan was changed; a note of changes on the rosette sheet would help. An example is cast #82 for which there were no DO samples analyzed although 7 were shown on the log sheet.

##### Compare

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

When outliers were excluded based on residuals the fit found was:

CTD DO Corrected = CTD DO * 1.0208 + 0.0377

Other cruises that used this sensor since it was last recalibrated produced the following fits:

CTD DO Corrected = CTD DO * 1.0151 + 0.0426	(2015-10)

CTD DO Corrected = CTD DO * 1.0225 + 0.0306	(2015-21)

CTD DO Corrected = CTD DO * 1.0208 + 0.0376	(2015-54 no DO<2mL/L)

CTD DO Corrected = CTD DO * 1.0066 +0.0652	(2016-41 free offset, no DO<2mL/L)

Or CTD DO Corrected = CTD DO * 1.0202		(2016-41 offset forced to 0)

The results from this cruise are remarkably close to 2015-54 and similar to 2015-21.

There were no significant outliers.

A plot was made that showed the points below 1100m in green with others in red and this showed no signs of hysteresis.

For more details see document 2016-47-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Note: The DO samples from cast #131 were accidentally excluded from this comparison - this was a shallow cast with a very high vertical DO gradient at most depths. A test run including the cast showed that most bottles would have been rejected as outliers. The fit with the cast included had a slope of 1.0211 and offset of 0.0371, which increases the correction by no more than 0.001mL/L. These fits are not so precise as to make this difference significant. (2016-47-dox-comp1x.xlsx.)

Salinity

Compare was run with pressure as reference channel.

There were 3 outliers that were not readily explained by high standard deviations in the CTD salinity or being very close to the surface where high salinity gradients together with incomplete flushing would explain large differences.

- Sample 241, event 67 – The bottle value is about 0.01 higher than expected and higher than any salinity seen anywhere during this cast even allowing for recalibration of salinity. A flag 3 was added by the analyst.

- Sample 378, event 105 – This looks likely to be due to incomplete flushing as salinity that high was seen during the bottle stop. No flag is needed as this value likely reflects the contents of the Niskin bottle.

- Sample 473, event #190 – This was a bottom bottle and the descent rate of the CTD was very quiet for this cast, so incomplete flushing would lead to a lower bottle value than expected. No flag was added as the value likely represents the contents of the Niskin bottle.

When these outliers and all values above 450m were excluded the primary salinity was found to be low by 0.0074 and the secondary low by 0.0033. The standard deviations were 0.0024 and 0.0025.

The samples were analyzed within 6 weeks, so evaporation should not be large though adsorption could account for bottle salinity being high by ~0.0015. Flushing is expected to be quite efficient as the descent rate was fairly noisy for most casts, though there may be a small effect. So it is probably reasonable to estimate that the primary salinity is low by ~0.006 and the secondary by ~0.002. The difference between these two estimates is similar to the differences reported in section 9.

There is quite a large difference between the temperature channels, with the secondary reading lower than primary by ~0.002 which would account for a difference of ~0.002 between the two salinity channels. It is not known if the primary temperature is reading too high, the secondary too low or a combination of both. However, if the secondary is corrected by adding 0.002, the secondary salinity would be even further from bottles. If the primary is corrected by subtracting 0.002 the primary salinity would move closer to bottles. So it looks more likely that the primary temperature is the principal source of the difference. Correcting primary temperature and recalculating primary salinity would make it look low by about 0.004 allowing for errors in the bottle salinity ~0.0015.

The analysis precision was good at Sp=0.001 but there were only 5 samples and 1 was rejected as a Chauvenet outlier.

Plots of differences against time and against salinity do not suggest any trend that cannot be explained by different depths and local gradients.

For full details for the COMPARE run see file 2016-47-sal-comp1.xls.

After the analyst added a flag and comment to sample 241, the merge process was rerun for event #67.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor. For cast #12 there was only 1 sample and the fluorometer was reading off-scale at the relevant depth, so it was removed from the comparison.

The plot of FL/CHL vs CHL below shows the usual pattern with CTD fluorescence reading higher than the extracted CHL for CHL<1ug/L and then falling relative to CHL until it is about 50% of CHL at high CHL values.

A fit forced through the origin has a slope of about 0.6.

For full details of the comparison see file 2016-47-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on a few casts and results with a setting of +2 to +5s; +3s was best overall and was the value used for 2016-41.

ALIGNCTD was run on all casts using +3s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. During other recent cruises, the default setting of (α = 0.0245, β=9.5) looked better for both conductivity channels than others tested. One cast was checked for this cruise and the default setting does improve the data.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. The shaded values are from a previous cruise that used these sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2015-21-0034 | 1000 | -0.0013 | -0.00014 | -0.0002 | High, Mod |
|  | 1900 | -0.0014 | -0.00006 | +0.0008 | Med, Noisy |
| 2015-21-0068 | 1000 | -0.0013 | -0.00008 | +0.0002 | High, VNoisy |
|  | 1900 | -0.0012 | -0.00004 | +0.0009 | “ |
| 2015-21-0099 | 1000 | -0.0014 | -0.00008 | +0.0004 | High, XNoisy |
|  | 1900 | -0.0013 | -0.00002 | +0.0012 | “ |
| 2016-47-0039 | 1000 | -0.0019 | +0.0001 | +0.0032 | High, V Noisy |
|  | 1900 | -0.0018 | +0.0002 | +0.0040 | “ |
| 2016-47-0069 | 1000 | -0.0019 | +0.0001 | +0.0036 | High, V Noisy |
|  | 1900 | -0.0017 | +0.0001 | +0.0039 | “ |
| 2016-47-0094 | 1000 | -0.0015 | +0.0002 | +0.0039 | High, Mod |
|  | 1900 | -0.0017 | +0.0002 | +0.0045 | “ |
| 2016-47-0105 | 1000 | -0.0017 | +0.00015 | +0.0038 | High, Noisy |
|  | 1900 | -0.0018 | +0.0002 | +0.0045 | “ |

The temperature differences were unusually noisy. The temperature differences for both cruises reported are higher than we expect. The differences have increased since 2015-21. There were 3 cruises between these two that are not reported on because they had only shallow casts.

Conductivity differences are also a little higher, but are within expected levels.

Salinity differences appear to be increasing with time though the differences are so noisy that such a conclusion is weak.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

A header check was run separately for each fluorescence cable to ensure that there were no off-scale fluorescence data. There were problems for some casts with the 10X cable, so for casts #1 to 12 CLEAN was also set to replace fluorescence values >14.82ug/L with pad values. There were no off-scale values with the 3X cable.

The first 2219 records were removed from the CLN file for event #184 so that DELETE will not select data from the initial drop to 30m.

##### Checking Headers

There is a note in the log that the computer clock was running 3 minutes fast for the first 8 casts, after which it was corrected. The log times are entered at the beginning of acquisition and normally would agree the “system upload times” in the headers, but for the first few casts they differ by 3 minutes. The time in the file headers is the NMEA time from when the computer is turned on, so no change is needed to the files.

A cross-reference list was checked against the log book. No errors were found.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.9db which is reasonable for the Tully. There was a deck measurement of pressure at 1.4db, which after the change to the pressure offset used in processing would be 0.8db.

The altimeter and water depth readings from the headers of the CLN files were exported to a spreadsheet. A quick calculation (Pressure*0.99 -Depth +Altimeter Reading) was made; where it differs from 0 by more than 4, the cause was investigated. The water depths were checked against log entries and discrepancies were found for a few casts. One was a case where the log was clearly wrong but for 2 others the log entry looks better, so that value was entered in the header. For the other casts identified as having a problem, plots were made and all the header entries for altimetry were found to be good. There were notes in the log indicating that for many of these casts there were doubts about the sounder readings. So an estimate of depth was made by adding Pressure*0.99 to the altimetry entry in the header. Those values were then used to replace the header depths that are clearly wrong. Such corrections were made to the CLN files for events #33, 35, 38, 45, 69, 71, 73, 88, 91, 94, 100, 102, 162, 171 and 184.

Similar data were exported from the SAMAVG files. Checks were made to ensure that for casts with only near-surface sampling there were no faulty altimetry headers due to misinterpreted spikes. There were such entries for 3 casts (events 18, 20, and 176) which were removed from the headers. Another cast was missing an expected altimetry header but there was only noise from the altimeter at the time of bottle firing. The depth corrections made for the CLN files were also made to the SAMAVG files. The MERGE with MRGCLN1s files was repeated and CLEAN rerun on the bottle files.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 3 casts to see what shift in conductivity does the best job of removing noise in the salinity channels The best primary shift was -0.4, while -0.6 worked best for the secondary.

SHIFT was run on all casts using -0.4 records for the primary conductivity.

SHIFT was run on all casts using -0.6 records for the secondary conductivity.

Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

The sensors were used during 4 cruises in 2015 and 1 just before this cruise in 2016. The pressure offset was updated based on near-surface sampling on the previous cruise. Salinity calibrations were mostly not trusted due to flushing problems, delayed analysis and/or noisy comparisons. The secondary salinity has been closer to bottles than the primary for all cruises. The dissolved oxygen comparisons have been reasonably consistent.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Salinity values were mostly within the climatology except for a few low surface values and some low values at depth in Rivers Inlet and some slightly high values around 50m at some casts just offshore of the 100m isobath. Temperatures were frequently above the maxima in the top 20m and in deeper waters in the northern Strait of Georgia. A month earlier similar deep features were observed in the southern Strait of Georgia using different sensors. None of these observations suggest a calibration problem.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

An initial examination to decide which channels to archive shows little difference in noise in the two salinity channels. The secondary channels were selected for the 3 cruises run just before this one and the secondary salinity is closer to bottle values, so the secondary channels were selected for this cruise for most casts. However, for event #30 the secondary salinity looked very bad in the top 100m and poor below that. For the 2 casts that followed the secondary salinity was noisy in places where there were sharp temperature changes, so the primary T and S were selected for events 30, 32 and 33. For cast #34 the secondary salinity appeared to have recovered and looks better than the primary.

Since the dissolved oxygen sensor was mounted on the secondary pump, a check of the bottle comparison was made to see if excluding casts #30 and 32 would affect the fit found, and the difference in fit was found to be insignificant, affecting the corrections by ~0.0001mL/L at the low end of the DO range and ~0.0005mL/L at the high end.  The only samples that looked significantly out of line were already excluded in the original fit. There was no salinity sampling for the affected casts. The fluorescence is relatively high for casts #29 and 30 and transmissivity quite low, so the likelihood is that biological material got into the intake tube for the secondary pump. Conductivity and temperature were both affected though the conductivity looks worst. The primary is also slightly affected.

The primary T and S channels were selected for casts #30, 32, 33 and editing was fairly heavy for those casts as there were many unstable features.

The secondary T and S channels were selected for all other casts.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. Some bad salinity points were removed. All files required some editing.

After editing T-S plots were examined for all casts. There were a few small unstable features in the T-S plots that look reasonable for the areas in which they were found. But for cast #30, the data look unbelievable at mid depths with a large unstable feature. The secondary temperature, dissolved oxygen and secondary salinity all look bad. The large excursion resolves by about 100m of the downcast, but there remain small spikes. The upcast secondary T and S channels look better but they are noisier than the primary.

Casts #32 and 33 also have noisier secondary T and S than the primary, though the problem is just small spikes rather than the large excursion seen in cast #30.

The problem would appear to be due to something getting into the secondary intake tube early in cast #30 and gradually flushing out.

So editing was rerun on casts #30, 32 and 33 using primary temperature and salinity.

##### Recalibration

There is no evidence to suggest that the pressure channel needs further adjustment.

Based on the comparisons with bottles described in section 5, 2016-47-recal1.ccf was prepared to subtract 0.002 from the primary temperature, to recalculate primary salinity and to apply the following correction to channel Oxygen:Dissolved:SBE:

CTD DO Corrected = CTD DO * 1.0208 + 0.0377

Then a second run of CALIBRATE was made using file 2016-47-recal2.ccf to add 0.004 to the primary salinity and 0.002 to the secondary salinity.

These corrections were first applied to the SAM and MRGCLN2 files. COMPARE was rerun and the results confirm that the recalibrations were applied properly. When outliers were removed based on residuals (~4% of samples were excluded) the average of differences in the DO fit was +0.0003mL/L and the standard deviation was 0.021mL/L. The two salinity channels are still lower than bottles, on average, but by no more than can be explained by adsorption and evaporation of samples.

(See file 2016-47-DO-comp2.xlsx for details.)

CALIBRATE was then run on the EDT files using 2 runs.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When all data were included the CTD DO was higher than the titrated samples by an average of ~0.05mL/L with a standard deviation of 0.09mL/L. There are two known sources of error that do not involve calibration drift: slow response of the DO sensor in high gradients and poor flushing of Niskin bottles. Both will lead to the downcast DO from the sensor reading higher than the bottles above the DO minimum (found ~800-1200m). Examination of cases with DO<1.5 mL/L shows that for those above the DO minimum the sensor reads slightly high (average +0.016mL/L). Below 1200m it reads very close to the bottles, with an average difference of 0.0005mL/L. So those data are consistent with the 2 error sources discussed above, though the data from below the minimum may suggest the DO values are slightly high. Near the surface the DO sensor is close to bottles, but between about 10m and 500m where vertical DO gradients are highest, the DO:SBE values are mostly higher than bottles.

No further recalibration is justified. See 2016-47-dox-comp3.xlsx for details.

##### Fluorescence Processing and special files for Dr. Peña

The COR2 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The newer PAR sensor has readings that are up to 10% higher than the older sensor except at very low values when it is sometimes lower than the older model.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts except #30, 32 and 33 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

The PAR1 channel was removed from casts for which it was not mounted.

REMOVE was run on cast #30, 32 and 33 to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

For cast #30, channel Oxygen:Dissolved:SBE was also removed.  For casts #32 and #33 that channel looks ok, though the quality may be lower than usual since the secondary temperature and salinity look noisy.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and unedited except that

some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

Downcast Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 100db

±0.2 mL/L from 100db to 300db

±0.1 mL/L from 300db to 800db

±0.02 mL/L below 800db

There were 2 Biospherical/Licor PAR sensors. The one identified as PAR refers to

a new sensor (S/N 70613) and PAR1 refers to an older sensor (S/N 4615).

For details on the processing see document: 2016-47_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

A cross-reference list was produced.

The sensor history was updated.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values ranged from 80% to 150% with most between 100% and 110%. Saturation was between 103% and 106% for casts #30 to 46 which were furthest from shore where such values look reasonable. The lowest values were in Juan de Fuca Strait and Rivers Inlet and Johnstone Strait where strong vertical mixing is likely. The values suggest that the DO calibration is reasonably good.

##### Final Bottle Files

The MRGCOR2 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts except #30, 32 and 33 to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate, Voltage:6 and Voltage:7  and Flag.

The PAR1 channel was removed from casts for which it was not mounted.

REMOVE was run on cast #30, 32 and 33 to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

Note that the Oxygen:Dissolved:SBE was not removed since the upcast data looked reasonably good as judged by examining the full profile and looking at the comparison with bottles.

A second SBE DO channel was added for both the CTD DO and titrated DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2016-47-bottles-final.xlsx. The entries were compared with the rosette log sheets to ensure no samples had been missed. An error was found in file 131. The dissolved oxygen samples had been identified as event 129 and should have been 131. (This error had been discovered earlier and fixed in the file, but there was an error in the rosette log). The merge process was repeated for that cast and the data were exported again and no further errors were found.

Standards check and a header check were run on all files. No problems were found.

The track plot looks ok.

Plots of each file were examined to ensure no further problems were found.

A cross-reference listing was produced for the CHE files.

##### Thermosalinograph Data

Major changes were made in the set-up of the TSG for this cruise.

There were 8 hex files. There were 10 extracted chlorophyll and salinity samples taken.

a.) Checking calibrations

The configuration files were all identical. One file was renamed as 2016-47-tsg.xmlcon. No errors were found in the calibration parameters.

b.) Conversion of Files

The 8 files were converted to CNV files using configuration file 2016-47-tsg.con.

Those CNV files were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

A time-series plot showed that the data look good. There are small one-sided spikes in the salinity but not many. There are sections of noisy salinity The flow rate was steady in sections but the values vary from 0.8 to 1.3, with a few spikes to low values including one brief drop out on May 31st at about 2300 hours.

The track plot looks fine. The plot was added to the end of this report.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db and exported to a spreadsheet which was saved as 2016-47-ctd-tsg-comp.xls. There were 91 casts which overlapped with TSG files. There was no overlap with the 8th TSG file.

The TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, lab temperature, salinity and fluorescence and the files were reduced to the times of CTD files and loop samples.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude and longitude were all ≤0.0008° and the median differences were 0.0000° and 0.0002°. This shows both the times and positions are reliable for both systems.

d.) Comparison of T, S and Fl from Loop & Rosette Samples and TSG and CTD data

- T1 vs T2 The intake temperature sensor worked throughout the cruise. The median difference was 0.189C° during stops for CTDs. Three TSG files were examined and the median differences were 0.18C°, 0.20C°, 0.19C° and 0.22C°. There was no obvious relationship to the flow rate with the largest heating associated with the highest flow rate. There was the usual dependence on intake temperature with warming in the loop reducing as the intake temperature increased and approached the ambient temperature of the ship.

- Flow Rate The flow rate was steady in sections but there were shifts between 0.8 to 1.3.

- TSG vs CTD The spreadsheet comparing CTD and TSG files was then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. Well-mixed casts were identified by finding the depth at which the salinity had increased by 0.005 from the 4m reading. For 13 of the casts this did not occur until the CTD was below 10m.

|  |  |  | Lat Diff | Long Diff | Temp (TSG Intake-CTD) | Temp (TSG lab-CTD) | Salinity (TSG-CTD) | FL Ratio (TSG / CTD) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| all casts |  | median | 0.0000 | 0.0002 | 0.0053 | 0.1992 | -0.0822 | 0.159 |
|  |  | stdev | 0.0002 | 0.0002 | 0.4260 | 0.5034 | 1.3503 | 0.064 |
|  |  | max | 0.0004 | 0.0008 | 1.7765 | 2.9883 | 4.3309 | 0.322 |
|  |  | min | -0.0005 | -0.0004 | -0.8583 | -0.1880 | -5.4508 | 0.047 |
| 13 well-mixed casts |  | median | 0.0000 | 0.0001 | 0.0008 | 0.1872 | -0.4863 | 0.214 |
|  |  | std dev | 0.0002 | 0.0002 | 0.0094 | 0.0218 | 0.6103 | 0.045 |

1. Intake Temperature The intake temperature is higher than the CTD temperature by a median of 0.007C° using all casts but the range was large (standard deviation 0.40C°). The average difference was very small for the 13 casts that were well-mixed below 4m, though the standard deviation was 0.009C°. Looking at only casts #58 to #82 the temperature was high by a median of 0.009C° but the standard deviation was 0.4.

2. SALINITY TSG salinity data are lower than the CTD salinity by a median value of 0.087 but the standard deviation was very high, at 1.35. There were a few areas where the differences were large and variable. When only casts from #58 to #82 are included the salinity is low by a median value of 0.06. For the 13 well-mixed casts the salinity is lower than the CTD salinity by a median of 0.5 but the standard deviation is 0.6. Some of those well-mixed casts occur when the TSG salinity traces suggest the presence of bubbles which may explain why salinity looks so low.

3. FLUORESCENCE

The TSG fluorescence data are uncalibrated and expressed in volts. The ratio of the TSG fluorescence to that from the CTD has a median value of 0.16 (standard deviation 0.07). Based on this comparison a very rough estimate of TSG fluorescence in chlorophyll units can be made by multiplying by 7.8.

A few casts were examined in detail to see if the TSG values match CTD values from a different level:

- For cast #16, the TSG values look like they come from shallower waters than measured by the CTD, with temperatures higher and salinity lower than TSG data at 3 or 4m.

- For cast #53 the TSG temperature was very close to the 4m CTD data but the salinity was about 0.5 lower than any measured by the CTD. This could be due to bubbles affecting the TSG salinity, but there is no spiking in the time-series plot, and careful examination of the TSG record for 10 minutes before and after the reading, the values are all low compared to the CTD. Perhaps the TSG was drawing water from above 3m though we would expect higher temperature in that case.

- For cast #55 the T and S gradients are low between 3 and 4 and the TSG temperature reads higher than that of the CTD by about 0.015C° and the TSG salinity reads lower than the CTD by about 0.6. This looks like the TSG was drawing water from above 3m.

- For cast #195 the TSG temperature and salinity values are similar to CTD values at 3m rather than 4m.

(See 2016-47-ctd-tsg-loop-comp.xls.)

- Loop Bottle - TSG Comparisons

All loops were taken while the ship was underway.

The TSG salinity is lower than the salinity sample from the loop by a median value of 0.017 but when 5 out of 10 samples are excluded due to high standard deviation in the TSG salinity the TSG is found to be low by a median value of 0.009, but the standard deviation is 0.008.

For the 10 loop samples the ratio of the TSG fluorescence (median over 2 minutes) to the extracted CHL from the loop is 0.15. A fit of Loop CHL versus TSG fluorescence is shown below. When the trend-line is forced through the origin the loop CHL is found to be higher than the TSG in volts by a factor of ~7.4 which is close to the results of the comparison with the CTD fluorometer.

(See 2016-47-ctd-TSG-loop-comp.xlsx.)

- Calibration History

The temperature and conductivity sensors were recalibrated in November 2015 and were used during 2016-01, but the salinity was spiky from that cruise so comparisons with loops and CTD were not trusted. There were also variations in flow rates that complicated comparisons.

The TSG fluorometer values were about 24% of the CTD fluorometer. Chlorophyll values were low for that cruise. The fluorescence data were archived in volts.

Conclusions

1. The TSG clock worked well.

2. The TSG flow rate was steady in sections but it stepped from one value to another at times, with values ranging from 0.8 to 1.3

3. The temperature in the loop increases by roughly 0.2Cº.

4. The TSG intake temperature was higher than the CTD temperature by a median of ~0.005Cº using all casts. It was higher by <0.001 for the 13 well-mixed casts, but the standard deviation was high at 0.009Cº.

5. The TSG Salinity is lower than the CTD salinity by between 0.08 and 0.5 depending on which data are selected. Compared to loop samples it is low by 0.02 if all samples are included and by 0.009 if half the cases are excluded due to high standard deviation in the TSG salinity. The standard deviation is 0.008 for the latter case.

6. The fit of CTD fluorescence in ug/L against TSG fluorescence in volts indicates that the TSG fluorescence in ug/L.:

TSG FL(ug/L) = 8.23 * TSG FL (volts) -0.29

Or a rough estimate can be made by multiplying by 7.4 based on comparison with loop CHL samples or 7.8 based on comparison with CTD fluorescence.

7. The changes to the TSG set-up appear to have led to better salinity data. There appear to be some drops in salinity that may be due to bubbles, but the problem is not as severe as seen in the past.

8. The comparison between TSG and CTD data suggests that the TSG may sometimes draw water from higher in the water column than 4m and likely from a depth for which we have no CTD data. While the comparison with loop samples also shows a lot of variation, it does show a better correspondence, as expected given the two measurements are from water sampled at the same place and at almost the same time.

9. No recalibration will be applied as the comparisons are subject to too many sources of error. TSG fluorescence will be left in voltage units, but a multiplier of 7.5 will can be applied for a very rough estimate in chlorophyll units, though that value is likely too high when chlorophyll is <0.5ug/L.

f.) Editing

The ATC files were copied to *.EDT.

Each file was opened in CTDEDIT and the salinity channels were cleaned where there were single-point spikes not seen in the temperature trace.

In the course of editing it became clear that there are areas where salinity gradually falls and then recovers. This is assumed to be due to bubbles accumulating and then clearing. These are impossible to edit except by removing large sections of salinity data. They were left in place as it is far from clear which data are bad.

File #1 – Salinity cleaned. Large patches of salinity that look suspicious were left alone.

File #2 – Just a few salinity spikes cleaned.

File #3 – Salinity cleaned and some records removed because flow was off briefly.

File #4 – Just a few salinity spikes cleaned.

File #5, 6, 7 and 8 – No editing applied.

.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Difference, Conductivity:Primary, Flag and Position:New channels.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats.

Those files were saved as TOB files.

The Standards Check and Header Check were run and no problems were found.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

# Particulars

PAR 4615 ON: 1-33, 46-67, 74-88, 97, 111-113, 119- 133, 142-195.

Out of Order firing: 2, 26, 34, 42, 43, 66 & 82.

Bottle file not needed: 1, 71, 73, 116 & 119.

1. Test cast – all bottles fired, no sampling.

1-8. Computer time 3 minutes ahead of UTC.

9. Computer time corrected before this cast.

10. Top caps on Niskins #5 and 7 rotated since leaking – did not leak after that.

12. Fluorescence gain corrected in configuration file – should have been 10X for casts 1 to 10, was entered as 30X. Fluorescence saturated >15ug/L so cable changed to 3X for next cast.

13. Fluorescence gain cable 3X and con file correct.

28. Pressure on deck before 1.4db.

37. Niskin #17 replaced with spare using some parts of original.

38. Test of Niskin 17 – no leaks. 30s stop at 457m to correct wire angle.

43. Used bottles 2, 17-23 but all bottles fired to get PAR readings for incubation experiment.

46. PAR 4615 on. Fluor trace weird upcast 250-150db.

63. Stall at surface on way down for ship position correction.

69. PAR off.

71. Niskin 7 closed just to test – no sampling. Cast started 5min before rosette went in the water b/o winch issues.

73. Niskin #7 closed for test and #1 by accident, neither needed. No leak on #7.

94. Flushed CTD with DI H2O following cast.

110. Brief stop on way down ~100m to lift boom and ~400m to lower boom.

116. Change O-ring Niskin 8 before cast and fired to test – no bottle file needed. About 8.5 min between start of file and start of acquisition. Stops at 1940 and 1650 downcast due to wire angle.

119. Test bottle fired, no bottle file needed.

126. Station name needs changing in CTD header.

142. Stopped 282 down to straighten wire angle, started coming back up to 200m, then back down ( resumed in wrong direction) stopped at bottom to straighten wire angle.

144. Stopped at 16m down to correct aft wire lead and at 142m.

162. No wait at surface after 10m soak.

184. Stop at ~30m on way down, came back up because operator went in wrong direction, rosette out of water briefly, back in and down. Salinity difference ~0.25 on way up, difference improved towards surface. Rosette and CTD had red stringy stuff all over it. CTD flushed several times with DI H22O, CTD and rosette wiped.

190. Down 10 but came up w pump on and out of water briefly. Returned to water quickly so left pump on and continued 10 and up. Sat at surface until temp and sal difference stabilized then started archiving, so brought back to surface, started archiving and continued.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 17Dec2014 | Factory |  |  |
| Conductivity | 2280 | 19Dec2014 | Factory |  |  |
| Secondary Temp. | 2663 | 15Jan2015 | Factory |  |  |
| Secondary Cond. | 2754 | 19Dec2014 | Factory |  |  |
| Transmissometer | 1185DR | 27Apr016 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 17Jan2015 | Factory |  |  |
| PAR | 70613 | 21Mar2016 | Factory |  |  |
| PAR | 4615 | 16Mar2011 | IOS |  |  |
| SeaPoint Fluor. | 3640 | n/a |  |  |  |
| Pressure Sensor | 0443 | 17Dec2014 | Factory |  |  |
| Altimeter | 62354 | n/a | Factory |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3411       Cruise ID#:	2016-47

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3411 | 7Nov15 | Factory |  |  |
| Conductivity | 3411 | 7Nov15 | Factory |  |  |
| WetLabs Fluorometer | Ws3s-953p |  |  |  |  |
| Temperature:Secondary | ? | 0842 |  |  |  |
| Flow meter | ? | n/a |  |  |  |