## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2016-49

Agency: PBS, Marine Ecosystems and Aquaculture Division, Nanaimo, B.C.

Party Chief: Wyeth M.				Location: Johnstone Strait

Project: Hard Bottom Longline Hook Survey	Platform: Neocaligus

Date: 1 August 2016 – 23 August 2016

Processed by: Germaine Gatien

Date of Processing: 13 July 2017 –14 July 2017

Number of original HEX files:	70		Number of CTD files: 71

# INSTRUMENT SUMMARY

SeaBird Model SBE-19+ CTD (s/n 7580) was mounted with oxygen sensor #0249.

# SUMMARY OF QUALITY AND CONCERNS

There was a digital version of the logbook available and a spreadsheet with position and time information and some comments. The file names were non-standard, but included information that made it easy to change them to standard. The event numbers run from 1001 to 1071.

Most hex files had been broken into individual cast files; one file had to be separated into 2 casts.

The CTD was soaked between 1db and 1.5db for about 2 minutes. A soak of 3 to 5 minutes is recommended by the manufacturer but that is to allow the instrument’s temperature to equilibrate. For this area and season a 2-minute soak should be fine.

There was no calibration sampling for either salinity or dissolved oxygen. The calibrations were from February 2016, so calibration drift was likely small.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea. The file names were non-standard.

##### Preliminary Steps

A digital version of the log book and a spreadsheet with positions and comments for casts were provided.

The event numbers go from 1001 to 1071.

The cruise summary sheet was completed.

A single configuration file used for all casts.

There were no parameters on file at IOS for the temperature, conductivity, dissolved oxygen and pressure calibrations, so the values could not be checked.

The configuration file was saved as file 2016-49-ctd.xmlcon.

##### Conversion of Raw Data

All casts were converted using file 2016-49-ctd.xmlcon.

Casts 1047 and 1048 were joined in a single file. The first 1346 records were saved as file 1047 and records 6000 to the end were saved as file 1048.

The file names were edited to change them to standard format.

The soak time at the surface was about 2 minutes.

##### FILTER

A low-pass filter with a 0.5s time constant was applied to the conductivity to force it to have the same response as the temperature. The pressure is smooth and does not require a filter.

##### ALIGNCTD

Alignctd was run to advance the DO signal by +3 seconds which brought the DO voltage traces into better alignment with the temperature traces.  Temperature was advanced by 0.5s relative to pressure.

##### DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

Plots of temperature, salinity and dissolved oxygen versus pressure were also examined to see that data look reasonable and they do.

CLEAN was run to replace pad values in the pressure channel with interpolated values.

##### Checking Headers

Some header information was missing: latitude, longitude, station name and water depth. A spreadsheet was provided with that information. After some reformatting, the spreadsheet was used to prepare file 2016-49_hdr_merge.csv.

IOS SHELL program “MERGE:CSV file to headers” was used to add information from the csv file to the IOS files with output MRH including the start time for each cast.

A cross-reference listing was produced and looks fine.

A preliminary header check was also run and no errors were found. The minimum pressure found in the data was -0.05db; that is fine given the initial accuracy quoted for this type of pressure gauge is ±1db.

The surface check is not very useful for this type of sampling as there may be records that are in water but with the pump not on. The average surface pressure was +0.1db, but the associated salinity values were near zero. However, the salinity begins to take on “in water” values at about +0.2db, so the pressure looks ok.

A track plot was produced and added to the end of the report.

##### SHIFT

Conductivity

Small shifts to conductivity were tested to see if they improved the salinity, as judged by T-S plots. None of the settings appeared to improve the data, so no shift was applied.

Oxygen

The dissolved oxygen channel was aligned earlier in processing, but was checked again. There are few distinct features to allow a good test, but there are no signs that further adjustment is needed.

##### DELETE, CLIP and CTDEDIT

DELETE was run on all casts using the following parameters:

Remove Surface Records in the top 10m.

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

Plots were made to see what editing was needed and it is clear that some bad near-surface data remain.

CLIP was run to remove data with pressure <1.5db which come from the soak period or the very beginning of the cast, but even with that step some poor data remain near the surface. There are a few unstable features at depth and a few casts appear to have very noisy profiles. However, such casts generally have a very small range due to strong vertical mixing, so the instabilities are actually very small and most of those are likely real.

Casts that may need editing were examined in CTDEDIT and some editing was applied to 30 casts. This involved removing a few surface records or records corrupted by shed wakes near the bottom and/or light cleaning of salinity where unstable features appear to be due to minor misalignments between temperature and salinity.

All CLIP files were copied to EDT and then the CTDEDIT output files were copied to EDT to produce a complete set of files.

##### Inter-comparisons

Sensor History – The sensors have been recalibrated since they were last used. They were all new in 2015.

Comparison of repeat casts –There were no repeat casts and nearby casts are too shallow to expect good repetition.

Historic Ranges – There is local climatology for only the last few casts in the southern part of the survey. Plots of temperature and salinity with the local climatology superimposed show all data fell within the normal range except for a few high temperature values near the bottom of one cast. The 3-standard-deviation climatology is too severe for near-shore data, so a slight excursion is not considered indicative of a calibration problem.

Post-cruise calibrations – None available.

##### CALIBRATE

There was no calibration sampling and no useful history to enable estimates of calibration drift.

##### Bin Average and REMOVE

The files were bin-averaged using 1db bins.

T-S plots were examined. Some unstable features remain but are likely to be real.

CHANGE UNITS was run to derive dissolved oxygen in mass units.

REMOVE was run to remove Scan_Number, Conductivity:Primary, Oxygen:Voltage, Descent Rate and Flag channels from all casts.

##### Dissolve Oxygen Saturation

Plots were made to check the DO saturation rate at the surface. Values ranged between 55% and 135% with the most common values being about 70% and 110%.  This range is not surprising given that the survey is in an area that is subject to strong tidal mixing where low saturation values are expected and some other areas with higher vertical stratification so that higher DO saturation is likely. We can conclude little about the calibration of the DO sensor except to say that the range looks reasonable.

##### HEADER EDIT and final checks.

Header Edit was used to fix headers, fix formats and channel names and to add the following note to the headers:

Data Processing Notes:

----------------------

The Oxygen:Dissolved:SBE channel is nominal and unedited except that some

records were removed in editing temperature and salinity.

There was no salinity or dissolved oxygen calibration sampling and no other

information available upon which to base a recalibration. However, the

sensors had been recalibrated recently, so calibration drift is expected

to be small.

For details on the processing see the report: 2016-49-proc.doc.

A cross-reference listing was produced.

A header check was run on the CTD files and no errors were found.

The sensor history was updated.

Particulars:

1003. Steep with tide running, so CTD went down on an angle. Don’t know how close cast got to the bottom.

1042. CTD hit bottom.

## CRUISE SUMMARY

| Cruise ID#: 2016-49 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Dates: Start: 1 August 2016 End: 23 August 2016 |  |  |  |  |
| Location: Strait of Georgia |  |  |  |  |
| Party Chief: Wyeth M. |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? |
| 1 | SEABIRD | 19+ | 7580 | No |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE19+/7580	Cruise ID#:		2016-49

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 7580 | 23Feb2016 | Factory |  |  |
| Conductivity | 7580 | 23Feb2016 | Factory |  |  |
| SBE43 Oxygen | 0249 | 11Feb2016 | Factory |  |  |
| Pressure | 7580 | 19Feb2016 | Factory |  |  |