## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2016-58

Agency: PBS, Salmon and Freshwater Ecosystems Division, Nanaimo, B.C.

Party Chief: Neville C.				Location: Qualicum River

Project: Big Qualicum Purse Seine		Platform: F/V Ocean Venture

Date: 26 May 2016 – 15 July 2016

Processed by: Germaine Gatien

Date of Processing: 27 September 2016 – 28 September 2016

Number of original HEX files:	13		Number of CTD files:  13

# INSTRUMENT SUMMARY

Two CTDs were used for this cruise:

- SeaBird Model SBE-25 CTD (s/n 0334) was used with temperature sensor #2968, conductivity sensor #2173, Wetlabs ECO Fluorometer #2215, oxygen sensor #766 and pressure sensor 0464.

- SeaBird Model SBE-25 CTD (s/n 0404) was used with temperature sensor #2095, conductivity sensor #1764, Wetlabs ECO Fluorometer #4185, pH sensor #0852, oxygen sensor #1176 and pressure sensor 0482.

# SUMMARY OF QUALITY AND CONCERNS

The log book had equipment lists and there was a good record of which CTD was in use for each cast as well as problems encountered. The times were in UTC. A spreadsheet was provided with positions, station names, water depths and CTD serial numbers; this was very helpful since position information is not in the headers.

The CTD was soaked at about 1 to 1.3db, generally for about 1.5 minutes. A soak of 3 to 5 minutes is recommended by the manufacturer but that is to allow the instrument’s temperature to equilibrate. If the temperature difference between air and water is not large, a 1.5 minute soak may be sufficient.

A pH sensor was mounted on CTD #0404 but the cap was left on so no useful data were acquired.

CTD #0404 had poorer pressure resolution than expected from a SeaBird 25 so the CTD technician was informed that there may be a problem with the settings for the instrument.

CTD #0334 performed as expected.

There was no calibration sampling.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea.

2. Preliminary Steps

This cruise was one of 3 that used the same sensors in spring/summer 2016, with CTD #0334 used during 2016-31, 2016-32 and 2016-58 and CTD #0404 used for parts of 2016-31 and 2016-58. All attached sensors were the same for each CTD. The steps followed for this cruise are based on analysis of the earlier cruises with checks made to ensure they were appropriate for this one.

The Daily Log was obtained. There was a pH sensor on CTD #0404 but the log indicates that the sensor cap was left on throughout. There was no pH sensor on CTD #0334.

The cruise summary sheet was completed.

A single configuration was file used for each CTD. The calibration parameters were checked and no errors were found.

There is a question about what offsets to use for the pressure sensors. The pressure for CTD 0404 is too crude to make an offset estimate, but there is no obvious evidence of a problem.

A study of pressure for CTD #0334 was done when 2016-32 was processed. This was based on data from all 3 cruises that used sensor #0334 in early 2016: 2016-31, 2016-32 and 2016-58. It appeared that pressure might be reading a little low. There is difficulty in judging where in the record the surface is encountered: there is slow equilibration of pressure at the beginning so it may appear to rise while there is actually no motion, and at the end water will remain in the plumbing for a little while after the CTD rises out of water. Despite the issue of pressure equilibration at the beginning of casts, records from the end of the soak period can generally be identified. The intent was to soak with the CTD fully submerged but close to the surface. Conditions were calm so shallow soaking was practical. We might expect a pressure of 0.7db to 1.5db.

- During cruise 2016-31 there were surface soaks of 1 to 1.5minutes at fairly consistent pressures ranging from 0.1db to 0.8db, with an average of 0.5db.

- For 2016-32 the soak pressures varied from 0.3db to 5.2db and the length of the soak period was sometimes too short to trust the pressure. There are enough low readings to suggest that either the top of the CTD was out of water or the pressure is reading low by ~0.5db, but the evidence is weak.

- Similarly, for cruise 2016-58 the soak pressures ranged from 0.4 to 0.8db and the readings at the end of casts suggest that the CTD left the water at about -0.4 to -0.6db.

Adding +0.5db looks like a reasonable step and would produce surface soaks at between 0.5db and 1.3db for 2016-31 and 2016-58. This may be an under-correction but it should ensure that no good data are lost near the surface because of negative pressures. We can only trust pressure from this type of sensor to within ±1db.

Configuration file 2016-58-ctd0334.xmlcon was changed by entering 0.5db for the pressure offset.

## 3. Conversion of Raw Data

All casts were converted using the appropriate configuration files.

Plots show that the channels all have reasonable profiles except for the pH data from CTD #0404. Notes in the log explain that the pH sensor cap had not been removed for any of those casts, so that channel will be removed later.

The times in the log are in UTC.

As was found during cruise 2016-32 the times in the headers of the casts from CTD#334 differ from UTC by 10 hours and 10 minutes.

As for 2016-31 and 2016-32 the times in the headers CTD #0443 are in PDT but a day ahead, so actually 17 hours ahead.

4. WILDEDIT

This step was skipped as it does not usually work well on SBE25. There were no obvious spikes to remove from pressure, temperature or conductivity.

5. FILTER

Tests were run on a few casts to choose a filter for these data.

SBE25 0334– There is a lot of fine-scale noise in the pressure, temperature and conductivity channels. But the pressure resolution looks normal with steps of ~0.15db. WFILTER was run with a filter width of 5 and this worked well without over-smoothing. When salinity is derived it looks much better with this filter. A test was made using a low-pass filter, but it was no better than the window filter which is the usual choice for this CTD type.

SBE25 0404 -The pressure looks steppy for all widths tested using WFILTER but the filter worked well on the temperature and conductivity and the resulting salinity looks good. The pressure readings from this sensor are odd with the steps being ~1.3db, much grosser resolution than expected from an SBE25. This same problem was noted in 2015.

While the WFilter is the usual choice for the SBE25s, a test was run using a low-pass filter with a 0.5s time constant for pressure and a 1s time constant for conductivity. The results were smoother for pressure but worse for salinity and temperature, with many unstable features in T-S space. The WFILTER approach produced better salinity.

So WFILTER was run on all casts using a width of 5 for pressure, temperature and conductivity.

6. ALIGNCTD

For SBE25 #0334 a setting of +2.5seconds looks best at bringing temperature and dissolved oxygen traces into alignment though there was little variation between 2s and 3s. The +2.5s setting was applied to all files.

For SBE25 #0404 the DO traces have few distinctive features to enable a good test of the best setting, but +2.5s does improve the alignment.

ALIGNCTD was run on all files using +2.5s.

7. CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25. CELLTM was run on all casts using that setting.

8. DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers. This was done separately for each CTD to get the correct information in the headers.

CLEAN was run to replace pad values in the pressure channel with interpolated values. There are negative values in the fluorescence and dissolved oxygen channels, so during the first run of CLEAN all negative fluorescence values were replaced with pad values.

CLEAN was run a second time to replace the negative dissolved oxygen values with pad values.

Time was adjusted using program ADD TIME CHANNEL as follows:

- For CTD #0334 the times were adjusted by subtracting 10 hours 10 minutes (-10.17hrs). This is the same adjustment used for this CTD during cruise 2016-32.

- For CTD #0404 the times were adjusted by subtracting 17 hours.

10. Checking Headers

Some header information was missing: latitude, longitude, station name and water depth.

A spreadsheet was provided and that was used to prepare file 2016-58-header_mrg.csv. IOS SHELL program “MERGE:CSV file to headers” was used to add information from the csv file to the IOS files with output MRH.

A track plot produced all positions on land! An error had been made in calculating the longitude for the CSV file used for merging. When that was fixed the plot looked fine.

A cross-reference listing was produced and checked against the log entries. There are some minor discrepancies in time that are not significant. However, there are two cases where the dates differ:

- Event #22 is given in the log as occurring at 01:18 on May 28th, which looks correct, but the spreadsheet shows it as occurring on May 27th which does not make sense since the activity at the previous site occurred late on May 27th. The spreadsheet date is wrong but the log is fine.

- Event #49 is given in the log as occurring at 00:35 on July 15, but the previous activity was a SET cast at 23:53 on July 15th so event #49 must have been on July 16th as is found in the file headers. So the log is wrong and the spreadsheet is wrong.

A header check was also run and no problems were found.

The surface check was run:

- For the SBE25 # 0404, the average surface value was 1.2db. The associated salinity values show the CTD was in water although the first few records often had very low salinity as the pump would not have been on. At the end of casts pressure falls to ~0.2db and salinity drops quickly between 0.4db and 0.2db; fluorescence at 0.2db looks like the CTD was in water.  Since the pressure steps are large for this particular sensor the surface check data are not very useful, but the pressure certainly look likely to be good to ±1db, which is the specification.

- For the SBE25 #0334 the average surface pressure was 0.4db with associated salinity values being very low but the pumps would not be on at that point. Most casts had a 1.5m soak around 1 to 1.3m. One soak lasted more than 2 minutes and from that we can see that the upcast data at 1m has similar values to the soak period. The salinity looks about the same until pressure is 0.04db and the pumps appear to have been on when the file ended. For another cast the fluorescence went off-scale at -0.05db which suggests that the sensor was right at the surface. While it is possible that the pressure is reading very slightly low, it is likely good to ±0.2db, which is well within the specifications.

##### 11. SHIFT

pH

- CTD #0334 did not have a pH sensor.

- The cap was left on the pH sensor on all CTD 0404 casts, so those data will not be archived.

No shift was applied to pH data.

Conductivity

For the SBE25 #0334 a shift of +0.5 records improved the data. Salinity was recalculated.

For the SBE25 #0404 every shift tested makes the data look worse. No adjustment is needed.

Fluorescence

The fluorometers on the SBE25s were not pumped, so a shift in alignment is expected to be small or unnecessary. Profile plots of temperature and fluorescence were examined but the steps in the fluorescence are hard to relate to variations in temperature which was filtered so the steps are gone. No alignment adjustment will be made.

Oxygen

The dissolved oxygen channel was aligned earlier in processing, but was checked again. As usual, the results vary from feature to feature as local vertical gradients and descent rates affect the alignment. No further shift will be applied.

##### 12. DELETE, CLIP and CTDEDIT

DELETE was run on all casts using the following parameters:

Remove Surface Records in the top 10m.

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: The only warning concerned the last record so is of no concern.

CLIP was run to remove data with pressure <1.5db.

All files were examined in CTDEDIT and some editing was applied to 8 casts. This mostly involved removing a few near-surface records or a few records from the bottom of the cast.  There was also some light cleaning of salinity.

14. Inter-comparisons

Sensor History – There is no useful history for these sensors since they have not been used on any other recent cruises with calibration sampling.

Comparison of repeat casts –There were casts that were close in space but not time. They are not useful for judging the repeatability of the CTD.

Historic Ranges – The local climatology is a 3-standard deviation climatology covering 3 months and a fairly large region, so it is considered a severe criterion for casts near shore. Nonetheless, profile plots were examined with the climatology superimposed.

- Temperatures were above the climatology maximum for all the May casts. However, the May temperature values are all within the summer climatology (June 1 – August 30), so these excursions could well be because spring 2016 was warm. A similar result was found in May 2016 in Cowichan Bay from cruise 2016-31. For the July casts the temperatures near the surface were within the climatology but below 15m most temperatures were above the maximum. When this sensor was used during 2016-31 temperatures were a little high in May but within the climatology in mid-June and August. During 2016-32 temperature data fell within the climatology. Excursions from climatology in this area are likely real and not due to sensor calibration problems.

- The only salinity values that fell outside the climatology were a few points at the bottom of casts #33 and 36 which were at or below the minimum.

Post-cruise calibrations – None available.

15. CALIBRATE

There was no calibration sampling and no useful history to enable estimates. Pressure appears to be ok with the offset applied in conversion. No recalibration was applied.

16. Fluorescence Filter

This step will be skipped as the fluorescence is not very spiky.

17. Bin Average and REMOVE

The files were bin averaged using 1db bins.

T-S plots were examined and no further editing was found necessary.

REMOVE was run to remove Scan_Number, Conductivity:Primary, Oxygen:Voltage, Descent Rate and Flag channels from all casts.

For casts 2-22 channel pH:SBE was also removed.

Dissolved oxygen% surface saturation was calculated and plotted. Near-surface values ranged from 80% to 90% during the May leg while it was about 115% for the July casts. The surface waters were better mixed and fluorescence was much higher in May than in July. The 2 legs had different sensors. The sensor used in July is the one found to have saturation between 90% and 110% in May in Cowichan Bay. The sensor used in May had a wide range of surface saturation rates in Cowichan Bay in late May and August, but they were almost all higher than those observed during this cruise. There is insufficient information to comment on DO sensor accuracy.

18. HEADER EDIT and final checks.

Header Edit was used to fix headers, fix formats and channel names and to add the following note to the headers:

For CTD #0334

Data Processing Notes:

Fluorescence and Oxygen:Dissolved:SBE channels are nominal and unedited except

that some records were removed in editing temperature and salinity.

There was no salinity or dissolved oxygen calibration sampling and no other

information available upon which to base a recalibration.

For details on the processing see the report: 2016-58-proc.doc.

For CTD #0443

Data Processing Notes:

----------------------

Fluorescence and Oxygen:Dissolved:SBE channels are nominal and unedited except

that some records were removed in editing temperature and salinity.

There was no salinity or dissolved oxygen calibration sampling and no other

information available upon which to base a recalibration.

The quality of data from CTD #0404 is lower than usual for this cruise due to

low resolution in the pressure.

Channel pH:SBE was removed because the sensor cap was accidentally left on.

For details on the processing see the report: 2016-58-proc.doc.

---------------------------------------------------------------------------------

A cross-reference listing was produced.

A header check was run on the CTD files and no errors were found.

The sensor history was updated.

Profile and T-S plots were examined and no problems were found.

Particulars

2-22. pH buffer bottle not removed during cast.

19. Change of CTD to #0334

## CRUISE SUMMARY

| Cruise ID#: 2016-58 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Dates: Start: 4 May 2016 End: 21 August 2016 |  |  |  |  |
| Location: Discovery Islands |  |  |  |  |
| Party Chief: Neville C. |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? |
| 1 | SEABIRD | 25 | 0334 | No |
| 2 | SEABIRD | 25 | 0404 | No |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0334	Cruise ID#:		2016-58

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 23Dec2014 | Factory |  |  |
| Conductivity | 2173 | 19Dec2014 | Factory |  |  |
| ECO Fluorometer | 2215 | May 2015 | ? |  |  |
| SBE43 Oxygen | 0766 | 23Dec2014 | Factory |  |  |
| Pressure | 464 | 23Dec2014 | Factory |  |  |

Make/Model/Serial#:	SEABIRD/SBE25/0404	Cruise ID#:		2016-58

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2095 | 1Apr2011 | Factory |  |  |
| Conductivity | 1764 | 3Jan2014 | Factory |  |  |
| pH | 0852 | 18Jan2016 |  |  |  |
| ECO Fluorometer | 4185 | 2Nov 2015 | ? |  |  |
| SBE43 Oxygen | 1176 | 19Dec2015 | Factory |  |  |
| Pressure | 482 | 4Jan2016 | Factory |  |  |