## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 20 March 2025 | Updated channel names & formats in TOB files. G.G. |
| ### 25 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| 16 Aug 2020 | Added HPLC Data. S.H. |

## PROCESSING NOTES

Cruise: 2016-62

Agency: OSD

Location: WCVI

Project: La Perouse

Party Chief: Yelland D.

Platform: John P. Tully

Date: 31 August 2016 – 12 September 2016

Processed by: Germaine Gatien

Date of Processing: December 9, 2016 – February 7, 2017

Number of original HEX files: 113 	Number of CTD files: 113

Number of bottle files: 	55		Number of bottle casts processed: 55

Number of original TSG files: 4		Number of processed TSG files:	 4

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1185DR), a SBE 43 DO sensor (#997) on the secondary pump, a SeaPoint Fluorometer (#3640) on the primary pump, a Biospherical QSP-400 PAR sensors (#70613), and an altimeter (#62354).

A thermosalinograph (Seacat 21 S/N 3411) was mounted with a Wet Labs WETstar fluorometer (S/N ws-3s 953p), remote temperature sensor and a flow meter.

Seasave version 7.26.1.8 was used for acquisition.

The data logging computer was the Tully CTD Laptop.

The deck unit was a Seabird model 11+, serial number 0425.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial #58879.

The oxygen kit was a Scripps kit.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log, rosette log sheets and analysis logs were generally in good order with detailed comments on problems.

The secondary temperature and salinity were chosen for archiving since there was a lot of noise in the primary data and the secondary have been chosen for other recent cruises that used the same equipment.

The comparison of salinity samples with CTD salinity was marked by a lot of scatter and the precision study showed poorer results than usual. There was a delay of 7 to 8 weeks before samples were analyzed which has been found to lead to increased salinity values due to adsorption and evaporation of samples, the latter factor being random depending on how well the bottles were sealed. Since the salinity bottle inserts in use for some recent cruises have been found to be of lower quality than usual, seals may well have been poor. Another factor was that the salinometer used was a Portasal which is ±0.001 PSU less accurate than the 8400B Autosal which is usually used. There may also be some error due to incomplete flushing of Niskin bottles which also leads to increased salinity, though this is not expected to be a large error for open waters. The secondary CTD  salinity was found to be lower than bottles by about 0.009, but based on the history of the sensor and a rough estimate of errors, it is likely much more accurate than that. No recalibration was applied based on the history of the sensor and low confidence in the bottle comparison. If a post-cruise calibration shows this choice to have been inappropriate, salinity can be recalibrated later.

As usual the CTD fluorescence values are higher than the extracted chlorophyll for CHL<1ug/L and then falling relative to CHL until they are about 50% of CHL at high CHL values.

There were many dissolved oxygen samples including some deep casts, so a good comparison with the DO sensor on the CTD was possible; results were in line with those from other cruises using this sensor.

Downcast Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5  mL/L from 0 to 200db

±0.25 mL/L from 200db to 400db

±0.08 mL/L from 400db to 800db

±0.02 mL/L below 800db

The intake thermistor malfunctioned throughout the cruise and the data are not usable. Otherwise the thermosalinograph worked well throughout the cruise with few salinity spikes. A correction was made to the TSG salinity based on comparison with CTD salinity and previous experience with the sensor. While there is no evidence in the salinity traces of the presence of large bubbles, small ones are likely present, which would at least partly explain why the TSG salinity is lower than that of the CTD. This should improve the overall accuracy but may overcorrect for times when the ship is stopped or where there are no bubbles, and under-correct where bubbles are more significant. The TSG fluorescence is given in voltage units; no estimate in concentration units is provided because there were no loop samples. Based on comparisons with CTD data and the history of the sensor, a proxy for intake temperature was created by subtracting 0.18Cº from the lab temperature.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Nutrients, extracted chlorophyll, dissolved oxygen, NH4 and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors on the CTD were checked. All were used during 10 previous cruises since they were last calibrated. The TSG sensors have been used for 4 other cruises since last used.

The same configuration was used throughout the cruise. The calibration constants were checked for all instruments and no changes were required.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2016-62-ctd.xmlcon to create CNV files. The hysteresis and Tau functions were selected. There are some deep casts.

A few casts were examined and all expected channels are present. The temperature channels usually track reasonably well during downcasts but for some casts they are further apart than expected with the secondary traces looking much smoother than the primary. There were comments in the log about sudden shifts in dissolved oxygen. Most often the traces look normal, but there are cases where gradients seem higher than expected or upcast and downcast traces differ somewhat, though it is not obvious that the data are bad.

Altimetry, PAR, fluorescence and transmissivity traces look normal.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2016-62-ctd.xmlcon

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. For casts 25, 67, 123 and 133 a few spiky points were removed from channel Salinity:T1:C1 using CTDEDIT.

The output files were copied to *.bot.

A preliminary header check and no problems were found. CTD fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

Sort was used to ensure that the ADDSAMP file was in sample number order.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

The SAM files were bin-averaged on bottle # and called SAMAVG.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2016-62-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2016-62chl*.xls. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2016-62chl.csv. The event numbers in the file were incorrect – they were repaired in the simplified file and the analyst was informed of the errors. There are also inconsistencies in the CHL labels for event #2 - the depth recorded is not in agreement with the sample # and Niskin #. (This was investigated after the files were merged.) The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2016-62oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and was then saved as 2016-62oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2016-62SAL.xls. The analysis was done between 48 and 56 days of collection and there is a precision study. The files were simplified and saved as 2016-62sal.csv. File 2016-62sal.csv was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2016-62_nutrients*.xlsx. This includes a precision study. The file was simplified, reordered on sample numbers and saved as 2016-62-nuts.csv. The file was converted to individual NUT files.

NH4

NH4 data were obtained in file 2016-62_NH4.xls which included a report on precision. The file was simplified (removing the filtered samples) and saved as 2016-62NH4.csv and converted to individual NH4 files.

The SAL, CHL, OXY, NUT and NH4 files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. The only problems found were:

- Nutrient data for event #104 had been entered as from event 38; this was a reversal of event number and station name.

- CHL sample 149 had flag 1 but the rosette log shows that the duplicate was never drawn because there was not enough water. Since there is a sample, the flag was changed to 2.

There are a number of casts with no samples analyzed at IOS. CHE files will be prepared anyway as they may be required later.

##### Compare

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

When outliers were excluded based on residuals the fit found was:

CTD DO Corrected = CTD DO * 1.0248 + 0.0233

When only deep casts were included (maximum pressure >950db) the fit was:

CTD DO Corrected = CTD DO * 1.0254 + 0.0303

The last time this sensor was used was during 2016-08 when there were many deep casts with many bottles below the OMZ, the fit was:

CTD DO Corrected = CTD DO * 1.0176 + 0.051

When the offset for 2016-62 is forced to be 0.051 the result is:

CTD DO Corrected = CTD DO * 1.0180 + 0.051

While the slope looks very close to the earlier cruise, the fit is not as good.

The fit is very close to that for 2016-12 which was based on a few deep casts in the same region in July 2016:

CTD DO Corrected = CTD DO * 1.0256 + 0.0324

The only significant outliers are clearly due to bad CTD data, which is reflected in both high standard deviation in the SBE DO while the CTD was stopped and by examining plots and seeing some very strange features that are most prominent near the surface. It was noted at sea that the dissolved oxygen trace looked strange at times. Yet, the comparison overall looks normal.

Sample #263 from event #79 was flagged 46 due to poor replicates. The CTD DO reads 3.668 which would be 3.782mL/L after recalibration.  Compared to the CTD DO the lower duplicate value, 3.638mL/L looks out of line. Using the other value, 3.814mL/L, the difference falls within the scatter that was included in the fit. The CTD DO looks ok for that bottle. The analyst rejected the outlier dup and changed the flag and comment.

For more details see document 2016-62-dox-comp1.xlsx.

Two cases were identified during which the CTD DO values were bad in the bottle files:

- Event 31: The Oxygen:Dissolved:SBE value for Niskin #16 was replaced with a pad value as the sensor clearly malfunctioned during the bottle stop.

- Event #123: The Oxygen:Dissolved:SBE value for Niskin #3 was replaced with a pad value as the sensor clearly malfunctioned during the bottle stop.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Salinity

Compare was run with pressure as reference channel.

When outliers were excluded based on differences and standard deviation in the CTD salinity, the primary salinity was found to be low by an average of 0.0106 and the secondary by 0.0094 with standard deviations of ~0.003psu for both. Both fits against pressure were very flat with the secondary being especially so.

Ignoring the cases in which the standard deviation in the CTD salinity was >0.0008 there were 5 significant outliers:

- Sample 32, event 22, ~5m – The standard deviation in the CTD salinity is low but the local gradient is high and even slightly poor flushing could easily account for the bottle being higher than the CTD by 0.04psu.

- Sample 82, event 31 – The bottle value is about 0.02 higher than the CTD at ~1000m. The analyst had flagged this value as 46 since there was poor agreement between the replicates. Looked at separately, the replicates would lead to the bottles being higher than the CTD primary salinity by 0.03or 0.01, so the lower value falls into line with the general comparison. The analyst decided to use the lower value with flag 2 and explanation.

- Sample 105, event 40 – The bottle is higher than the CTD by 0.02. There is no evidence of trouble in the CTD salinity. This could be due to a poor seal leading to more evaporation of sample than seen in other bottles. There could be some effect from incomplete flushing, though the CTD descent rate suggests that the Niskin was being shaken up quite a lot and the local gradient is not large. There is no evidence that the problem is in the analysis. A flag 36 was used with a note that there may have been significant evaporation of the samples.

- Sample 295, event #98 – The bottle is higher than the CTD by ~0.02psu. Though the standard deviation in the CTD salinity was low during the 10s around firing time, there was quite a lot of variability if a wider window is examined. No flag was recommended as the value likely represents the contents of the Niskin bottle.

- Sample 334, event #113 – The bottle was higher than the CTD by ~0.02psu, the standard deviation in the CTD salinity was low but there was quite a lot of variability if a wider window is examined. The value may reflect the Niskin contents. No flag is justified.

Pressure dependence is usually seen in such comparisons as any small variations in alignment and local gradients become less significant with depth. For these data the trendline looks flat in differences against pressure for both sensor pairs, but the scatter makes that an unconvincing result.

Plots against event numbers suggest some slight time-dependence. This was investigated in the secondary salinity. To ensure that the trend was not caused by different pressure ranges in the different casts, a plot was made selecting only data between 450 and 1100db. When 2 values are excluded that were found to be outliers in the fit against pressure, the average difference showed the secondary salinity low by 0.010psu, but the standard deviation is 0.004psu. The trendline shows the first cast being low by 0.014 and the last by 0.007psu, but the scatter is high. The difference looks like too much to be explained by the first bottles waiting for analysis by about a week longer than the later ones. For a few cases with large differences there is evidence that a shed wake may have filled the Niskin bottle with water from deeper in the water column, though the waits were long enough and there was some vertical motion that should help flush the bottle. There may also be variations in local gradients that would make poorer flushing more or less significant, but there are insufficient data to resolve such factors.

The salinity bottle inserts in use for some recent cruises have been found to be of lower quality than usual, so seals may well have been poor leading to more evaporation than expected.

The evidence is too weak to justify recalibration of salinity.

For full details for the COMPARE run see file 2016-62-sal-comp1.xls.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor.

The plot of FL/CHL vs CHL below shows the usual pattern with CTD fluorescence mostly reading higher than the extracted CHL for CHL<1ug/L and then falling relative to CHL until it is about 50% of CHL at high CHL values. There were few very low chlorophyll values, with only 3% of the values having CHL<0.2ug/L compared to 23% being <0.2ug/L for the August Line P cruise.

A fit forced through the origin has a slope of about 0.64, which is close to the results found during 2016-47 in May 2016.

For full details of the comparison see file 2016-62-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on a few casts to see what alignment setting would produce a vertical difference between upcast and downcast dissolved oxygen traces and those from temperature. A setting of +2.5s looked best for some features and +3s for others, but in general the temperature was too noisy and temporal changes too large for a good comparison. During the most recent use of this equipment +2.5s was found best.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. During other recent cruises, the default setting of (α = 0.0245, β=9.5) looked better for both conductivity channels than others tested. Two casts were checked for this cruise and the default setting does improve the data.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. The shaded values are from a previous cruise that used these sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2016-47-0039 | 1000 | -0.0019 | +0.0001 | +0.0032 | High, V Noisy |
|  | 1900 | -0.0018 | +0.0002 | +0.0040 | “ |
| 2016-47-0105 | 1000 | -0.0017 | +0.00015 | +0.0038 | High, Noisy |
|  | 1900 | -0.0018 | +0.0002 | +0.0045 | “ |
| 2016-06-0029 | 1000 | -0.0021 | +0.0003XN | +0.0045 VN | High, Moderate |
|  | 1900 | -0.0017 | +0.0003 | +0.0050 | “ |
| 2016-06-0068 | 1000 | -0.0016 | +0.0002 | +0.0045 | High, Moderate |
|  | 1900 | -0.0016 | +0.0003 | +0.0052 | “ |
|  | 3800 | -0.0020 | +0.0003 | +0.0059 | “ |
| 2016-08-0054 | 1000 | +0.0003 | +0.0003 | +0.0038 | F. High, Noisy |
|  | 1900 | +0.0007 | +0.0004 | +0.0042 | “ |
|  | 2900 | +0.0013 | +0.0004 | +0.0040 | “ |
|  | 3400 | +0.0016 | +0.0005 | +0.0038 | “ |
| 2016-08-0088 | 1000 | +0.0001 | +0.0003 | +0.0041 | High, Noisy |
|  | 1900 | +0.0007 | +0.0004 | +0.0040 | “ |
|  | 2900 | +0.0012 | +0.0004 | +0.0038 | “ |
|  | 3400 | +0.0015 | +0.0004 | +0.0036 | “ |
| 2016-62-0039 | 1000 | +0.0001 XN | -0.0002 XN | -0.0022 XN | High, V Noisy |
|  | 1900 | +0.0009 | -0.0001 | -0.0015 | High, Noisy |
| 2016-62-0067 | 1000 | -0.0001 | 0 | -0.0005 | High, Noisy |
|  | 1900 | +0.0008 | 0 | -0.0004 | High, Noisy |
| 2016-62-0092 | 1000 | +0.0003 VN | 0 | +0.0001 | High, V Noisy |
|  | 1900 | +0.0005 VN | +0.0001 | +0.0003 | High, Noisy |
| 2016-62-0125 | 1000 | +0.0001 | 0 | +0.0002 | High, Noisy |
|  | 1900 | +0.0008 | +0.0001 | +0.0005 | “ |

Note that the temperature sensors were switched between 2016-06 and 2016-08.

The differences were very noisy so the values are very rough estimates. The temperature differences are similar to those during 2016-08 from the same depths. Conductivity differences are smaller. The salinity differences are smaller than during 2016-08; they start with the opposite sign to the earlier cruise and then gradually shift to increasingly positive differences as time passes.

The salinity comparison suggests that the difference between primary salinity and bottles is similar to that seen during 2016-08 while the secondary is slightly farther from bottles. Given the randomness introduced by the wait for analysis, it is not justified to consider this evidence that the secondary calibration has drifted more than the primary.

The CTD is lower than bottles by more early in the cruise than later, which could be partly due to more deep casts later. However, when only data from 450m to 1100m are included that trend is still seen. This could partly be due to the fact that it was a long cruise so that early bottles had even longer to suffer from evaporation and adsorption than the later ones. The plot against time has a slightly higher slope for the secondary salinity than for the primary. Using the fits the difference (secondary salinity – primary salinity) is -0.0016 during the first cast with salinity sampling and +0.0006 for the final bottle cast. This fits the results of the downcast differences quite well. During 2016-08 when the same pressure range is selected, there is less time-dependence and the two sensors behaved in a similar way with the primary having slightly less time-dependence.

The primary salinity shows a slightly higher pressure dependence in the salinity comparison when the same points are used for both fits. The difference between the fit at 0db and 2000db is 0.002 for the primary and 0.0002 for the secondary. Using the fits we find (Sal1 – Sal0) = +0.002psu at the surface and +0.00008psu at 2000m. During 2016-08 a similar calculation leads to a difference of 0.005psu at the surface and +0.002psu at 2000db. The pressure dependence could be due to a problem with the conductivity cell of the primary sensor but the pressure dependence in the temperature sensors suggests that one of them may be the problem. It could also reflect a problem with flow that becomes less serious as vertical gradients decrease. During 2016-08 when the sensor was last used the primary sensor showed slightly less pressure dependence than the secondary.

So while the salinity traces are often noisy, the bottle comparisons are in reasonable agreement with the differences noted during downcasts. The upcasts are very noisy when the CTD is not stopped but the results while stopped do look comparable to downcast data.

While the evidence suggests more of a change to the primary than the secondary, the bottle comparison is not terribly reliable. There is another source of trouble that may also be affecting the primary channels. The primary salinity was found to be very noisy during 2016-08 and it looks even worse now. This may be evidence of eddies around the primary intake, particularly during upcasts, or some other cause of poorer flow that leads to the primary salinity reading slightly lower during downcast and higher during upcasts. That might explain why the downcast salinity differences are smaller than during the previous cruise and why upcast salinity differences are mostly of the opposite sign. While stopped, the sensors would have time to equilibrate, so that differences look similar to those from the downcast. The standard deviations in the 2 salinity channels during bottle stops are similar.

The secondary sensors appear to be the better choice overall. There are some spikes in the secondary salinity but they usually involve just a few scans and are mostly found in upcast data.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

Header Check was run and no problems were found.

A cross-reference list was checked against the log book. One wrong station name and one wrong file name were found and fixed,

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.8db which is reasonable for the Tully. The only data available from above 1m had pumps off, so was not very useful in determining if the pressures are reasonably accurate.

The altimeter and water depth readings from the headers of the CLN files were exported to a spreadsheet. A quick calculation (Pressure*0.99 -Depth +Altimeter Reading) was made; where it differs from 0 by more than 4, the cause was investigated. (Below 1200m 0.985 works better.) Where that value is >4m, a check was made of water depths against the log and a plot of altimetry near the bottom to see if the altimetry header needs to be removed or water depth headers needs alteration. Depth entries were changed to fit the log if the log entry produced a better result. In one case the depth entry was obviously an estimate and it looks like it was not appropriate, so it was removed. In a few cases the sounder was varying, which may be due to real changes or sounder problems – the one that best suited the data from the bottom of the cast was chosen. Altimetry was spiky for one cast, so the header entry does not look reliable and was removed. Corrections were made to the CLN files for events #29, 65, 97, 100, 109, 111 and 197.

Similar data were exported from the SAMAVG files. Checks were made to ensure that for casts with only near-surface sampling there were no faulty altimetry headers due to misinterpreted spikes. The altimetry header was removed from casts 7 and 32, and the bottom depth was corrected for cast #29 in the SAM and SAMAVG files. The merge of SAMAVG with MRGCLN1s was repeated followed by CLEAN.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 5 casts to see what shift in conductivity does the best job of removing noise in the salinity channels There were many large unstable features in T-S space for the primary channels, and no setting appeared to help those. A shift of +0.3 records did improve areas with small-scale unstable features. For the secondary a setting of -0.6 records worked best; there were some large unstable features that were not improved, but many fewer than in the primary channels.

SHIFT was run on all casts using +0.3 records for the primary conductivity.

SHIFT was run on all casts using -0.6 records for the secondary conductivity.

Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

All sensors EXCEPT temperature sensors were used during 4 cruises in 2015 and 6 before this cruise in 2016, one of which had not been processed yet. The pressure offset was updated in June 2016 based on near-surface sampling on 2 cruises. Salinity calibrations were mostly not trusted due to flushing problems, delayed analysis and/or noisy comparisons for most of the cruises, but cruise 2016-06 did provide more reliable information. The secondary salinity has been closer to bottles than the primary for all cruises, reading lower by from 0.002psu to 0psu. The primary salinity read lower than bottles by from 0.006 to 0.008psu. There were large differences between temperature sensors that account for the primary reading lower than the secondary by 0.002. The temperature sensors were replaced before cruise 2016-08 and provided somewhat better data than during 2016-06. The dissolved oxygen comparisons have been reasonably consistent when sampling similar regions.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Salinity values were mostly within the climatology except for slightly low values at about 700m at station LD08 and LJ05. Temperatures were above the climatology maximum at 2 stations in the northern Strait of Georgia between 30db and 70db; higher temperatures were observed in June 2016 in the same area but at that time the excursions were deeper and more extensive. None of these observations suggest a calibration problem.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

As discussed in earlier sections the primary channels are very noisy. The secondary also has spikes, but overall is much better. It was selected for editing and archiving.

The secondary T and S channels were selected for all casts.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. Some bad salinity points were removed. All files required some editing.

After editing T-S plots were examined for all casts which led to a little more editing being applied to one cast.

Next, because some problems had been noted in the dissolved oxygen traces during upcasts, plots were made of the downcast dissolved oxygen traces to see if any editing was required for that channel. Where there were features in DO that were not seen in temperature plots or in other nearby casts, plots of the full casts were made. In each case the upcasts showed the same features. No editing was applied to the DO channel.

##### Recalibration

There is no evidence to suggest that the pressure channel needs further adjustment.

The comparison with bottles suggests that the secondary salinity is low by about 0.009 with a standard deviation of 0.003, but due to analysis being delayed, it is likely that bottle values are high by about 0.005 due to evaporation and adsorption, or possibly by even more given that the bottle inserts may have been of poorer quality than usual. The analysis was run on a Portasal which is not as accurate as the Autosal. Because the precision result was Sp = 0.0025 the analyst reported salinity values with only 3 decimal places rather than the usual 4. There are too many sources of error to justify recalibration of salinity.

Based on the comparisons with bottles described in section 5, 2016-62-recal1.ccf was prepared to apply the following correction to channel Oxygen:Dissolved:SBE:

CTD DO Corrected = CTD DO * 1.0248 + 0.0233

This correction was first applied to the SAM and MRGCLN2 files. COMPARE was rerun and the results confirm that the recalibrations were applied properly. When outliers were removed based on residuals the average of differences in the DO fit was -0.0001mL/L and the standard deviation was 0.015mL/L.

(See file 2016-62-DO-comp2.xlsx for details.)

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When all data were included the CTD DO was higher than the titrated samples by an average of ~0.10mL/L and when some outliers are removed based on residuals, it is high by 0.05mL/L with a standard deviation of 0.05mL/L. There are two known sources of error that do not involve calibration drift: slow response of the DO sensor in high gradients and poor flushing of Niskin bottles. Both will lead to the downcast DO from the sensor reading higher than the bottles above the DO minimum (found ~800-1200m). The differences for 16 samples from below 1000m show the CTD reading very close to the bottles or a little lower, with an average difference of -0.012mL/L and standard deviation of 0.009mL/L Near the surface the DO sensor is close to bottles, but between about 10m and 400m where vertical DO gradients are highest, the DO:SBE values are mostly higher than bottles.

No further recalibration is justified. See 2016-62-dox-comp3.xlsx for details.

##### Fluorescence Processing and special files for Dr. Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and unedited except that

some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5  mL/L from 0 to 200db

±0.25 mL/L from 200db to 400db

±0.08 mL/L from 400db to 800db

±0.02 mL/L below 800db

For details on the processing see document: 2016-62_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

A cross-reference list was produced.

The sensor history was updated.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values ranged from 70% to 130% with most between 100% and 110%. For the casts farthest from shore values ranged from 102% to 105% except for one cast at 109%; typical values for that area are ~103% so that range looks good. The highest values were mostly very close to the west coast of Vancouver Island. The lowest were seen in shallow waters a little offshore and in Rivers Inlet and at some casts in the northern Strait of Georgia. The values suggest that the DO calibration is reasonably good.

##### Final Bottle Files

The MRGCOR2 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and titrated DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2016-62-bottles-final.xlsx.

Standards check and a header check were run on all files. No problems were found.

The track plot looks ok.

Plots of each file were examined to ensure no further problems were found.

A cross-reference listing was produced for the CHE files.

##### Thermosalinograph Data

There were 4 hex files. There were no loop samples taken.

a.) Checking calibrations

The configuration files were all identical. One file was renamed as 2016-62-tsg.xmlcon. No errors were found in the calibration parameters.

b.) Conversion of Files

The 4 files were converted to CNV files using configuration file 2016-62-tsg.con.

Those CNV files were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

A time-series plot showed that the differences between the intake and lab temperature are much larger than usual. This was also seen during the previous cruise and the problem was found to be with a malfunctioning intake temperature sensor. The salinity data look good with just a few small spikes and some sections with small-scale noise that could be due to small bubbles. The flow rate was low and variable at the beginning of the first file, so temperature, salinity and fluorescence should be removed for that section. The flow thereafter was quite steady and the rate was close to 1.

The track plot looks fine. The plot was added to the end of this report.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db and exported to a spreadsheet which was saved as 2016-62-ctd-tsg-comp.xls. There were 110 casts which overlapped with TSG files.

The TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, lab temperature, salinity and fluorescence and the files were reduced to the times of CTD files and loop samples.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude and longitude were all ≤0.0007° and the median differences were 0.0002° for both. This shows both the times and positions are reliable for both systems.

d.) Comparison of T, S and Fl from Rosette Samples and TSG and CTD data

- T1 vs T2 The intake temperature sensor worked throughout the cruise, but the values are unbelievable as they are higher than the lab temperatures. This would only make sense if the intake waters were warmer than the ambient temperature of the ship, which was not the case as the maximum temperature seen in CTD records was about 16.5°C and most near-surface water temperatures were much lower than that, so there should be some warming in the loop. The same problem was noted when the same equipment was used during 2016-08. The differences between the lab and intake are highly variable, but the intake is almost always higher.

- Flow Rate The flow rate was steady in sections with values between 0.98 and 1.04 other than the first part of the first file when it was very low.

- TSG vs CTD The spreadsheet comparing CTD and TSG files was then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. Using all data the results were:

|  | Tlab-Tctd | Tint-Tctd | SALtsg-SALctd | FLtsg/FLfl |
| --- | --- | --- | --- | --- |
| median | 0.218 | 0.903 | -0.012 | 0.121 |
| average | 0.378 | 1.016 | -0.499 | 0.129 |
| stdev | 0.440 | 0.532 | 1.963 | 0.043 |

The intake temperature was higher than the CTD temperature by a median of 0.9C° which is clearly a bad result. Intake temperature will not be archived.

When 20 casts were selected which had the lowest standard deviation in the TSG temperature the lab temperature was higher than the CTD temperature by a median of 0.183C° (standard deviation 0.240). When the 20 casts with the lowest standard deviation in the TSG Salinity were selected Salinity:Lab was low by a median of 0.007psu (standard deviation was 0.008.) The ratio of TSG fluorescence to CTD fluorescence was ~0.15 for the 20 cases with standard deviation in the fluorescence being low, but low standard deviation is likely to be associated with low fluorescence values, so this result may not be generally applicable.

The TSG fluorescence is in volts while the CTD fluorescence is in chlorophyll units.

Forcing a fit through the origin provides a reasonable fit though it overcorrects the TSG fluorescence at low values and may undercorrect at high values. But the CTD Fluorescence itself has a complex fit against extracted chlorophyll, so this is of little value. Without loop samples, making an estimate of chlorophyll content based on the TSG fluorometer is unwise.

(See 2016-62-ctd-tsg-comp.xls.)

- Loop Bottle - TSG Comparisons

There were no loop samples.

- Calibration History

The temperature and conductivity sensors were recalibrated in November 2015 and were used during 2016-01, 2016-47, 2016-06 and 2016-08. For the first two there was too much scatter in comparisons to justify recalibration of salinity. There were also variations in flow rates that complicated comparisons for 2016-01.  Salinity was recalibrated for 2016-06 by adding 0.015psu based on loop sampling while underway; the differences were somewhat lower at ~0.008psu while stopped. For 2016-08 salinity analysis was delayed so the results of the previous cruise were used for recalibration. Because the intake thermistor malfunctioned a proxy was derived by subtracting 0.18Cº from the lab temperature. The fluorescence data were archived in volts for all 4 cruises.

Conclusions

1. The TSG clock worked well.

2. The TSG flow rate was very low for the first half hour of file 1, but was otherwise very steady in sections with small jumps between sections with a range from 0.98 to 1.06.

3. The temperature in the loop is higher than CTD temperatures by between 0.18 & 0.22Cº with the results considered most reliable near the bottom of that range.

4. The TSG intake temperature values are bad, either due to malfunction of the sensor, or possibly something covering it. The data were also bad during the previous cruise, 2016-08. The correction applied to the lab temperature in order to create a proxy for intake temperature for cruise 2016-08 looks appropriate for this cruise. While the correction might be slightly low, there was a lot of scatter in the 2016-62 results, likely due to the surface gradients being highly variable.

5. The TSG Salinity is lower than the CTD salinity by a median value between 0.007psu and 0.012psu depending on which data are selected. Based on loop samples during the previous 2 cruises it is likely that it reads low by more than that while underway. A correction of -0.0015 was used for those cruises. If, however, bubbles account for the lower TSG salinity while underway, then this cruise could conceivably have been less affected due to either different sea states or steaming speeds. In the absence of loop samples and given the large standard deviations in the results, it does look best to use the results from 2016-08.

6. The relationship between TSG fluorescence and CTD fluorescence is complex as is that between the CTD fluorescence and extracted chlorophyll values. There were no loop chlorophyll samples, so no estimate is offered as to how to convert from voltage units to chlorophyll units for the TSG.

f.) Editing

The ATC files were copied to *.EDT.

Each file was opened in CTDEDIT and temperature, salinity, flow rate and fluorescence channels were examined. Only file #1 required editing: T, S and FL points were removed from the first 68 records due to very low flow rate and salinity was cleaned very lightly.

g.) Recalibration

ADD CHANNEL was used to add channel Temperature:Lab and to set it equal to Temperature; Primary. File 2016-62-tsg-recal1.ccf was used to add 0.015 to channel Salinity:T0:C0 and to subtract 0.18 from Temperature:Primary.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Secondary, Temperature:Difference, Conductivity:Primary, Flag and Position:New.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats.

Those files were saved as TOB files.

The Standards Check and Header Check were run and no problems were found.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

# Particulars

PAR OFF:  None

Out of Order: 31& 65

9. Stop on way up at 110db because of wire angle.

40. Sounder jumping from 1410 to 1460.

49. DO drop from 3 to 1at 100m.

67. Stop on way down at 233db because of wire angle.

94. Wrong station name in file – should be LBP8, NOT LBP4. Changed.

109. Wrong depth in hex – should be 1334. Changed.

111. Wrong depth – Log entry 1384, then comment giving 1520; file suggests about 1420. Changed to 1420.

112. File name wrong – should be event 111. Changed to 111.

123. Odd DO data between 50 and 40m of upcast.

159. Odd jump in oxygen trace on upcast. Following downcast profile but offset above 250m.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4752 | 16Dec2015 | Factory |  |  |
| Conductivity | 2280 | 19Dec2014 | Factory |  |  |
| Secondary Temp. | 2710 | 16Dec2015 | Factory |  |  |
| Secondary Cond. | 2754 | 19Dec2014 | Factory |  |  |
| Transmissometer | 1185DR | 27Apr016 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 17Jan2015 | Factory |  |  |
| PAR | 70613 | 21Mar2016 | Factory |  |  |
| SeaPoint Fluor. | 3640 | n/a |  |  |  |
| Pressure Sensor | 0443 | 17Dec2014 | Factory |  |  |
| Altimeter | 62354 | n/a | Factory |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3411       Cruise ID#:	2016-62

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3411 | 7Nov15 | Factory |  |  |
| Conductivity | 3411 | 7Nov15 | Factory |  |  |
| WetLabs Fluorometer | Ws3s-953p |  |  |  |  |
| Temperature:Secondary | ? | 0842 |  |  |  |
| Flow meter | ? | n/a |  |  |  |