## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2016-67

Agency: Ocean Networks Canada / NRCan

Location: Strait of Georgia

Projects: ONC/NRCAN Strait of Georgia

Party Chiefs: Macoun P.

Platform: John P. Tully

Date: October 5, 2016 –October 11, 2016

Processed by: Germaine Gatien

Date of Processing: 24 January 2017 – 26 January 2017

Number of original HEX files: 13	Number of CTD files: 13

Note: 6 BOT files were prepared for possible later use – not intended for OSD_DATA_LIBRARY

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0443) was used for this cruise. It was mounted in a rosette and attached were a WetLabs CSTAR transmissometer (#1185DR), a SBE 43 DO sensor (#0997) on the secondary pump, a SeaPoint Fluorometer (#3640) on the primary pump, a PAR sensor (#04565) and an altimeter (62354).

The data logging computer was the Tully CTD Laptop (Acer).

The data acquisition program was Seasave.

The CTD deck unit was an SBE model 11+, #425.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log book lacked a list of Scientific Personnel. There was a complete equipment list.

There were no bottle samples available at the time of processing, but some sampling had been done, so 6 BOT files were prepared with CTD data from a 10s window centred on the time of bottle firing. These will not be placed in the OSD Data archive, but they are available, if needed.

There was no calibration sampling available. The result of the most recent cruise that used the same equipment was used to recalibrate dissolved oxygen. Pressure was recalibrated based on deck measurements.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. There was sampling from rosette bottles, but none of the analysis was done at IOS. There were dissolved oxygen samples analyzed at UVIC using old equipment.

The history of the pressure sensor, conductivity and DO sensors were checked.

The offset for the pressure sensor was decreased by 0.4db in June 2016 and the observations during this cruise show a further decrease of between 0.2db and 0.3db is appropriate. Rather than change the pressure at the conversion stage, an adjustment will be made later when more information is available.

The only change to the XMLCON files since this equipment was last used is the PAR sensor.

No errors were found in the configuration file used at sea, so one such file was renamed as 2016-67-ctd.xmlcon. There were no changes in configuration during the cruise.

(Note that a further pressure adjustment of -0.3db was made later in processing for a net offset of -1.3db.)

##### Conversion of Full Files from Raw Data

All hex files were converted using 2016-67-ctd.xmlcon to create CNV files.

A few casts were examined.  All expected channels are present. The primary and secondary temperature and conductivity channels are reasonably close during downcasts, but as usual the upcasts differ more due to noise in both channels. Dissolved oxygen, fluorescence, altimetry and transmissivity profiles look normal. The descent rate was generally fairly high and steady.

PAR values are low raising the question about whether it was really mounted on the CTD. However, most casts were at night and late enough that little light would be expected. This will be checked later.

The initial soak data are included in files 9, 12 and 13; the opening scans will need to be removed.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2016-67-ctd.xmlcon.

The ROS files were converted to IOS format.  They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and no problems were seen that required editing.

A preliminary header check was done and no problems were found.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run to determine which setting was best for aligning DO with temperature. During the last 2 cruises using this sensor a setting of 2.5s has been used while +3s was used for many other cruises. For this cruise +2.5s looks best. ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. Tests on previous cruises using these sensors showed the default setting of (α = 0.0245, β=9.5) did the best job and it does improve the data for both conductivity channels for these data.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

Tests are usually run at this point to see if the differences between sensor pairs have changed significantly. Unfortunately, the deepest cast was <300m, so the results are not very reliable for conductivity and salinity which are somewhat depth dependent. DERIVE was run a second time on the deepest cast to examine differences between sensor pairs. Differences from other cruises using the same equipment are also shown with dark shading.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2016-08-0088 | 1000 | +0.0001 | +0.0003 | +0.0041 | High, Noisy |
|  | 1900 | +0.0007 | +0.0004 | +0.0040 | “ |
|  | 2900 | +0.0012 | +0.0004 | +0.0038 | “ |
|  | 3400 | +0.0015 | +0.0004 | +0.0036 | “ |
| 2016-62-0067 | 1000 | -0.0001 | 0 | -0.0005 | High, Noisy |
|  | 1900 | +0.0008 | 0 | -0.0004 | High, Noisy |
| 2016-62-0092 | 1000 | +0.0003 VN | 0 | +0.0001 | High, V Noisy |
|  | 1900 | +0.0005 VN | +0.0001 | +0.0003 | High, Noisy |
| 2016-62-0125 | 1000 | +0.0001 | 0 | +0.0002 | High, Noisy |
|  | 1900 | +0.0008 | +0.0001 | +0.0005 | “ |
| 2016-67-0004 | 290 | +0.0004 | +0.0001 | +0.0007 | High, F. Steady |

The differences are all small.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers.

CLEAN files were copied to *.CLIP

CLIP was run to remove 5000, 7700 and 4400 records from casts 9, 12 and 13 respectively, to remove the data collected during the initial drop to 10m for the soak.

##### Checking Headers

A cross-reference list was checked against the log book. The times in the headers are a few minutes earlier than those in the log book, but the former would come from when the file was started and the latter is likely from when acquisition started; no changes were made. No problems were found.

A header check was run. A few potential problems were investigated:

- The 3X gain cable seemed odd for October, but the maximum value of 4.4ug/L seems reasonable. It was likely just left in place after the previous cruise in early September.

- The bottle files had extremely low PAR values which raised the possibility that the sensor was not actually mounted. However, the casts were mostly at night, and all are at times when light would be low. The maximum PAR reading in the CLIP files was 36 which shows that the sensor was mounted. That reading comes from the end of a cast when the CTD may have been out of water.

- There are some negative pressure readings towards the end of cast #12. If we apply an offset to reduce pressure there will be more negative readings, but they will disappear after DELETE is run. The negative values appear to be small spikes in a section where pressure is varying significantly, probably just as the CTD left the water. The readings taken on deck at the end of casts were 0.29 to 0.34db. For cast #12 when the bottles were out of the water but the CTD in the water, the reading was ~0.2db before archiving began.

The deck measurements made at sea as well as the observations from cast #12 all suggest that subtracting 0.3db is appropriate.

Cruise tracks were plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.2db before correcting pressure and after removing the initial soak data. This is lower than usual for the Tully, but reasonable for the area and weather conditions.

The altimeter and depth readings from the headers of the CLN files were exported to a spreadsheet to check that they are reasonable. A rough check was made by subtracting the altimeter reading from the depth entered in the header to find the maximum depth sampled. That was then compared with the maximum pressure recorded in the files using a rough estimate that depth is ~99% of pressure. When those two estimates differed by more than 2m, the bottom depth was checked against the log book. In 4 cases editing (or adding) the bottom depth in the file to match the log book entry was sufficient to make the check value less than 2db. In one case the altimetry header was found to be inaccurate – there had been a stop a little above the bottom which likely had an adverse effect on the algorithm. The value was changed from 3.6 to 2.5m. None of the changes involve files with rosette data. All other altimetry headers looked appropriate.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and overall the choice made earlier looks appropriate, so no further alignment was applied.

Conductivity

The choice for 2016-62 when these sensors were last used was a shift of +0.3 records for the primary and -0.6 records for the secondary. Tests were run on 2 casts using a variety of shifts of a similar magnitude. While the differences among the choices were small the values that worked best for these data were the same as for 2016-62.

SHIFT was run on all casts using +0.3records to the primary conductivity and -0.6records to the secondary conductivity; both salinity channels were recalculated.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

The DEL files were copied to *.EDT.

##### Other Comparisons

Previous experience with these sensors –

Salinity: The T/C sensors were used during 2016-08 and 2016-62. The primary and secondary salinity were both found to be lower than bottles, but for both there were delays in analysis and for 2016-62 the analysis was done on a Portasal, rather than the usual Autosal. An estimate was made for 2016-08 that the salinity was within 0.002psu. No recalibration was applied to either cruise as there was insufficient confidence in the results. When the conductivity sensors were used with different temperature sensors during 4 earlier cruises, the primary salinity was also considered to be within 0.002psu.

Pressure: During 2016-08 and 2015-62 there was little evidence and no adjustment was made.

Dissolved Oxygen: This sensor had been used for 11 other cruises since its last factory service. The corrections applied have varied. Cruises that include a lot of deep casts tend to smaller corrections than those that have shallower sampling. The correction for the most recent use during 2016-62 was based on many bottles and was similar to that of 2016-12; both included many shallow casts.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Local climatology was not available for Saanich Inlet. The only excursions from the local climatology found were some slightly low salinity values in the top 5m at 3 of the 4 casts in the Strait of Georgia. This is not considered evidence of calibration problems.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The primary channels were selected for archiving for only 1 of the other 11 cruises that used this equipment since it was last calibrated at the factory. The secondary channels have slightly less spiking in these files, so they were selected for editing and archiving.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts.

All casts required editing.

##### Initial Recalibration

The pressure offset will be adjusted by subtracting 0.3db based on deck measurements.

During 2016-62 no recalibration was applied to salinity because the results were not trusted; the results of other recent cruises suggested that Salinity:T1:C1 was good to ±0.002psu.

Oxygen:Dissolved:SBE will be recalibrated based on the results of cruise 2016-62.

CALIBRATE was run on the EDT files using file 2016-67-recal1.ccf to subtract 0.3db from the pressure channel and to correct the Oxygen:Dissolved:SBE channel using:

CTD DO Corrected = CTD DO * 1.0248 + 0.0233

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots look fine.

Profiles were examined alone and in groups of nearby casts and no significant outliers were seen.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and unedited except that

some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

Oxygen:Dissolved:SBE was recalibrated based on the results of cruise 2016-62.

For details on the processing see document: 2016-67_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

A file list was produced and looks fine. The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values ranged from 65% to 100%, with the highest values at the sill of Saanich Inlet and at 2 casts in the Strait of Georgia. Those casts had higher vertical DO gradients near the surface than the casts with lower saturation. Where vertical gradients were lower there was likely more mixing down of DO. The highest values suggest that the DO calibration is reasonable, but the evidence is much too weak to suggest how accurate the DO data are.

As expected, all casts that went deeper than 100m in Saanich Inlet have very low DO values at depth, ~0.1mL/L. For casts #7 and #9, in the middle of the BPS line, DO rises right at the bottom to ~0.4ml/L and ~0.2mL/L respectively.

##### Producing final files

A cross-reference listing was produced for CTD files.

The sensor history was updated.

# Particulars

7. The CTD was left running for a deck pressure reading 0.286db

8. Deck pressure 0.322db

9. Starting deck pressure 0.609, end 0.339

10. Deck pressure end 0.305

11. Deck pressure after cast 0.318

12. At 10:54 the CTD was at ~0.2m and bottles in air. At the end of the cast the pressure read 0.28 when the CTD was just above the water line and 0.287 on deck.

CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4752 | 16Dec2015 | Factory |  |  |
| Conductivity | 2280 | 19Dec2014 | Factory |  |  |
| Secondary Temp. | 2710 | 16Dec2015 | Factory |  |  |
| Secondary Cond. | 2754 | 19Dec2014 | Factory |  |  |
| Transmissometer | 1185DR | 27Apr016 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 17Jan2015 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | Factory |  |  |
| SeaPoint Fluor. | 3640 | n/a |  |  |  |
| Pressure Sensor | 0443 | 17Dec2014 | Factory |  |  |
| Altimeter | 62354 | n/a | Factory |  |  |