## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 20 March 2025 | Updated channel names & formats in TOB files. G.G. |
| ### 25 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| 16 Aug 2020 | Added HPLC Data. S.H. |

## PROCESSING NOTES

Cruise: 2016-71

Agency: Ocean Sciences Division

Location: Strait of Georgia and Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Strait Water Properties Survey

Party Chief: Fraser T.

Platform: Vector

Date: October 30, 2016 – November 5, 2016

Processed by: Germaine Gatien

Date of Processing: 11 January 2017 – 1 March 2017

Number of original HEX files: 64

Number of CTD files: 63	Number of bottle files: 23

Number of TSG files: 3

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#1119) on the secondary pump, a SeaPoint Fluorometer (#3865) on the primary pump with a 3X cable, a Biospherical QSP-400 PAR sensor (#4565), Surface PAR (#20518) a pH sensor (#0692) and an altimeter (#1252).

The data acquisition program was Seasave v7.26.2.13.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log book contained no equipment list. There were also no photos. Since the serial numbers of external sensors are not recorded in the hex files, there is no way to confirm that the calibration parameters entered for those sensors are correct. The equipment used is believed to be the same as used for cruise 2016-10 in late September 2016 based on the configuration file and the technician’s memory. But the conversion process does not fail if the serial numbers of external sensors do not match the configuration file, so there remains some chance that they could have been switched. There is also no way to determine on which pump the dissolved oxygen sensor and fluorometer were mounted for either this cruise or the previous one that used the equipment. When there is a problem with a pump this information is helpful in establishing what happened. The entries for events were clear with good notes in the log book about problems encountered.

The sampling/rosette log sheets were generally in good order although two casts had the wrong event numbers.

The start time was wrong in file 2016-71-0008. There was a net cast before the rosette cast. The NMEA data matches the log entries for the NET cast which was event #7, but not the ROS cast that occurred 30 minutes later. The acquisition time is present in the header of file #8 and matches the log time for event #8. The start time in the header was adjusted to match the acquisition time, but the positions are from the earlier time. Based on log entries, the error in positions is not significant. However, starting the file so early is a bad idea since it could lead to problems with positions if the ship drifts.

Because the last 2 bottles were fired in air for event #8, a second cast was run immediately after to get the missing samples. The bottle files for events #8 and 9 were combined to form a single file, named 2016-71-0008.CHE. The pumps were not turned on for the second cast so all pumped CTD channels were replaced with pad values.

Salinity analysis was done promptly. Based on 2 casts near the mouth of Juan de Fuca Strait where flushing of bottles is likely to be good, both CTD salinity channels were found to be within 0.001psu of bottle salinity. Using more casts, they are within 0.002psu. No recalibration was applied.

Channel Oxygen:Dissolved:SBE was recalibrated. The downcast Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.3 mL/L from 0 to 150db

±0.08 mL/L from 150db to 300db

±0.03 mL/L below 300db

As some of these errors may be due to incomplete flushing of Niskin bottles, this may underestimate the accuracy of the dissolved oxygen sensor during downcasts.

The fit of CTD fluorescence versus extracted chlorophyll showed the usual pattern with the ratio of CTD Fluorescence to Extracted CHL samples starting at about 2 for CHL ~0.1ug/L and gradually falling to about 0.5 for CHL=2.5ug.

Channel pH:SBE was removed because the sensor malfunctioned throughout the cruise.

There are 3 TSG files. The Vector TSG set-up has no loop sampling, flow rate meter or intake temperature sensor. Comparisons with CTD data are the only means of assessing accuracy. Comparisons were made but are limited by the fact that there are few CTD data available from the intake level of 2m and the quality of that data is limited by high variability and ship effects. The comparison of CTD and TSG data at 2m and 3m suggest that the calibrations of the TSG sensors are reasonable, but that the water in the loop is being drawn from higher in the water column than the intake level.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The sensor history was found for the pressure, conductivity and DO sensors. This was the 9th use of the pressure sensor and the 6th use for the conductivity and DO sensors since they were last calibrated at the factory.

The configuration file did not change through the cruise. The calibration constants were checked for all instruments and no problems were found. One file was saved as 2016-71-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2016-71-ctd.xmlcon to create CNV files.

A few casts were examined.  All expected channels are present. The primary and secondary temperature and conductivity channels are close during downcasts, but as usual the upcasts differ more due to noise in both channels. There are some spikes in both primary and secondary conductivity and temperature. Dissolved oxygen, fluorescence, PAR, SPAR, altimetry and transmissivity profiles all look normal.

The descent rate was mostly fairly high and steady, except near the mouth of Juan de Fuca Strait where it is noisy.

The pH signal looks bad with unbelievably high values. Problems were noted with this sensor late in 2016-45 and throughout 2016-10. The pH data should be removed later.

##### BOTTLE FILE PREPARATION

ROS files were created using file 2016-71-ctd.xmlcon.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and a few spiky values were seen in the primary salinity in 3 files and the secondary salinity in 1 file.

CTDEDIT was used to edit the BOT files for events #8, 12, 24 and 57. Comments about editing were added to the headers. The output files were then copied to BOT.

Because 2 surface bottles fired in air during cast #8, a second cast was run to fire close the bottles again – that event was saved as event #9. The BOT files for the two events files were renamed as 2016-71-008.BOTx and 2016-71-008.BOTy.  The data for Niskin bottles #14 and 15 were removed from file 2016-71-0008.BOTx. The bottle numbers were changed from 1 & 2 to 16 & 17 for file BOTy. The two segments were then joined as 2016-71-008.BOT. A header check was done and it was discovered that the pumps were not turned on for the second cast, so file 2016-71-0008.boty was edited to replace all pumped channels with pad values and the JOIN step was repeated.

The header check also confirmed that the pH values are unrealistic. This sensor malfunctioned for the last few casts of 2016-46 and all of 2016-10.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records. For event 1 all bottles were fired, but only the first 10 were assigned sample numbers, so the lines for Niskin bottles 11-20 were removed from the ADDSAMP file.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine. Those files were bin averaged on bottle number.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2016-71-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF 2016-71CHL*.xlsx. The file included comments, flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2016-71chl.csv, event numbers were added, and the file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2016-71oxy*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2016-71oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in QF2016-71SAL*.xlsx. There were no duplicates. The analysis was done within 18-23 days of collection. The files were simplified and saved as 2016-71sal.csv.

That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet NUTS_QF2016-71nuts.xlsx. This includes a precision study. The spreadsheet was simplified, reordered on sample numbers and saved as 2016-71nuts.csv.

The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only, with output files named MRGCLN1.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number as the merge channel.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. Two casts were missing analysis data because the files had the wrong event number. There was due to errors on the rosette log sheets:

- The event identified as 51 on the rosette sheet should have been 50.

- The event identified as 53 on the rosette sheet should have been 52.

The merge process was repeated and the data exported to a spreadsheet again. No further problems were found.

CLEAN was run on the MRG files to add 0 flags to empty flag channels and to update header limits.

##### Compare

During 2 casts, 39 and 79, the CTD touched bottom, but neither were rosette casts.

Salinity

Compare was run with pressure as reference channel.

The analysis was done promptly after collection so we do not expect significant evaporation or adsorption. The descent rate was very steady for some of the casts, so there may be incomplete flushing of some bottles.

When outliers were excluded based on standard deviations in the CTD salinity the primary salinity was found to be low by an average of 0.0008psu and the secondary was low by 0.0016psu. The standard deviations for both were ~0.002psu. Using only the 2 bottle casts closest to the mouth of Juan de Fuca Strait where the descent rate was quite noisy, the primary is found to be high by 0.0006psu and the secondary low by 0.0009psu. The 2 most significant outliers were both associated with very high standard deviations in the CTD salinity.

The fits were not very flat against pressure, but dropping a few more outliers did not help much. There are likely just not enough data to achieve such a fit, especially if the effect of incomplete flushing varies through the cruise.

For full details for the COMPARE run see file 2016-71-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The cast from Saanich Inlet was excluded from the comparison (as usual) because most bottles from that site do not fit the rest of the cruise because of the extreme gradients. The only other cases with DO<1mL/L come from Indian Arm where poor flushing is likely due to very steady descent rates, so the bottle values are likely a little low, which accounts for the CTD values looking high. The gradients are small so the error would be small, but near the origin the effect on a fit can be significant. Errors due to calibration drift are always such that the CTD looks low compared to bottles, so these high values near the origin do appear to be the result of poor flushing. When the offset is set equal to 0, the Indian Arm bottles do look out of line. Some other bottles were excluded due to high standard deviations in CTD DO and some as outliers based on residuals.  So the offset was forced to be 0 and the fit found was:

CTD DO Corrected = CTD DO * 1.0284

When only the 4 casts from near the mouth of Juan de Fuca Strait were used, the fit was:

CTD DO Corrected = CTD DO * 1.029

The DO profiles near the bottom in some Juan de Fuca casts show the CTD DO still falling during the beginning of the upcast. This may explain some cases with differences above the trendline at DO~2mL/L.

In the central Strait of Georgia there were large mid-depth DO reversals, so any effect of poor flushing would vary in sign. The fit there was:

CTD DO Corrected = CTD DO * 1.025

Similarly, in Indian Arm the DO gradient had a different sign above 100db from the one below it. The Indian Arm fit is too dependent on the method of choosing outliers to be trusted.

There is little difference among these fits. This may indicate that flushing is not a major problem either because bottles flushed well, or because vertical DO gradients were low and of variable sign.

The slopes are a little higher than in other recent uses of this equipment. Poor flushing in areas with DO increasing steadily with pressure would lead to underestimates of DO correction.  So this may be a better fit than we usually see. It may also be a result of calibration drift. The history of the sensors does show the slope generally increasing slowly.

Most of the significant outliers are explained by high standard deviation in the CTD dissolved oxygen channel which is generally due to high local vertical gradients which will lead to a mismatch between what the bottles contain and what the CTD is measuring. A few outliers that could not be explained in that way were:

- Event #29 at 2db. The standard deviation in the CTD DO was low, but there was a high vertical gradient in DO at this level, so minor inefficiency in flushing could explain the difference.

- Event #57. 75db. The standard deviation in the CTD DO was low. The CTD is lower than the bottle by ~2.5mL/L. The difference is of the wrong sign to suggest a flushing problem and the descent rate is noisy enough to suggest this is not a cast where flushing should be a big problem. The gradient is quite high above the bottle depth, but not high enough that the distance between bottle and CTD could produce such a large difference. All bottles from this cast were flagged 3 due to a problem in how samples were handled, but only this one stands out. Sample 202 was flagged 4.

- Event #77, 4.8db. The standard deviation in the CTD DO was low, but there was a high vertical gradient in DO at this level, so even slight inefficiency in flushing could explain the difference.

For more detail see document 2016-71-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further problems were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor. There were 58 samples (1 sample had no CTD fluorescence for comparison) with a range of ~0.09ug/L to ~2.4ug/L. The fit is:

SBE Fluorescence = 0.81 * CHL

As usual, for low CHL the fluorescence tends to be higher than CHL, reads close to CHL at 1ug/L and then drops relative to CHL. For details see file 2016-71-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on ALIGNCTD to determine the best setting to advance the DO signal. When this sensor was used in June the best results were +2.5s for one cruise and +1.5s for the other and in September +2.0s was used. Tests were run using a variety of settings and the best results were with 1.5s.

ALIGNCTD was run on all casts using +1.5s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. The default setting of was selected and it does improve the data for both conductivity channels for these data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 4 of the deeper casts to calculate differences between sensor pairs. None of the casts are very deep. The shaded entries are from 2 other recent cruises during which these sensors were used.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2016-05-0060 | 350 | -0.0002 | +0.0001 N | +0.0013 | High, Steady |
| 2016-05-0085 | 350 | -0.0001 | +0.0001 N | +0.0012 | High, Steady |
| 2016-05-0101 | 330 | -0.0002 | +0.0001 N | +0.0013 | High, F.Steady |
| 2016-07-0069 | 350 | -0.0005 | -0.0001 | -0.0010 | High, F. Steady |
| 2016-07-0108 | 350 | -0.0002 | -0.0001 | -0.0011 | High, F. Steady |
| 2016-10-0022 | 351 | -0.0003 | -0.00005 | -0.0003 | High F.Steady |
| 2016-10-0044 | 343 | -0.0002 | -0.00007 | -0.0004 | High, steady |
| 2016-10-0052 | 275 | -0.0004 | -0.0007 | -0.0003 | High, steady |
| 2016-71-0002 | 300 | -0.0004 | -0.0001 | -0.0006 | F.High, Steady |
| 2016-71-0055 | 340 | -0.0002 | -0.0001 | -0.0008 | High, Noisy |
| 2016-71-0072 | 325 | -0.0003 | -0.0001 | -0.0006 | F.High, F.Steady |
| 2016-71-0078 | 390 | -0.0004 | -0.0001 | -0.0008 | FHigh, steady |

The differences are small and similar to those seen in September. The salinity difference is slightly higher than those seen during 2016-10 but smaller than during 2016-05 and 2016-07. Salinity differences are close to the average of -0.0008psu deduced from the comparison with bottles.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. Problems found were:

- File 2016-71-0008 has a time that matches the one in the log for event #7. In the headers the NMEA time is 15:19:23 and the System UTC is 15:52. The latter matches the log. It is possible that the file was started for event #7 which was a net cast at the same site, but archiving was started for event #8. The time was changed in the header to match the System UTC. The NMEA positions were the only ones available in the header. The log shows only a small change in position between the start of the two casts, so the error should be small. The time was changed in the IOS and CLN files and a comment added to explain the change.

- Stations names were wrong for events 42 and 43. They were fixed.

The errors were fixed in the IOS and CLN files and an explanatory comment was added to the header of event #8. Corrections were also made to bottle files for events #8.

A header check was run. The only problem noted is the very high pH values already noted. No off-scale fluorescence values were found.

The cruise track was plotted and added to the end of this report.

The Surface Check gave the average surface pressure as 2.3db which is within the normal range for the Vector. The lowest pressure in the CTD files was about 0.4db and pumps were on then. The CTD data show that the CTD was definitely in water at that level.

The altimeter and bottom depth headers were exported from the SAMAVG and CLN files to spreadsheets. The bottom depths were checked against the log book. There were some minor discrepancies; some log entries had been changed so they may be a better reflection of the depth when the CTD was at the bottom.

The following calculation was made to see if the bottom depths and altimetry readings are good:

Check=maximum depth sampled (Max press*0.99) - water depth + altimeter reading

We don’t expect a 0 reading since the altimetry reading is an average over 2db and the bottom depth is subject to changes during the cast and the calculation of maximum depth is rough. Where the Check value was > 2 checks were made by seeing if changing the depth to that in the log improved the result. If not, then plots were examined to see if the entry was reasonable. Changes were made to the CLN, MRG and SAMAVG files where problems were found. The altimetry header was removed from cast #1 as the CTD clearly did not get within 15m of the bottom.  In some cases changing the bottom depth to match the log improved the result so those were changed in the header and for cast #47 the altimetry header was adjusted based on a plot.

Cast #9 will not be processed further as it only went to 2m to fire 2 bottles.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and as usual, there is variability with the up and downcast traces sometimes closer than temperature and sometimes further apart. This is likely due to varying vertical gradients and descent/ascent speeds. Overall the choice made earlier looks appropriate, so no further alignment will be applied.

Conductivity

For 2 recent cruises using the same equipment the best results varied from cast to cast and feature to feature, but overall -0.4 records for the primary and -0.6 records for the secondary looked best for 2016-010. Tests were run on 5 casts and the best results were with -0.4 records for the primary and -0.5 for the secondary conductivity, though other choices looked good as well.

SHIFT was run twice on all casts using -0.4 records for the primary and -0.5 for the secondary conductivity.

pH

The data from the pH sensor look bad, with all values >9.7. This sensor malfunctioned at the end of 2016-45 and throughout 2016-10.  The data will not be archived.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were warnings for 11 casts but they pertained to upcast parts of the casts so are of no concern.

A check was made of casts #1 and 5 for which archiving started at the end of the soak, so the files contained the rise from 10db to the surface. DELETE worked well with no data from the upcast section selected.

##### Other Comparisons

Previous experience with these sensors –

The temperature, conductivity and dissolved oxygen sensors were used during 2 cruises in April 2016 and 2 during June 2016. Two cruisse had delayed salinity bottle analysis and flushing of Niskin bottles was likely poor during all the cruises. While the CTD salinity looked low compared to bottles, it is likely because the bottle salinity was too high. During 2016-07 near the mouth of Juan de Fuca Strait where flushing was likely good, the primary salinity was found to be low by 0.0003psu and the secondary salinity low by 0.0013psu, so both calibrations appear to be ok. The 2016-10 comparison was poor but suggested that both salinity channels were ±0.002psu. Dissolved oxygen comparisons also showed signs of poor flushing. Recalibration was based on results from an area where more ship motion likely improved flushing.  No recalibration was found necessary for pressure.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed.  Temperature was high between 60 and 140db for casts #8 and 10 at the mouth of Juan de Fuca Strait. As has been seen during several other 2016 cruises there are some temperatures above the climatology maximum around 50 to 60db in the northern part of the Strait of Georgia. No other temperature data was outside the climatology. Salinity was low in many areas, most often between 20 and 80m where the profiles suggest much deeper mixing than usual, especially in the northern parts of the Strait of Georgia. For the two casts at the mouth of Juan de Fuca the salinity was low in the top 120db and at the bottom for cast #8 and at all depths except between 40db and 50db for cast #10. Those 2 profiles do not suggest deeper mixing. Saanich Inlet salinity was low from 50m down and near the bottom of a few casts in the southern Strait of Georgia. Salinity was low between about 20 and 80db near Texada Island with profiles showing deeper mixing than indicated by the climatology.  Given the unusual weather and evidence of deep mixing these excursions do not suggest problems with sensor calibration. The climatology is severe for near-shore and somewhat dated.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

There is little to distinguish between the primary and secondary T and S channels. Both have some spikes though the secondary looks slightly better. The comparison with bottles does not make it clear which salinity is more accurate. The secondary was slightly further from bottles but a small effect from incomplete flushing might reverse that observation. The secondary has been archived for the 5 earlier 2016 cruises that used these sensors. The secondary data were selected for editing and archiving.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes. All but 2 files (#5 and 18) required some editing, but mostly it was removal of records from near the top and bottom of casts and light cleaning of salinity.

##### Initial Recalibration

- There is no evidence that pressure needs recalibration.

- Salinity:T1:C1 appears to be within ±0.001 of bottles, so does not need recalibration.

- Based on the comparisons discussed in section 5 , CALIBRATE was run using file 2016-71-recal1.ccf to correct the Oxygen:Dissolved:SBE channel in the SAM and MRGCLN2 files using:

- CTD DO Corrected = CTD DO * 1.0284

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly. But they also show that for CTD DO<2mL/L the differences are all positive. For the Indian Arm data poor flushing seems likely due to a very steady descent rate; this was quite a small effect because the gradient was low. For one cast in Juan de Fuca Strait with values close to 2mL/L, it appears that the CTD DO value was still falling as the CTD started to move upwards. Whether this is due to slow sensor response or conditions changing during the upcast is not known but in either case it is not a surprise that these data fall a little outside the general fit. When DO<2mL/L is removed from the fit the CTD DO reads high by an average of 0.0012mL/L.

See file 2016-71-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed based on residuals and values <2mL/L were excluded, the CTD DO was higher than the bottles by an average of ~0.03mL/L and standard deviation of 0.025mL/L.

No further recalibration of DO is justified.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

The CTD files from rosette casts were clipped to 50m; sigma-T was derived and the data were exported to a single file, 2016-71-SOG.csv.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some small unstable features, but those are from sites where active mixing is expected. The only other problem is the bad pH data set.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, PAR and PAR:Reference data are nominal and unedited

except that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.3 mL/L from 0 to 150db

±0.08 mL/L from 150db to 300db

±0.03 mL/L below 300db

Channel pH:SBE was removed because the sensor malfunctioned throughout the cruise.

For details on the processing see document: 2016-71_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and a few problems were found:

- The speed check showed a high speed between cast #6 and 8. The header for event #8 had been changed to match the system upload time because the file was started 30 minutes before the cast. However, the system upload time in the file was in PDT, so 7 hours was added to the header time of event #8 and the Header Check was rerun. The time in the bottle file was also corrected.

- The file for event #15 was missing. A header error in the EDT file was fixed and steps from CALIBRATE onwards were run on that file.

Header Check was rerun and no further errors were found.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values were low ranging from 60% to 96%. As usual values were lowest in Haro Strait and in the southern Gulf Islands region and at a few casts in narrow channels to the north-west. The highest values were in Juan de Fuca Strait and the central part of the Strait of Georgia. The weather in October 2016 was unusual with many storms, which may account for the deep mixing of dissolved oxygen. The values at 50m ranged from 3mL/L to 6mL/L whereas in late September 2016 they ranged from 2.8mL/L to 3.8mL/L. So the low saturation likely does not derive from problems with the DO sensor calibration.

##### Final Bottle Files

CALIBRATE was run on all files and then the MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

A note was added to the header of file #8 to explain that it is a combination of events 8 and 9 and that the pumped channels were removed from the last 2 records in the file because the pumps were not on.

Plots were made of all casts to look for problems and none were found.

Standards check and a header check were run on all files. The chief scientist’s name was wrong, so Head Edit was rerun to fix that. No other errors were found.

The track plot looks ok.

Data from the CHE files were exported to a spreadsheet and compared with rosette sheets. No problems were found.

Plots of each file were examined and no problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

##### Thermosalinograph Data

There were no loop samples, flow meter or intake thermistor. The intake is at about 2m. The only method to check calibration is to compare with the CTD casts. There are many more CTD data available at 3m than 2m, so data from 3m will be used as well as 2m.

a.) Checking calibrations

There were 3 files. The configuration files were identical.

One configuration file was renamed as 2016-71-tsg.xmlcon. No errors were found in the calibration parameters.

b.) Conversion of Files

The hex files were converted using file 2016-71-tsg.xmlcon.

They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots were produced. The traces are odd at the beginning of the first file but settle down after a few hours. Temperature and salinity traces look ok in file #2 and also in #3 until the end when it looks like the flow was turned off at about 0400 on the 5th of November.

The track plot looks fine. File #1 is short covering Saanich Inlet to northern Haro Strait. File #2 covers Haro Strait to the mouth of Juan de Fuca and a short part of the return trip. There is a long break between files #2 and #3 during which time the ship moved from the western Juan de Fuca Strait until it got to the southern Strait of Georgia

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.3db of 3db. These were exported to a spreadsheet which was saved as 2016-71-ctd-tsg-comp.xls. There were 52 casts which overlapped with TSG files. The data from 2db were also extracted though there were only 32 casts that overlapped with TSG data and contain data as shallow as 2db. They were added to the spreadsheet with the 2db data.

The 3 TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for lab temperature and salinity and the files were reduced to the times of CTD files. Those data were added to as 2016-71-ctd-tsg-comp.xls.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. There were no differences >0.00095º and the median differences were 0.0001º. When one cast is excluded there were no differences >0.0006 º. For that one cast the position recorded by the CTD cast is found in the TSG record 1.5min earlier, so the difference does not reflect a significant error in either clock.

d.) Comparison of T and S from TSG and CTD data

In the table below T3 and T2 refer to CTD data from 3db and 2db, respectively.

Tmed and SALmed refer to median TSG temperature and salinity, with the median taken over 5 records.

|  | Tmed-Tctd3 | Smed-Sctd3 | Tmed-Tctd2 | Smed-Sctd2 |
| --- | --- | --- | --- | --- |
| Average using all available casts | 0.0512 | -1.3510 | 0.0478 | -1.2657 |
| Median using all available casts | 0.1190 | -0.1103 | 0.1165 | -0.1589 |
| Std Dev using all available casts | 0.1447 | 2.3010 | 0.1328 | 1.9595 |
| Median of casts with low vertical gradients | 0.1312 | -0.0499 | 0.1310 | -0.0501 |
| # of casts in median above | 16 | 17 | 10 | 9 |
| Median of 20/11 casts with lowest TSG std dev | 0.1312 | -0.0497 | 0.1313 | -0.0584 |
| # of casts in median above | 20 | 20 | 11 | 11 |
| Median of well-mixed and low std dev in TSG | 0.1320 | -0.0499 | 0.1316 | -0.0528 |
| # of casts in median above | 10 | 9 | 5 | 5 |

The differences between the CTD and TSG temperatures did not vary greatly once outliers were removed using any of three criteria (low standard deviation in TSG data, low CTD temperature gradient or both of the latter being low.) They all suggest that the lab temperature is higher than the CTD by about 0.13C°. There is no significant difference between the 3m and 2m comparisons.

The differences between the CTD and TSG salinity at 3m are consistent showing the TSG salinity to be lower by about 0.05psu. The 2m comparison shows slightly larger differences but there are fewer data points in the comparison. For the 2m cases with both low standard deviation in the TSG data and low near-surface salinity gradients, the TSG salinity was low by ~0.053psu but only 5 casts met those criteria.

The criteria applied ensure that the effects of the mismatch between the intake level and the CTD sampling level is small so that we can see if there are problems with the TSG. The difference of 0.05psu sets an upper limit to the errors due to small bubbles in the loop and/or conductivity calibration drift. But even though the gradients are relatively low, some of this error is may be due to water being drawn into the loop from above 2m.

When we look at the casts with larger salinity gradients, the TSG tends to read much lower than the CTD. The areas with large salinity gradients were in Indian Arm, Saanich Inlet and in the Strait of Georgia in the area most influenced by the Fraser River. While it is possible that there is something about being near rivers that increases bubbles, it seems much more likely that the loop is drawing water from above 2m. There is too little evidence to support recalibrating the TSG salinity. It may well reflect actual conditions closer to the surface than measured by CTD data. This would also explain some of the temperature difference, so no attempt will be made to calculate a proxy for intake temperature.

(See 2016-71-ctd-tsg-comp.xls.)

Calibration History

The TSG temperature and conductivity were recalibrated in December 2015 and have been used for 2 other cruises: 2016-07 and 2016-10 in the same region. For the earlier cruise in June the lab temperature was higher than the CTD temperature by between 0.1C° and 0.2C°.  Salinity was found to be low by 0.62 using all casts, but the standard deviation was very high at 3.0psu. The median difference was 0.03psu when only the 10 casts with the lowest standard deviation in the TSG salinity were included. For 2016-10 in September the CTD temperature was higher than the CTD temperature by a median of 0.127C°. The TSG salinity was lower than the CTD salinity by a median value of 0.93psu using only the 20 casts with the lowest standard deviation in the TSG salinity. The differences were slightly lower when the 2db data were used but there were fewer data available.

Conclusions

1. The TSG clock worked well.

2. The TSG lab temperature was higher than the CTD temperature by about 0.13Cº. We have little experience with the Vector loop system and no knowledge of the flow rate during the cruise. From the Tully we might expect heating of ~0.16Cº given similar intake temperatures, but the Vector loop is shorter so that would reduce the amount of heating. During the previous use in the same region the difference was 0.12Cº, but the intake temperatures were higher which would lead to less heating since the intake waters would be closer to the ambient temperature of the ship. Some of the difference between TSG and CTD temperature could be due to water in the loop coming from higher than 2m. Calibration drift is not likely to be a significant factor.

3. The TSG Salinity is lower than that from the 3m CTD by about 0.05psu. Calibration drift is unlikely to be that large. It is more likely that the difference is due to small bubbles in the loop and or that water is being drawn from above 2m. Bubbles are not obvious in the TSG traces, but small ones could be present. The near-surface gradients are very high for some casts and the differences between CTD and TSG are very high for those casts with high near-surface gradients.

4. Overall, the quality of the data was good, but there are a few bad sections at the beginning of file #1 and the end of file #3. File #1 looks very odd with no variability at first followed by a spike during which temperature drops suddenly. Then for about 1 hour when the ship does not appear to be moving the salinity is extremely noisy. This is assumed to be during the test cast in Saanich Inlet. There are then some believable records that relate well to the file that followed. The 3rd file had a section at the end where the temperature gradually rises and salinity and position do not vary; the pump had obviously been turned off so some data need to be removed.

5. There is insufficient evidence to justify making an estimate of intake temperature or to recalibrate salinity.

f.) Editing

CTDEDIT was used to remove records from both files:

2016-71-0001 – Temperature and salinity data were removed from the first 325 records before flow was well established.

2016-71-0002 – Records from the end of the file (records 10893-11296) were removed since the ship did not move and a steady rise in temperature showed that flow in the loop had stopped.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Conductivity:Primary, Flag and Position:New channels.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats. The Temperature:Primary was renamed as Temperature:Lab to make it clear that it is not the intake temperature.

Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series and all look fine.

The cruise plot was added to the end of this report.

Particulars (mostly notes from log)

1. Archiving started at 10m.

4. pH cap not removed.

5. Archiving started as CTD rose from soak level.

8. Bottles 14 and 15 fired in air during swell.

9. Bottles 14 and 15 fired to replace those missed during cast 8. JOIN to #8. Do not need CTD file

12. pH cap may have been on.

13. Forgot to fire bottle #15 – no particle sample.

19. Bottle fired by mistake – no sampling. No CHE file needed.

39. May have touched bottom; altimeter did not read bottom on downcast, but does read 0 for a few records at bottom.

40. Very windy during sampling.

42. Altimeter did not kick in until bottom.

46. Sample #174 not used. Raining heavily during sampling.

51. Approximate time and location in log. Raining heavily during sampling and very windy.

53. Raining heavily during sampling.

55. Bottle fired for nutrient reference, no sample number, no CHE file needed.

60. Chains on.

61. High winds ~40knots – mixed layer 75m deep.

62. Windy ~40knots

64. Chains off. Pump not on until return of CTD to surface after soak.

70. pH cap left on.

74. pH cap may have been left on.

79. Hit bottom.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 18Dec2014 | Factory |  |  |
| Conductivity | 1766 | 20Aug2014 | Factory |  |  |
| Secondary Temp. | 4054 | 20Aug2014 | Factory |  |  |
| Secondary Cond. | 3531 | 6Feb2013 | Factory |  |  |
| Transmissometer | 1396DR | 2Feb2016 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 19Dec2015 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| Surface PAR | 20518 | 21Mar2016 |  |  |  |
| pH | 0692 | 12Jan2016 | Factory |  |  |
| SeaPoint Fluor. | 3865 | n/a |  |  |  |
| Pressure Sensor | 0550 | 17Dec2014 | Factory |  |  |
| Altimeter | 1253 | n/a |  |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3363       Cruise ID#:	2016-71

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3363 | 17Dec15 | Factory |  |  |
| Conductivity | 3363 | 17Dec15 | Factory |  |  |