## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: CFMETR Nanoose

Agency: Canadian Forces

Project: Canadian Forces Maritime Experimental Test Range

Area: Nanoose

Date: 6 January 2016 – 16 December 2016

Platform: CFAV Stikine (and others)

Processed by: Germaine Gatien

Date of Processing: 1 March 2017 – 3 March 2017

Number of original Hex files: 80 	Number of casts processed: 66

INSTRUMENT SUMMARY

Six Sea Bird Model SBE 19 SEACAT CTDs were used – serial numbers #1449, 1450, 1451, 1511, 1512 and 1513. Each CTD had an altimeter attached.

SUMMARY OF QUALITY AND CONCERNS

Six different CTDs were used. The calibration information could not be confirmed. The CTDs are believed to be serviced annually but only 2 had calibration coefficients updated in 2016. The others have calibration coefficients from 2014.

Four casts have positions that are suspicious being near shore or on land. If the positions are correct they are outside the area of interest and if they are wrong we have no way to correct them, so those casts will not be archived.

To assist in analysis, event numbers are helpful and the software in use will only allow 4 digits which suits most cruises. However a 7-character system is used for this project. A 4-character format was derived from those IDs based on the month and day of each cast. This means the same event # will appear if there is more than one cast on one day as frequently occurred.

The soak period for the CTD varied greatly from about 25s to 6 minutes, but was usually about 1.25minutes. It is recommended that the SBE19+ be soaked for at least 2 minutes to improve salinity quality, though in the absence of large differences between air and water temperatures it is likely that 1 minute is sufficient.

There was no salinity calibration sampling. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors are expected to be as high as 0.05units in areas of high gradients. Away from high gradients errors in salinity are expected to be on the order of 0.005.

The times given in the SeaBird headers do not include a time zone, but the times strongly suggest that UTC was not used. All times were converted to UTC by adding 8 hours which is suitable if the original was in PST. If PDT was used for some casts, then those will have an error of 1 hour.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

80 hex files, 78 con configuration files and 77 hdr files were delivered.

The file names are non-standard and contain the opening (U) followed by the date in MMDD format and 2 digits to indicate the sequential casts on a single day. There are no event numbers as this is a time-series rather than a cruise.

- Conversion of Raw Data

- The file names were changed to the usual format for Nanoose data with a leading “n16” instead of (U) so that the format is then nYYMMDD##.

- The raw data were converted by matching the configuration file to the input file name. This produced 78 CNV files.

- In one case, n16031101 there was a hex file with associated hdr file but no con file. From the header file it was determined that the CTD was #1449, so the con file from another cast using that sensor was used to convert the hex file.

- In another case, n16110101 there was a hex file with no associated hdr or con files. From the hex file header it was determined that the CTD was #1451, so the con file from another cast using that sensor was used to convert the hex file.

- A spreadsheet “NanooseCtdStations2016.xls” was provided with positions for the casts but there were only 67 stations. For 10 of the 13 casts missing positions the converted files have data that does not look useful either because there are only surface data or they are shallow with large spikes and were obviously aborted and a second cast run that was successful.

- The 3 casts with data that look good were followed by second casts for which positions are available, so those can be used since there was not much time between the casts so the ship couldn’t have moved far:

- N16040502 –Use position from file n16040503.

- N16040701 –Use position from file n16040702.

- N16090801 –Use position from file n16090802 (Both are shallow casts)

- The spreadsheet provided was adjusted by adding the 3 missing positions and saved as 2016-99-merge-hdr.csv. Some editing was required:

- The format of latitude and longitude in the spreadsheet needed to be changed to DD MM.MM N and DDD MM.MM W. Concatenate was used to do this.

- A column LOC:Station was added with every row having “Nanoose” entered.

- A column of event numbers was adding to the CSV file using MMDD part of the file name as event number.

- The soak period ranged from 25s to 6 minutes, with the most common wait was for 1.25minutes. A 2-minute soak is generally advised for the best salinity data, but in most cases the conductivity seems to have reached equilibrium after 1 minute.

- Temperature and conductivity generally look ok.

- The altimetry looks useful near the bottom but the CTD did not always get close enough for the altimetry to get entered into the IOS Headers.

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature.  Tests were run on a few casts to ensure this was the best choice.

The pressure was not filtered since it is quite smooth already and tests showed that filtering made things worse in sections where pressure was constant.

- ALIGNCTD

An initial run with no alignment applied to the temperature channel produced poor results with instabilities in T-S plots. Tests were run on a few casts with different CTDs and the best results were found when temperature was advance by 0.73s relative to pressure. That is the value used for 2015 data and close to the setting used in 2012 (0.7s). There are small instabilities left in many casts. Alignment can be fine-tuned later using SHIFT if further adjustment is found useful.

- DERIVE

Program DERIVE was run to calculate salinity

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers.

The program MERGE CSV FILES TO HEADERS was then used to add positions, station name, event numbers and project name to the headers.

8. Checking Headers

Track Plot was used to ensure positions look normal. There were a few problems that were found to be due to errors in fixing the formats, but there remain 4 outliers, one on land, with no obvious explanation. These were referred back to the chief scientist who decided they should be removed from the data set since there is no way to correct the positions if they are wrong and the data fall outside the area of interest if they are right.

The times are said to be UTC but are clearly in local time, so ADD TIME CHANNEL was used to add 8 hours. It is not known if PST was used throughout the year, so the times may be off by one hour if and when PDT was used at sea.

Header Check was run and there was no evidence of problems. There were some negative pressures but all at the beginning of casts. There were 6 different CTDs in use (S/N 1449, 1450, 1451, 1511, 1512, 1513).

Surface check was run and values ranged from -0.7 to +3db. The average surface pressure was 1.1db.The 4 negative pressures were all from records with 0 or near-zero conductivity values and those came from 4 different CTDs. The 7 casts with the highest surface values came from 4 different CTDs as well, so there is no evidence of significant differences among the pressure sensors.

- The altimetry header information from each file was exported to a spreadsheet. Approximately half of the casts did not get within 15m of the bottom, so there is no header entry. Plots for a selection of casts were checked and the altimetry header values look appropriate. A few with no header entry were also checked and the CTD did not get with 15m of bottom.

9. SHIFT

These data were aligned earlier in the process, but tests were run to see if a further shift might reduce unstable features. The results were not consistent with only slight improvements with varying settings. SHIFT was not run.

10. CLIP

CLIP was used to remove the first 50 scans; by eliminating the early soak period editing will be simpler. Plots were examined to ensure no good data had been removed. One cast was put through CLIP again to remove more scans.

11. DELETE

Then DELETE was run on the clipped files using the following parameters:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

There were warnings for events #405 and 407. There are jumps in pressure. Tests to adjust the filter applied earlier did not improve this, nor did using the pressure filter within DELETE. While there are jumps it appears the DELETE worked properly.

12. CTDEDIT

The DEL files were copied to EDT.

DEL files were opened in CTDEDIT. 13 of the casts were edited, mostly to smooth cases where unstable features appeared to be due to minor misalignments of T and C in the presence of temperature gradients. A few records were removed from the top and bottom of some casts.

The edited files were copied to EDT.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure      	Averaging interval = 1.000

Minimum bin value = .000	Average value will be used.

Interpolated values are NOT used for empty bins

Profile plots were examined on-screen and no problems were found.

T-S plots were examined and look ok with only a few very slightly unstable features.

14. Intercomparisons

Historic ranges – The only excursions from the local climatology were:

- Temperatures were often a little high near the surface down to about 25m in May and between 140 and 180m for one cast in March.

- Salinity was slightly low around 40m for one March cast and near the surface down to about 30 to 65db in November and December.

These do not look like evidence of calibration problems.

Repeat casts – Casts were all close in space, and on a few occasions there were 2 or more casts on one day but not necessarily very close together. Two casts that were very close in space and about 3 hours apart were examined. For most of the profiles there were fairly large differences but from 350m to the bottom temperatures were within 0.005Cº of each other along sigma-T lines and salinity values were within 0.002psu. This is good agreement.

15. CALIBRATION

No recalibration was applied.

16. REMOVE and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity, Descent_Rate Altimetry and Flag.

REORDER was used to place Depth after Pressure.

The 4 casts with suspicious positions mentioned in section 8 were not processed further.

The HEADEDIT routine was used to fix formats and to add the following comment to the headers:

SeaCat salinity is prone to error due to a mismatch of conductivity and

temperature response times, especially when the descent rate is non-uniform.

Such errors are likely to be larger than calibration errors. SeaCat salinity

errors are expected to be as high as 0.05 units in areas of high gradients,

but on the order of 0.005 where gradients are low and descent rate steady.

There were no salinity calibration data available for this cruise.

The CTDs are serviced annually.

The standards check routine was run and no problems were found.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing and header check were produced. No errors were found.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	Nanoose Test Range

Dates: 	Start:	6 January 2016	    End:	16 December 2016

Location: Nanoose Test Range 	Vessel:    Stikine (and others)

NOTE: Six SEACATs were used for this cruise.

The calibration information used in conversion can be found in the headers of the CTD files and a summary of which CTDs were used for which casts can be found in document 2016-99-ctd-final.doc which is in the archive folder Cruise_Data_Processing\2016-99\Archive\DOC.