## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2017-03

Agency: WVL, Ecosystem Sciences Division

Location: Baynes Sound

Project: Coastal Ecosystem Research Project

Party Chief: Sutherland T.

Platform: Vector

Date: April 14, 2017 – April 19, 2017

Processed by: Germaine Gatien

Date of Processing: 7 June 2017 – June 13, 2017

Number of original HEX files: 40

Number of CTD files: 40	Number of bottle files: 30

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#11185DR), a SBE 43 DO sensor (#1119), a SeaPoint Fluorometer (#3865) with a 3X cable, a Biospherical QSP-400 PAR sensor (#4565), Surface PAR (#20518) a pH sensor (#0691) and an altimeter.

The data acquisition program was Seasave v7.26.2.13.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log book was generally in good order with notes about errors in some station names. This is helpful because it is better to correct them early in the processing. But there was no list of science crew which can delay processing if problems are found and the chief scientist is not available to answer questions.

There were some problems with the sampling log sheets. For some casts there were no entries of what sampling was done. This can be a problem; samples sometimes go astray in analysis or processing, but we will have no way of knowing that we need to look for them. There were two cases of sample numbers being used twice. We cannot use alphabetic names like 1a and 1b, so the numbers were changed using a leading 9 for one of each pair. One log sheet had the wrong event number.

The casts were shallow and in a region where Niskin bottles are likely to flush poorly and vertical gradients are apt to be high. So comparisons between CTD and bottle data are of limited value for judging calibration. Fortunately, the cruise immediately followed cruise 2017-63 which had the same equipment and included some sampling in Juan de Fuca Strait where more reliable comparisons could be done. So the results of the earlier cruise were used for calibration purposes.

Salinity analysis was done promptly so there should be little error due to evaporation or adsorption. However, poor flushing of Niskin bottles in the presence of high vertical salinity gradients make the sampling unsuitable for assessment of the accuracy of the CTD sensors. Results of cruise 2017-63 which immediately preceded this cruise and used the same equipment indicate that the primary salinity did not require correction. The secondary salinity was recalibrated to match the primary in case it is ever needed, but the primary salinity was chosen for archiving for all casts.

Channel Oxygen:Dissolved:SBE was recalibrated using the results of cruise 2017-63.  During that cruise the downcast Oxygen:Dissolved:SBE data were considered, to be, very roughly:

±0.3 mL/L from 0 to 150db

As some of the error found for that cruise may be due to incomplete flushing of Niskin bottles, this may underestimate the accuracy of the dissolved oxygen sensor during downcasts.

During the two previous cruises that used this transmissometer #1185, the data looked odd. During 2017-01 in offshore waters it was very clear that there were sudden shifts in values that usually persisted through the remainder of a cast. Upcast and downcast data were often further apart than usual. During 2017-63 which immediately preceded this cruise there were some sudden shifts in values and there were significant differences between downcast and upcast profiles for many casts. For the current cruise the casts are all shallow and there is no evidence of a problem. There is a lot of variability in transmissivity at repeat stations, but the chief scientist expected that due to the nature of the sampling and the season.

The transmissivity data were included in the final files, but should be used with caution.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. Notes from the log book indicate a problem with station names in files #46 to 51:

Nutrients, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

There only sensor history for the pressure, conductivity and DO sensors comes from 2017-63 because that is the only other cruise since their factory recalibration in February or March 2017.

The calibration constants were checked for all instruments and no problems were found. One file was saved as 2017-03-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2017-03-ctd.xmlcon to create CNV files. The file names were corrected in the headers of files #46 to 51

A few casts were examined.  All expected channels are present. The primary and secondary temperature and conductivity channels are close during downcasts, but as usual the upcasts differ more due to noise in both channels. Dissolved oxygen, fluorescence, PAR, SPAR, pH and altimetry profiles all look normal.

Transmissivity is hard to judge – it has had problems in recent cruises and will be examined closely later.

The descent rate was high and steady.

##### BOTTLE FILE PREPARATION

ROS files were created using file 2017-03-ctd.xmlcon.

At the time of initial processing no files were found for casts 53 to 57, but they were found later and put through the same steps.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers.

CTDEDIT was used to clean Salinity:T1:C1 to remove 1 large spike in each of casts 28 and 39.

The *.ED1 files were copied to *.BOT.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv.

Sample numbers were added to the file based on the rosette log records. There were two problems with repeated sample numbers.

- There was a sample for cast #22 given as #98 on the log sheet and one from cast 23 given as #98a. For cast #22 there was only a nutrient sample with that number. The samples were all called #98 in the analysts’ spreadsheets for both casts. Since there was only a nutrient sample with that number from cast #22 it was renamed as #998 to distinguish them.

- There was a sample for cast #53 given as #181 on the log sheet and one from cast 54 given as #181a. There was only nutrient sampling for both casts. Both samples were called #181 in the analyst’s spreadsheet so the one from cast #53 was renamed as #998 to distinguish them.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine. Those files were bin averaged on bottle number.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2017-03-bot-hdr.txt which will be updated as needed during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2017-03oxy*.xlsx which includes flags, and comments. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2017-03oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in QF2017-03SAL*.xlsx. There were no duplicates. The analysis was done within 7 to 12 days of collection. The files were simplified and saved as 2017-03sal.csv.

That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet NUTS_QF2017-03nuts.xlsx. The spreadsheet was simplified and saved as 2017-03nuts.csv. The file was converted to individual NUT files.

The SAL, OXY and NUT files were merged with CST files in 4 steps.

After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only, with output files named MRGCLN1.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number as the merge channel.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. There were many discrepancies, mostly cases where no sampling is shown on log sheets but samples exist. There was one case of a DO sample noted on the log sheet but not found by the analyst. Given the many discrepancies, no flag was added as this may have been a deliberate choice.

CLEAN was run on the MRG files to add 0 flags to empty flag channels and to update header limits.

##### Compare

Salinity

Compare was run with pressure as reference channel. The bottle from cast #57 is not included because the CTD data were not available when the comparison was done.

There was a large scatter in the comparison. The two most extreme differences were removed based on residuals. The primary CTD salinity was found to be higher than the bottle salinity by an average of 0.050 with a standard deviation of 0.029. The secondary CTD salinity was found to be higher than the bottle salinity by an average of 0.057 with a standard deviation of 0.030. This implies a difference between the two CTD salinity channels is roughly 0.007. That is slightly higher than the differences found between the channels as reported in section 9, but the standard deviations during bottles stops are quite high. All the samples came from bottom bottles so flushing errors lead to bottle values being lower than expected. The differences are much higher than usual, but vertical gradients were high and flushing likely was incomplete.

The outliers do not look like they are due to problems in sampling or analysis so no flags should be added. They are likely due to incomplete flushing of Niskin bottles in the presence of large vertical salinity gradients. This is to be expected in calm waters such as these.

For full details for the COMPARE run see file 2017-03-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel. The bottle from cast #57 is not included because the CTD data were not available when the comparison was done. Because the casts are shallow, there is a larger proportion of values from the bottom of casts than usually seen in such comparisons. This is an area where we expect poor flushing, so the bottom bottles are likely to have DO values that are higher than ambient waters, while bottles fired on the way up are likely to have values that are lower than ambient waters. This type of DO sensor almost always reads too low, but if the bottles are also low the fit will suggest a lower correction than usual. The bottom bottles will have the opposite effect.

There is a lot of scatter in the fit of (CTD DO – Bottle DO) versus CTD DO. When cases of the CTD reading higher than titrated samples were excluded and some other outliers based on residuals, the fit found was:

CTD DO Corrected = CTD DO * 1.0352    R² = 0.33 (1)

The fit was forced through the origin because there were no DO values <4mL/L. This correction is similar to that found for 2017-63 using all casts:

CTD DO Corrected = CTD DO * 1.0342 + 0.0712    R² = 0.80 (2)

Given the scatter this is a little surprising until it is noted that most of the bottom bottles were excluded as outliers. The fit using bottom bottles and excluding one outlier was:

CTD DO Corrected = CTD DO * 1.0742    R² = 0.06 (3)

This is shown only for information as the range of DO values in the fit is very low, so we can’t expect a good estimate. It does, however, show that the bottom bottles are quite different from the rest in their effect on the fits.

The fit for cruise 2017-63 using only casts from near the mouth of Juan de Fuca Strait and excluding some outliers was:

CTD DO Corrected = CTD DO * 1.0422 + 0.0358    R² = 0.91 (4)

This fit was the one chosen to recalibrate all data from 2017-63 because the Juan de Fuca casts appeared to have better flushing of Niskin bottles because the CTD had more vertical movement during stops. For a CTD sensor reading of about 6mL/L the corrected values using fit(1) was 6.21 and for fit (4) it was  6.29mL/L. Fit(4) is likely to provide the best results, though it may still underestimate the correction.

None of the outliers are clearly due to sampling or analysis results. They are in areas with large DO vertical gradients where even small mismatches in levels of sampling or small flushing inefficiencies can explain large differences.

For more detail see document 2017-03-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further problems were found.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

The results of tests during 2017-63 were used to determine the best setting to advance the DO signal.

ALIGNCTD was run on all casts using +1.5s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. The default setting of was selected and it does improve the data for both conductivity channels for these data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on the deepest cast to calculate differences between sensor pairs. The differences found during the previous cruise are listed as well.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2017-63-0074 | 400 | -0.0001 | +0.00037 | +0.0040 | F.High, Steady |
| 2017-63-0083 | 400 | 0 | +0.00038 | +0.0039 | F.High, Steady |
| 2017-63-0107 | 320 | -0.0006 N | +0.00040 N | +0.0045 N | F.High, Steady |
| 2017-63-0129 | 300 | -0.00045 N | +0.00048 N | +0.0050 N | F.High, F.Steady |
| 2017-03-0027 | 90 | ~0.0001 | ~0.0005 | ~0.0054 | High, Steady |

The differences in conductivity and temperature during 2017-63 were of moderate size, though they have cumulative effects on salinity. The salinity differences were larger than we would expect from newly calibrated sensors. The results for 2017-03-0027 are reasonably close to the earlier results. The salinity difference implied by the salinity comparison with bottles was ~0.007; the standard deviations in the CTD salinity during bottle stops were very high so this does not look like a very reliable result.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. The only errors found were in station names, one was not complete and there was inconsistent formatting in many others, so formats were changed to eliminate spacing as that is the most common format used on other cruises and it was used for many of these casts as well.. So BS23 A because BS23A. this was done for both the CLN and SAMAVG files. After the corrections the MRG and MRGCLN2 files were recreated.

There is no evidence of the time problems that were found during 2017-63.

Cruise tracks with station names and event numbers were plotted and added to the end of this report.

The Surface Check gave the average surface pressure as ~1.7db which is within the normal range for the Vector.

Header Check was run and showed that the lowest pressure recorded was 0.35db with pumps running; the salinity was ~27.47 so clearly the CTD was in water.

The altimeter and bottom depth headers were exported from the SAMAVG and CLN files to spreadsheets. The bottom depths were checked against the log book. There were many minor discrepancies but a few were large enough to need investigation. The altimetry header entries were checked by calculating

Check Value = (Bottom depth – maximum depth of sampling + altimetry header)

The following cases were examined because either the Check Value >2m or bottom depths differed significantly from log entries:

- Casts 5, 46, 47, 50 and 57 had bottom depths that did not agree with the log. In some cases it was an obvious error. In other cases there may have been a change during the cast. For these casts changing the depth to match the log led to an improved Check Value, so they were changed in the headers,

- For casts 5 and 52 plots showed that the altimetry headers were not appropriate, having been contaminated by spikes. An estimate was made based on the difference between the maximum depth sampled and the water depth recorded in the log book.

For some of the bottle files there was no altimetry header due to spikiness; no attempt was made to estimate the distance from the bottom since there may well have been drift between the CTD being at its maximum depth and the bottle being fired. None of the errors found in the CTD files are relevant to the bottle files.

##### Shift RERUN AFTER HEADER FIXES

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. There were few distinctive features to enable a judgment, but the choice made earlier looks reasonable, so no further alignment will be applied.

Conductivity

There are few obvious unstable features in these casts. Tests were run to see if the shifts used for 2017-63 have any effect and there are some very minor improvements and nothing looked worse, so those values were used.

SHIFT was run twice on all casts using -0.6 records for the primary and -0.5 for the secondary conductivity.

pH

Checks were made to see if the results of 2017-63 look reasonable. There are few distinct features to aid assessment, but the shifted values look much better than the unshifted data.

SHIFT was run to advance the pH channel by 50 records for all casts.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

All sensors were used during 2017-63. The pressure looked ok, the primary salinity was very close to bottles in Juan de Fuca Strait where flushing of Niskin bottles was considered better than for other casts. The secondary salinity was higher than bottles by about 0.0045 in Juan de Fuca Strait.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed.  All data were within the climatology except for a few low salinity values near the surface of a few casts. This does not suggest calibration problems.

Repeat Casts – There were repeat casts but they are shallow and taken over a tidal cycle, so not suitable for judging the repeatability of the sensors.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

There primary T and S channels were chosen for editing since that was found best for 2017-63.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes. All files required some editing.

##### Initial Recalibration

- There is no evidence that pressure needs recalibration.

- Salinity:T1:C1 is reading higher than the primary by ~0.004psu, so subtracting -0.004 will bring it into agreement with the primary which is considered more realistic

- Dissolved oxygen will be recalibrated based on the 2017-63 Juan de Fuca comparison:

CTD DO Corrected = CTD DO * 1.0422 + 0.0358

File 2017-03-recal1.ccf was prepared to apply the above corrections to salinity and dissolved oxygen in the MRGCLN2 and SAM files.

COMPARE was run using the recalibrated salinity data to check that the correction was applied correctly; the average differences for primary and secondary salinity are closer but still differ by 0.003. The recalibration was based on 2017-63, so a good match is not expected.

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly. The CTD values are mostly higher than bottle values as expected due to incomplete flushing of Niskin bottles.

See files 2017-03-sal2.xlsx and 2017-03-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is usually made of downcast CTD data to bottle data from the same pressure. This was not appropriate for this cruise since the recalibration was based on another cruise and there was huge scatter in the DO comparison. A second calibration step is generally not found necessary and was not found necessary for 2017-63.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots show no significant instabilities.

At this point the transmissivity was examined. The data from 2 other recent cruises showed sudden shifts in values and the data were considered of doubtful value. For this cruise it is harder to judge individual casts, especially the upcasts because there are usually shifts during stops for bottles. But there were many repeat casts. Plots were made of casts at a single site were made:

- BS2 – There were 8 casts and one stands out as having extremely low transmissivity; temperature and fluorescence do not stand out as unusual.  At 30m cast #8 had transmissivity ~5%. The other 7 casts vary from 30 to 40% at the bottom -30m.

- BS17 – 20% to 50% at bottom -50m

- BS23 – 35% to 50% at bottom 50 m. Fluorescence is very different for the final cast BS23I which is a bit to the north so perhaps is near an aquaculture site. Oxygen and temperature vary little.

Data from the previous year in this region were examined to see if this is typical of the area. During April 2016 transmissivity did not vary much ranging from 40 to 60% near the bottom, mostly around 50%.  So the 2017 values are often lower than seen the year before and the shifts noted during 2017-01 appeared to be towards lower values than expected. So these data do not look reliable. However, the chief scientist reports:

Given the delay in spring conditions this year, we would anticipate differences in freshwater runoff and phytoplankton blooms. In addition, last spring was very different to previous years as we were in t-shirts on the back deck of the Vector – it was warmer than our June cruise.

The repeat stations that were collected over a tidal cycle were positioned at each entrance where there is significant freshwater runoff (north entrance) and water exchange (south entrance). It was anticipated to see some change over the tidal cycle.

Once we process the water filters we could regress the SPM with the turbidity of the bottle samples and see if they match.

During the earlier cruises there were no observations of transmissivity shifts in the top 40m of the downcasts and few above 100m, so these shallow casts may well be ok. In light of the nature of the experiment, the channel will not be removed, but a warning will be entered in the headers. This decision can be revisited if further research indicates that the transmissivity data are of poor quality.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence:URU:Seapoint, pH:SBE, PAR and PAR:Reference

data are nominal and unedited except that some records were removed

in editing temperature and salinity.

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. There was no extracted chlorophyll sampling during

this cruise.

Dissolved oxygen was calibrated based on the results of cruise 2017-63 when the

same equipment was used.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large.

During cruise 2017-63 when the same equipment was used, the Oxygen:Dissolved:SBE

data were considered, very roughly, to be ±0.3 mL/L from 0 to 150db

WARNING: The transmissometer used for this cruise malfunctioned during two previous

previous uses. There is no clear evidence of problems during this cruise but

the data should be used with caution. There is a lot of variability in the

transmissivity during repeat casts, but this was expected due to the

nature of the sampling.

For details on the processing see document: 2017-03_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values ranged from 100% to 160%. Given the time of year and location the variability is to be expected.

##### Final Bottle Files

CALIBRATE was run on all files and then the MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Plots were made of all casts to look for problems and none were found.

Standards check and a header check were run on all files. No errors were found.

The track plot looks ok.

Data from the CHE files were exported to a spreadsheet and compared with rosette sheets. No problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars

27. Deep basin CTD – 94m

46-51. Station names need fixing – just in CNV files, no bottles for these casts

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 22Feb2017 | Factory |  |  |
| Conductivity | 1766 | 30Mar2017 | Factory |  |  |
| Secondary Temp. | 4484 | 22Feb2017 | Factory |  |  |
| Secondary Cond. | 3531 | 30Mar2017 | Factory |  |  |
| Transmissometer | 1185DR | 27Apr2016 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 19Dec2015 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| Surface PAR | 20518 | 21Mar2016 |  |  |  |
| pH | 0691 | 15Jan2015 | Factory |  |  |
| SeaPoint Fluor. | 3865 | n/a |  |  |  |
| Pressure Sensor | 0550 | 24Feb2017 | Factory |  |  |
| Altimeter | ? |  |  |  |  |