## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 20 March 2025 | Updated channel names & formats in TOB files. G.G. |
| ### 26 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| 01 Sept 2020 | Added HPLC data. S.H. |
| 8Oct2020 | Corrected HPLC data for event 174. G.G. |

## PROCESSING NOTES

Cruise: 2017-05

Agency: OSD

Location: WCVI

Project: La Perouse

Party Chief: Yelland D.

Platform: John P. Tully

Date: 23 May 2017 – 4 June 2017

Processed by: Germaine Gatien

Date of Processing: 14 August, 2017 – 29 August

Number of original HEX files: 96	Number of CTD files: 95 (2 joined)

Number of bottle files: 69		Number of bottle casts processed: 61

Number of original TSG files: 4		Number of processed TSG files:	 4

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1201DR), a SBE 43 DO sensor (#1119) on the primary pump, a SeaPoint Fluorometer (#3685) on the secondary pump, a Biospherical QSP-400 PAR sensor (#70613), and an altimeter (#4565).

A thermosalinograph (Seacat 21 S/N 2488) was mounted with a Wet Labs WETstar fluorometer (S/N ws-3s 953p), remote temperature sensor and a flow meter.

Seasave version 7.26.1.8 was used for acquisition.

The data logging computer was the Tully CTD Laptop.

The deck unit was a Seabird model 11+, serial number 0425.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial #58879.

The oxygen kit was Scripps kit #2.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log, rosette log sheets and analysis logs were generally in good order with detailed comments on problems, but there was no list of personnel.

The sample numbers from event #174 were renamed because the numbers were also used for event #177. Samples 9414, 9415 and 9416 will be used for event #174 and samples 414, 415, 416 for event #177.

Event #67 is noted in the log but no data file was found. There were BL and XMLCON files.

Events #76 and #172 contained only upcast data. There were computer problems noted at the bottom of the casts and it looks like the computer started new files with the same names, thus erasing previously recorded data. Files were created using upcast data, but the quality are considered lower due to the CTD rising through wakes created by the cable and rosette.

The transmissivity data were bad below 450m and continued to be bad for most of the upcasts. All values below 400m have been removed. The Transmissivity channel was removed from all the bottle files. The channel was left in the CTD files for events #76 and #172 even though they contain upcast data since the values above 400m do look like the downcast values seen in nearby casts.

The secondary temperature and salinity were chosen for archiving for most casts since there was more noise in the primary salinity. The primary channels were selected for 2 casts that had poor secondary salinity. The secondary salinity was recalibrated by subtracting 0.0054psu based on a comparison with bottle data, while the primary salinity required no adjustment. The bottle comparisons had low scatter and salinity analysis precision was excellent.

As usual the CTD fluorescence values are higher than the extracted chlorophyll for CHL<2ug/L and then fall relative to CHL until they are about 50% of CHL at high CHL values.

There were many dissolved oxygen samples including some deep casts, so a good comparison with the DO sensor on the CTD was possible; results were in line with those from other cruises using this sensor.

Downcast Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5 mL/L from 0 to 75db

±0.3 mL/L from 75db to 200db

±0.08 mL/L from 200db to 400db

±0.04 mL/L below 400db

The intake thermistor worked well throughout the cruise. The salinity trace was full of large one-sided spikes to lower values that are likely due to bubbles. Values are given with only 2 decimal places due to the accuracy being lower than when such spiking is not seen. While salinity is known to be low, on average, comparisons only show what was happening while the ship was stopped and the errors vary greatly, so recalibration is inappropriate. The TSG fluorometer readings were fairly close to the CTD fluorometer, but varied from being higher for CTD fluorescence values <1ug/L and lower above that. There was no loop sampling. TSG fluorescence values were lower than extracted chlorophyll samples from rosettes, though the latter mostly come from a few metres below the TSG intake.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

A correction was made to the station name for event #135 based on the log notes.

No rosette files are needed for events #1, 139 and 149.

Nutrients, extracted chlorophyll, dissolved oxygen, NH4 and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors on the CTD were checked.

All were used during 2 previous cruises since they were last calibrated.

The configuration file used for the first 2 casts had some errors. There were 2 PAR sensors listed but only one was in use. The parameters for the PAR sensor that was in use had errors in both the calibration constant and offset. The date of the transmissometer calibration was wrong. Those errors were discovered when test casts were sent to IOS to be checked. The parameters were corrected for event #4.

Configuration file 2017-05-ctd1.xmlcon is appropriate for events #1 and #2.

Configuration file 2017-05-ctd2.xmlcon is appropriate for events #4 to #182.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2017-05-ctd1.xmlcon or 2017-05-ctd2.xmlcon to create CNV files. The hysteresis and Tau functions were selected. There are some deep casts.

A few casts were examined.  All expected channels are present. The primary and secondary temperature and conductivity channels are close during downcasts, but as usual the upcasts differ more due to noise in both channels. There are many spikes, some are clearly due to shed wake corruption, but others are not.

The transmissivity channel has some clear problems. The traces look “steppy” as though the data is under-sampling. This is a new transmissometer and it has been determined that the factory deliberately set it to sample at less than 24Hz. The second problem is seen only in casts deeper than 400m and is more serious since the instrument seems to stop working properly when it gets to about 450m of the downcast. At about that depth it starts producing a smooth trace with values steadily declining during the downcast. On the upcast values are steady or decline slightly for a few hundred metres, then start steadily climbing until they are close to downcast values at the surface. It looks like downcast values to 400m may be reliable, but checks will be made later to see if the values are reasonable.

Dissolved oxygen, fluorescence, PAR and altimetry profiles all look normal.

The descent rate was generally kept high but was often extremely noisy leading to frequent corruption of the data by shed wakes.

##### BOTTLE FILE PREPARATION

The ROS files were created using files 2017-05-ctd1.xmlcon and 2017-05-ctd2.xmlcon.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. For casts 48 and 150 a few spiky points were removed from channel Salinity:T0:C0 using CTDEDIT.

The output files were copied to *.bot.

A preliminary header check and no problems were found. CTD fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs. A few adjustments were needed. There were a number of cases where a bottle was fired as a test, to collect water or by accident with no sample number being assigned. Such bottles were removed from the ADDSAMP file for events #24, 28, 29, 38, 51, 79, 85, 103, 174 and 182.

Sample #s 414, 415 and 416 were used for both event 174 and 177. The samples from event 174 were renamed 9414, 9415 and 9416.

Sort was used to ensure that the ADDSAMP file was in sample number order.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

The SAM files were bin-averaged on bottle # and called SAMAVG.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2017-05-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2017-05chl*.xls. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2017-05chl.csv. The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2017-05oxy.xls which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and was then saved as 2017-05oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in 2017-05SAL.xls. The analysis was done between 16 and 35 days of collection and there is a precision study. The files were simplified and saved as 2017-05sal.csv. File 2017-05sal.csv was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2017-05_nutrients*.xlsx. This includes a precision study. The file was simplified and saved as 2017-05-nuts.csv. The file was converted to individual NUT files.

NH4

NH4 data were obtained in file 2017-05_NH4.xls which included a report on precision. The file was simplified and saved as 2017-05NH4.csv and converted to individual NH4 files.

The SAL, CHL, OXY, NUT and NH4 files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. Problems found were:

- Cast #2 – bottles 1-3 had no sampling, no sample number.

- Cast #18 – Niskin 1 had no sampling, no sample number.

- Cast #24 – Niskin 7 had no sampling, no sample number.

- Cast #28 – Niskin 2 had no sampling, no sample number.

- Cast #29 – Niskin 12 had no sampling, no sample number.

- Cast #38 had one bottle fired by accident. It will not be processed further.

- Some casts had no sampling for IOS but did collect samples for others, so bottle files will be created in case they are needed: events #16, 20, 64, 72, 76, 106, 110, 153.

- Cast #43 – out of order firing. CHL and NH4 planned at same depths, but sample numbers don’t agree. Analysts flagged suspect mislabelling of NH4.

- Cast #51 – drop Niskin 8 – no sampling.

- Cast #79 – Niskin 22 had only sampling for incubation, no sample number. Not needed.

- Cast #85 – Niskin 13 had no sampling, no sample number.

- Cast #103 – Niskins 21-24 closed just for water – no sampling.

- Cast #174 – Niskin 1 had no sampling, no sample number.

- Cast #182 – Niskins 14-24 closed but not sampled.

The unneeded lines identified above were removed from the SAMAVG files; MERGE and CLEAN were then rerun.

##### Compare

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

When 1 severe outlier was removed and others were excluded based on residuals the fit found was:

CTD DO Corrected = CTD DO * 1.0448 - 0.0527

There was only 1 severe outlier (more than 1mL/L off the fit). Examination of the CTD profile data shows that there was a serious problem with the CTD dissolved oxygen data between 20m and 30m of cast #65.

The last time this sensor was used was during 2017-63 in the Strait of Georgia and Juan de Fuca Strait, when the fit was:

CTD DO Corrected = CTD DO * 1.0342 + 0.0712

When 5 casts from Juan de Fuca Strait were used and some outliers removed the fit was:

CTD DO Corrected = CTD DO * 1.0422 + 0.0358

The flushing of Niskin bottles were considered to be much better in Juan de Fuca Strait than for other parts of that cruise.

Most of the outliers are from near-surface records where the fits do not usually work well due to subsurface maxima and high gradients.

One deeper sample was an outlier: event #103, sample 321. The difference looks a little out of line, but examination of the profile shows a local reversal in DO, so like the near-surface outliers, this is likely due to a slight mismatch in depths of the CTD and the contents of the Niskin bottle.

Another problem was found in examining plots of titrated DO and CTD DO against CTD salinity. For event #43 it appears that samples 134 and 135 were reversed. There was a problem during the cast and the 30m sample was taken after the 20m and 10m samples, so labels may have been inaccurate. This is complicated by the fact that the DO profile is complex with a local minimum at about 25m, but if those samples are reversed the comparison between CTD and titrated samples looks much better. The analyst agreed with this step and the results were good.

For more details see document 2017-05-dox-comp1.xlsx.

Salinity

Compare was run with pressure as reference channel.

When outliers were excluded based on differences or depths the primary salinity was found to be very close to bottles. From 500db downwards the average difference was +0.00001 with very little pressure dependence, especially in the secondary. While there is a large scatter near the surface, it is 2-sided with no sign of poor flushing of bottles or evaporation of samples. For the secondary salinity there is a similar scatter pattern with a very flat fit below 500db, but the differences are much higher with the salinity reading high by an average of 0.0054psu. The fit against time suggests that the secondary salinity might be increasing very slightly with time, but the scatter in that fit is too high to give much weight to the trend. The primary shows a smaller temporal change while the secondary varies less with pressure.

There were no cases in the primary salinity comparison with differences >0.06psu. All but 3 differences >0.005psu were associated with a high standard deviation in the CTD salinity. The other 3 showed the CTD reading low. They were all from within 30m of the surface in areas where vertical salinity gradients were high. The one with the largest difference was in a protected area with a steady descent rate suggesting quiet conditions conducive to incomplete flushing of Niskin bottles. It is likely that all 3 outliers were due to incomplete flushing in the presence of high vertical gradients.

For full details for the COMPARE run see file 2017-05-sal-comp1.xls.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor.

A fit forced through the origin has a slope of about 0.63, which is close to the results found during 2016-47 and 2016-62. A plot of FL/CHL vs CHL below shows the usual pattern with CTD fluorescence mostly reading higher than the extracted CHL for CHL<2ug/L and then falling relative to CHL until it is about 50% of CHL at high CHL values.

For full details of the comparison see file 2017-05-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on a few casts to see what alignment setting would produce a vertical difference between upcast and downcast dissolved oxygen traces and those from temperature. During the most recent uses of this equipment +1.5s was found best. But for these data a setting of +2.5s looked best overall.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. During other recent cruises, the default setting of (α = 0.0245, β=9.5) looked better for both conductivity channels than others tested. Two casts were checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. The shaded values are from a previous cruise that used these sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2017-63-0074 | 400 | -0.0001 | +0.00037 | +0.0040 | F.High, Steady |
| 2017-63-0083 | 400 | 0 | +0.00038 | +0.0039 | F.High, Steady |
| 2017-63-0107 | 320 | -0.0006 N | +0.00040 N | +0.0045 N | F.High, Steady |
| 2017-63-0129 | 300 | -0.00045 N | +0.00048 N | +0.0050 N | F.High, F.Steady |
| 2017-03-0027 | 90 | ~0.0001 | ~0.0005 | ~0.0054 | High, Steady |
| 2016-05-0038 | 1000 | +0.0001 | +0.0004 | +0.0046 | High, X Noisy |
|  | 1900 | +0.0003 | +0.0004 | +0.0040 |  |
| 2016-05-0101 | 1000 | +0.0006 | +0.0005 | +0.0055 | High, Moderate |
|  | 1900 | +0.0004 | +0.0005 | +0.0052 | “ |
| 2016-05-0139 | 1000 | +0.0002 | +0.0005 | +0.0051 | High, X Noisy |
|  | 1900 | +0.0004 | +0.0005 | +0.0055 | “ |

The temperature differences are fairly low, but the conductivity differences are a little larger than usual. The salinity differences are fairly large but in line with the earlier cruises, though those data are from much higher in the water column.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers, to replace pad values in the pressure channel with interpolated values based on record number and to replace Transmissivity values with pad values where pressure is >400db.

##### Checking Headers

Header Check was run. It shows large pressure spikes in file #132; this was checked to ensure the spikes were not introduced in processing, but they were found in the original converted file. This cast was interrupted by computer problems and the spikes occurred while stopped trying to resolve those issues. The remainder of the cast is in file #133. No other problems were noted.

A cross-reference list was checked against the log book. No errors were found.

The cruise track was plotted and added to the end of this report.

Surface check was run and shows an average surface pressure for the cruise was 2.2db which is slightly low for the Tully. There were only a few small groups of data with pressure <0db at the end of casts. In one cast there is about 1s during which pressures were as low as -0.35 but a few of the conductivity values were off-scale, so the CTD likely came out of water very briefly. During another cast there are a few pressures ~-0.01 with both conductivity channels going off-scale. So the pressure is likely quite accurate at the surface.

The altimeter and water depth readings from the headers of the CLN files were exported to a spreadsheet. A quick calculation (Max Depth Sampled – Water Depth Sampled +Altimeter Reading) was made; where it is >3m, the cause was investigated. First a check was made of water depths against the log and in a few cases adjusting the file depth entry brought the check value into the expected range. For other cases that needed checking, a plot of altimetry near the bottom was examined to see if the altimetry header needs to be removed or adjusted. Corrections were made to the water depth in CLN files for events #137, 138, 147 and to the altimetry header for event #108.

Similar data were exported from the SAMAVG files. Checks were made to ensure that for casts with only near-surface sampling there were no faulty altimetry headers due to misinterpreted spikes. The altimetry header was removed from casts #10, 16, 34, 106, and 180. The bottom depth was corrected for cast #138 in the SAM files and the altimetry header was adjusted for cast #108. The SAM files were averaged again. The merge of SAMAVG with MRGCLN1s was repeated followed by CLEAN.

The CLN files from the files #132 and 133 (split cast from station SS4) were renamed as 2017-05-0132.clna and 2017-05-0132.clnb. They were joined to create a single complete file which was named 2017-05-0132.CLN.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 4 casts to see what shift in conductivity does the best job of removing noise in the salinity channels. There were many large unstable features in T-S space for the primary channels, and no setting appeared to make much difference to those. A shift of -0.6 records did improve areas with small-scale unstable features. For the secondary a setting of -0.6 records worked best as well. Again there were some large unstable features that were not improved.

SHIFT was run on all casts using -0.6 records for the primary conductivity.

SHIFT was run on all casts using -0.6 records for the secondary conductivity.

Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were warnings for 2 files:

- Cast #51 – Because of a bad section in the secondary salinity, the primary channels should be chosen for the archive.

- Cast #76 - This was an upcast only. The SHFC1 cast was put through REVERSE and then DELETE with output extension REVDEL

- Cast #162 – This was a split cast and it looks like records from the end of file #162 should have been removed from the CLN files before joining the file with #163, as there were many spikes in the first file. When records 8485 to 11392 were removed and the join redone, SHIFT was rerun and then DELETE. There is one warning when this was done, but just notes that the gap in pressure between the two files is >2db. It is about 5db.

- Cast #172 – This was an upcast only. The SHFC1 cast was put through REVERSE and then DELETE with output extension REVDEL.

##### Other Comparisons

Previous experience with these sensors –

All sensors were used during 2 earlier cruises in 2017. The first was a cruise in the Strait of Georgia and Juan de Fuca Strait. The bottles from Juan de Fuca were trusted more than those from the Strait of Georgia because conditions were better for flushing Niskin bottles; they indicated that the primary salinity was close to bottles while the secondary was high by about 0.0045. Since flushing may not have been complete even in Juan de Fuca, both salinity channels could have been reading a little higher than those results suggest. The dissolved oxygen comparison from 2017-63 was also based on Juan de Fuca bottles.

Cruise 2017-03 was in Baynes Sound and the Strait of Georgia, with shallow sampling and poor flushing of bottles, so the results were not considered reliable enough to use for recalibration.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Temperature values were all within the climatology except for 2 minor excursions towards higher temperatures at LBP1 and NL02. There were more excursions from the salinity climatology that are not systematic, sometimes being fresher, sometimes saltier. There were more excursions near the coast and in the northern portion of the cruise. None of these observations suggest a calibration problem.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

Cast #139 has some bad primary salinity values. Problems were also found during the upcast of event #65. Examination of T-S plots shows little difference in noise levels between the 2 T/S pairs, with sometimes one looking better, sometimes the other. But there are occasional larger unstable features that are more common in the primary. The primary was chosen for the 2 previous cruises that used these sensors and it compares much better with bottles. The fact that we have a good comparison with bottles that will make recalibration of the secondary salinity reliable, makes choosing the secondary T and S channels look like a better choice overall.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. All files required some editing.

After editing T-S plots were examined for all casts which led to a little more editing being applied to 1.

Primary temperature and salinity were found to be better for casts #51 and 68.

##### Recalibration

There is no evidence to suggest that the pressure channel needs further adjustment.

The comparison with bottles suggests that the secondary salinity is high by 0.0054 with a standard deviation of 0.001. The fit is very flat against pressure and the analysis precision was excellent. So Salinity:T1:C1 should be recalibrated by subtracting 0.0054. The primary salinity requires no recalibration.

Based on the comparisons with bottles described in section 5, 2017-05-recal1.ccf was prepared to apply the following correction to channel Oxygen:Dissolved:SBE:

CTD DO Corrected = CTD DO * 1.0448 + 0.0527

File 2017-05-recal1.ccf was prepared to apply the corrections to Salinity:T1:C1 and Oxygen:Dissolved:SBE.

Theses corrections were first applied to the SAM files. COMPARE was rerun and the results confirm that the recalibrations were applied properly.

- Using the same selection of data as for the original fit, the secondary salinity was high by an average of 0.00008.

- When outliers were removed based on residuals the average of differences in the DO fit was <0.0009mL/L and the standard deviation was 0.012mL/L. CTD DO values are slightly high near the surface and low at depth.

(See file 2017-05-DO-comp2.xlsx for details.)

CALIBRATE was then run on the MRGCLN2 and EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When all data were included the CTD DO was higher than the titrated samples by an average of ~0.05mL/L and when some outliers are removed based on residuals, it is high by 0.04mL/L with a standard deviation of 0.13mL/L. There are two known sources of error that do not involve calibration drift: slow response of the DO sensor in high gradients and poor flushing of Niskin bottles. Both will lead to the downcast DO from the sensor reading higher than the bottles above the DO minimum (found ~800-1000m). The differences below 300m show the CTD reading very close to the bottles with an average difference of -0.0004mL/L and standard deviation of 0.009mL/L Near the surface the DO sensor is close to bottles, but between about 10m and 120m where vertical DO gradients are highest, the DO:SBE values tend to be higher than bottles.

No further recalibration is justified. See 2017-05-dox-comp3.xlsx for details.

##### Fluorescence Processing and special files for Dr. Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts except #51, 68, to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

REMOVE was run on casts #51, 68, to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Fluorescence and PAR data are nominal and unedited except that some

records were removed in editing temperature and salinity.

Transmissivity data are nominal. All values below 400m were removed because

the sensor malfunctioned at depth. Above 400m the data are of lower

quality than usual due to undersampling.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5 mL/L from 0 to 75db

±0.3 mL/L from 75db to 200db

±0.08 mL/L from 200db to 400db

±0.04 mL/L from 400db to 2000db

For details on the processing see document: 2017-05_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

A cross-reference list was produced.

The sensor history was updated.

The track plot looks fine.

Profile plots were examined and did turn up one problem – a large section of file #143 had bad transmissivity; those values were replaced with pad values in the REO file and Head Edit was rerun.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values ranged from 70% to 134% with most between 100% and 115%. For the casts farthest from shore values were mostly about 105%. Typical values for offshore are ~102% to 106% so that looks good. The highest and lowest values were close to shore. The values suggest that the DO calibration is reasonably good.

##### Final Bottle Files

The MRGCOR2 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts except #51, 68, to remove the following channels:

Scan_Number, Temperature:Primary, Transmissivity, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

REMOVE was run on casts #51, 68, to remove the following channels:

Scan_Number, Temperature:Secondary, Transmissivity, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and titrated DO, with mass units and REORDER was run to get the 2 SBE DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2017-05-bottles-final.xlsx.

Standards check and a header check were run on all files. No problems were found.

The track plot looks ok.

Plots of each file were examined to ensure no further problems were found.

A cross-reference listing was produced for the CHE files.

##### Thermosalinograph Data

There were 4 hex files. There were no loop samples taken. There is no file #1.

a.) Checking calibrations

There was a change to the configuration file after 2017-05-0003 because it was found that there was an error in the fluorometer calibration. One of the corrected files was renamed as 2017-05-tsg.xmlcon. No errors were found in the calibration parameters.

b.) Conversion of Files

The 4 files were converted to CNV files using configuration file 2017-05-tsg.con.

Those CNV files were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

A time-series plot showed that the differences between the intake and lab temperature were fairly steady with the lab temperature higher than the intake temperature by about 0.18C°. The salinity data is full of 1-sided spikes that look likely to be due to bubbles. The flow rate was high and fairly steady. The fluorescence is bad in the first, very short, file with constant and very high values. After that the traces look reasonable. At the end of the file #5, the flow was turned off so the last 19 hours of the record has no good data other than time and positions.

The track plot looks fine. The plot was added to the end of this report.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4db and exported to a spreadsheet which was saved as 2017-05-ctd-tsg-comp.xls. There were 90 casts which overlapped with TSG file 2017-05-0005. There was no overlap with the first 3 files.

The TSG file was opened in EXCEL and the files were reduced to the times of CTD files and loop samples. Those data were added to the comparison file.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. They were all ≤0.0003° except for a single event where a difference of 0.0023° occurred. The median differences were 0.0000° for both latitude and longitude. This shows both the times and positions are reliable for both systems. The one site where the differences were out of line was a case where the CTD file was started well ahead of the actual cast, so the ship would still have been moving when the position was recorded.

d.) Comparison of T, S and Fl from Rosette Samples and TSG and CTD data

- T1 vs T2 The intake temperature sensor worked throughout the cruise, and shows heating in the loop to be about 0.18C°. This is a reasonable value for the Tully in May/June.

- Flow Rate The flow rate was quite steady with a median value of 1.08. While values ranged from 1.04 to 1.22 values above 1.14 were only seen during events 170, 172 and 174.

- TSG vs CTD The spreadsheet comparing CTD and TSG files was then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. Using all data the results were:

|  | Flow Rate | TSG Intake-TSG Lab Temp | Latitude diff | Longitude diff | TSG intake – CTD Temp | SAL lab-SAL ctd | FL tsg / FL ctd |
| --- | --- | --- | --- | --- | --- | --- | --- |
| median | 1.08 | -0.1827 | 0.0000 | 0.0000 | 0.0118 | -0.1312 | 0.91 |
| average | 1.09 | -0.2092 | 0.0000 | 0.0001 | 0.1415 | -0.5862 | 0.97 |
| std dev | 0.03 | 0.1957 | 0.0001 | 0.0003 | 0.3915 | 1.7972 | 0.28 |
| max | 1.22 | 0.5110 | 0.0008 | 0.0023 | 2.2200 | 0.0038 | 1.85 |
| min | 1.04 | -0.9551 | -0.0003 | -0.0003 | -0.7627 | -14.1870 | 0.31 |
| Excluding outliers based on Tint-Tctd >0.1 or <-0.1 |  |  |  |  |  |  |  |
| median | 1.08 | -0.1805 | 0.0000 | 0.0000 | 0.0050 | -0.0947 | 0.93 |
| average | 1.09 | -0.1928 | 0.0000 | 0.0001 | 0.0078 | -0.1296 | 0.97 |
| std dev | 0.03 | 0.0815 | 0.0002 | 0.0003 | 0.0355 | 0.1189 | 0.22 |
| Excluding one position |  |  |  |  |  |  |  |
| median |  |  | 0.0003 | 0.0003 |  |  |  |

The intake temperature was higher than the CTD temperature by a median of 0.012C° (standard deviation 0.39C°) using all data, and 0.0050C° (standard deviation 0.04C°) when differences >0.1C° were excluded.

The salinity differences are large with a median of -0.131psu (standard deviation 1.797psu) and the time-series plots make it clear that this is due to problems in the TSG salinity. A typical pattern is that the salinity gradually drops ending in a spike to a much lower value followed by a sudden increase. Because of the one-sided nature of these drops and spikes, there is no obvious way to identify outliers. However, the largest outliers in the comparison do not look as though they are due to spikes, but rather due to very high local gradients (likely both vertical and horizontal gradients.) The deepest spikes are seen towards the end of the cruise when the ship was close to shore in areas where there are also large temperature differences. When records were removed that had been identified as temperature outliers, the median difference is only slightly lower at -0.095psu but the standard deviation is much lower at 0.119psu.

The TSG fluorometer was in reasonable agreement with the CTD fluorometer with a median ratio of 0.91 (standard deviation 0.28). Given a likely difference in depth of readings the variability is reasonable. However, when the TSG Fluorescence is plotted against CTD fluorescence it shows that the TSG readings are ~77% of the CTD readings. When the CTD Fluorescence is <1ug/L the TSG values are higher than the CTD and there are many values in that range.

(See 2017-05-ctd-tsg-comp.xls.)

- Loop Bottle - TSG Comparisons

There were no loop samples.

- Comparisons with rosette samples

No comparison was made with rosette salinity bottles since the TSG salinity is far too noisy for this to be useful. However, the extracted chlorophyll values were examined to see how they compared with the TSG fluorescence. There were few samples from close to 4m. The two closest showed the TSG fluorescence being between 60% and 80% of the CHL sample. Using all samples the ratio was ~70%.

- Calibration History

The temperature and conductivity sensors were recalibrated in March 2017 and this was the first use since then.

Conclusions

1. The TSG clock worked well.

2. The TSG flow rate was high and quite steady. It increased somewhat near the end of the cruise, but remained reasonably steady.

3. The temperature in the lab is higher than CTD temperatures by about 0.18Cº which looks reasonable based expected warming in the loop.

4. The TSG intake temperature is higher than the CTD temperature at 4m by a median value of 0.005Cº when outliers are excluded. This is may be partly due to the loop drawing water from slightly above 4m and/or there may be a slight warming of the water near the intake thermometer.

5. The TSG Salinity is reading lower than the CTD salinity and this appears to be due to bubbles. The largest differences are more likely due to local variability in areas of large horizontal and vertical salinity gradients. There is no reasonable way to edit these spikes and the variability means recalibration is not appropriate, especially since we only have comparisons while stopped. Fewer significant figures will be shown for the salinity to alert users to the fact that the data are less reliable than usual.

6. The TSG fluorescence is higher than the CTD fluorescence when the latter is <1ug/L and about 77% of the CTD fluorescence for higher values >1ug/L. There were no loop chlorophyll samples. Comparisons with rosette chlorophyll samples were limited but suggest that the TSG fluorescence was about 70% of chlorophyll values.

f.) Editing

The ATC files were copied to *.EDT.

Each file was opened in CTDEDIT and temperature, salinity, flow rate and fluorescence channels were examined. Only file 35 required editing. That was to remove values for the intake temperature, lab temperature, salinity and fluorescence during the last 19 hours when the flow was turned off.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Secondary, Temperature:Difference, Conductivity:Primary, Flag and Position:New.

The Fluorescence:URU:Wetlabs channel was removed from file #2 only.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats.

Those files were saved as TOB files.

The Standards Check and Header Check were run. The speed between TSG files #3 and 4 is high. But it was a very short break, hardly long enough to start a new file. It is possible that the computer clock was adjusted between the two files. According to the log book, time was adjusted at some point after the Haro59 stop by about 4 minutes. No other problems were seen.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

# Particulars

Out of Order: Cast #43.

Casts with bottles fired, but no sampling: 1, 22, 53, 84, 128, 139 and 149.

1. 24 Niskins fired at surface; no sampling. No bottle file needed.

2. Niskins #1-3 just tests; #4-6 sampled. After cast, configuration file changed and computer time adjusted.

6. Log said depth in header might be wrong, but it was correct.

10. Niskin #1 leaking.

11 - 20. Skipped Niskin #1 due to leak. Bot. cap on 7 leaky.

20. Niskin #1 closed but not sampled; Niskin #2 sampled.

22. Log doesn’t indicate bottles fired but 2 were fired – no sampling. No bottle file needed.

24. Fired bottle 17 as sample 42 due to possible leak in bottle 17.

27. Between this cast and next one – swap new viton O-rings, both caps, for old non-viton ones.

38. Niskin #1 fired accidentally.

47. After this cast top/bottom caps switched viton O-rings.

53. Bottle #9 fired just to test. No bottle file needed.

57. Up/down profiles very different.

65. Oxygen curve was strange on ascent at 30m.

67. Only bl and xmlcon files – no hex file. No mention in log about a problem.

76. Strange computer flips at 200m. No downcast data found.

81. Stopped a few times. Bottom depth crazy.

82. Closed all Niskins. Stop at 180 down because of wire angle. After cast bottom/top cap viton O-ring swapped.

84. Bottle #13 fired to test, not needed. No bottle file needed.

93. Bottom reading very unsteady.

97. Niskin 13 swapped top and bottom end caps; bottle continued to leak.

101. After cast fluorometer connector re-greased.

101. Adjustments to Niskins 13, 10, 5 to get the bottom cap to open more.

102. Fluorometer trace seems to be fixed.

103. Transmissometer cable changed after cast.

125. After cast top/bottom baps Viton O-rings replaced with old O-rings on Niskin #9.

128. Niskins 4, 22, 23 fired as test – all good. No Bottle file needed.

132. Slightly longer time than usual between file start and archive start. Split cast. Started cast 133 at 150m of downcast. Rename that file as #132 after join.

135. Station name in file wrong – should be SS3.

139. No sample number. No bottle file needed.

139. Problem in primary conductivity on downcast, cleared up at~315db. DO signal also affected. Upcast seems ok.

143. Transmissometer very spiky to 125db on downcast.

149. One bottle fired for incubations – no bottle file needed.

172. Computer blips – screen turned off at bottom of cast. Only upcast data acquired.

174. Niskin #1 closed but not sampled; other bottles sampled.

Extracts from Cruise Report on equipment problems

-TSG installed but not functional. Several connections and hose clamps not tight. Proper flow eventually established. Section of copper pipe use in section just before fluorometer. As copper is extremely toxic to phytoplankton, will attempt to replace this section with pvc or tygon. No supplies on board at present.

-CTD/Rosette: O-rings had been changed to Vitran, some issues with minor leakage. Attempts to solve ongoing.

-Transmissometer "semi-functional". Seems to stop working properly below about 300m. Spare on board had failed on previous cruise, not repaired before this cruise (so didn't use). Attempts to solve include changing and cleaning cables and connectors, to no good effect. --Occasional glitches with video on CTD computer, cuts out but returns. No cause found (yet). Also, the termination had to be retaped, as a new (more appropriate) termination has never been installed.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 16Dec2017 | Factory |  |  |
| Conductivity | 1766 | 19Dec2017 | Factory |  |  |
| Secondary Temp. | 4484 | 22Feb2017 | Factory |  |  |
| Secondary Cond. | 3531 | 30Mar2017 | Factory |  |  |
| Transmissometer | 1201DR | 27Apr2016 | Factory |  |  |
| SBE 43 DO sensor | 1119 | 28Mar2017 | Factory |  |  |
| PAR | 70613 | 21Mar2016 | Factory |  |  |
| SeaPoint Fluor. | 3685 | 13Mar2017 |  |  |  |
| Pressure Sensor | 0550 | 24Feb2017 | Factory |  |  |
| Altimeter | 62355 | n/a | Factory |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3411       Cruise ID#:	2017-05

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 4Mar17 | Factory |  |  |
| Conductivity | 2488 | 4Mar17 | Factory |  |  |
| WetLabs Fluorometer | Ws3s-953p | May2017 |  |  |  |
| Temperature:Secondary | ? | 0842 |  |  |  |
| Flow meter | ? | n/a |  |  |  |