## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 20 March 2025 | ### Updated channel names & formats in TOB & loop files. GG & SH |
| ### 18 April 2023 | ### Corrected DIC/ALK format, fixed few flags and comments. G.G. |
| ### 6 Dec 2021 | ### Corrected Salinity:Bottle precision lost during HPLC addition. Re-did the Total Nitrates addition for casts 8, 30, 57, 75 & 79. S.H. |
| ### 15 April 2021 | ### Added DIC and TA data, changed phosphate flags for samples 300& 302 in event 46. S.H. |
| 14 May 2021 | Added DOC/TOC/TDN/TN to 11 CHE files. S.H. |
| ### 2 Sept 2020 | ### Added HPLC data. S.H. |
| 27 Nov 2019 | Added DMSP –all values padded due to problem with blanks G.G. |
| 20 Dec. 2017 | Update CHL loop data flags/comments and update flag/comment CHE file sample 55,event 12 |
| 14 Dec. 2017 | Change TSG salinity format; corrected TSG comments. |
| 11 Dec. 2017 | Update to CHL spreadsheet and CHL precision statement in CHE files. |

## PROCESSING NOTES

Cruise: 2017-06

Agency: OSD

Location: North-East Pacific

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: 4 June 2017 – 20 June 2017

Processed by: Germaine Gatien

Date of Processing: 30 August 2017 – 25 October 2017

Number of original HEX files:  	49	Number of CTD files: 49

Number of bottle files: 		49	Number of bottle casts processed: 48

Number of original TSG files: 	3	Number of processed TSG files:

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1201DR), a SBE 43 DO sensor (#1119) on the primary pump, a SeaPoint Fluorometer (#3685) on the secondary pump, a Biospherical QSP-400 PAR sensor (#70613), and an altimeter (#62355).

A thermosalinograph (Seacat 21 S/N 2488) was mounted with a Wet Labs WETstar fluorometer (S/N ws-3s 953p), remote temperature sensor and a flow meter.

Seasave version 7.26.1.8 was used for acquisition.

The data logging computer was the Tully CTD Laptop.

The deck unit was a Seabird model 11+, serial number 0425.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial #68542.

The oxygen kit was SIO (nsB-6009) model 2.36a kit #2.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

Note that as of January 2017 a change has been made to the threshold levels for quality flags 3and 4 for Extracted Chlorophyll. Flag=3 will now be 10%CV (formerly 15%) and Flag=4 will now be 30% (formerly 50%).

The Daily Science Log, rosette log sheets and analysis logs were in good order with detailed comments on problems. Notes from the Chief scientist were very helpful as was the log of loop sampling.

The transmissivity data were bad below 250m of the downcast and continued to be bad for most of the upcasts. All values below 250m have been removed from the CTD files and the Transmissivity channel was removed from all the bottle files. Even above 250m these data are considered nominal and should be used with caution.

Both CTD salinity channels had spikes. There was excellent precision in the salinity analysis and it was done very soon after the cruise. The primary temperature/salinity pair was selected for archiving because the primary salinity showed little variation with pressure or time. The secondary salinity appears to have some pressure dependence or temporal drift, more likely the latter, but having the deepest sampling later in the cruise complicates that judgment.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples, especially when far from shore. The data show the usual pattern for this type of fluorometer with the ratio FL/CHL being high at very low CHL values and dropping as CHL rises. There were only 2 CHL samples from the rosette with values >0.7ug/L.

The comparison of titrated dissolved oxygen samples with CTD dissolved oxygen looks quite tight. There was notable hysteresis in the initial comparison, so the hysteresis correction factor E was adjusted and when the comparison was rerun the results were good.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.3 mL/L from 0 to 100db

±0.2 mL/L from 100db to 250db

±0.1 mL/L from 250db to 500db

±0.04 mL/L from 500db to 2500db

±0.02 mL/L below 2500db

Overall the TSG performed well with intake temperature available throughout the cruise, but there were some problems.   Channel Salinity:T0:C0 is given with only 3 decimal places due to the presence of      many large 1-sided spikes. While comparisons with CTD data, rosette samples and loop samples indicate that TSG salinity values were low by between 0.2psu and 0.5psu, recalibration is not justified due to great variability in the comparisons and in the spikiness of the traces. The flow meter malfunctioned during file #3, but the temperature trace suggests that the flow was fine. GPS positions are missing for about 11 hours during the 2nd file, but positions were available from a Fukawi file; where no Fukawi value was available it was assumed that the ship had not moved, so duplicate positions were used.

The TSG fluorometer readings were fairly close to the CTD fluorometer and chlorophyll samples from the loop. The TSG fluorescence read high for low chlorophyll, but as CHL rose the ratio fell. The TSG fluorescence was closer to CHL loop samples when the ship was underway (ratio FL/CHL ~0.9) than when stopped (ratio ~1.8), even though the median loop CHL values are close. The increase of fluorescence when the ship stops can sometimes be seen in the fluorescence traces. Varied flow in the loop might explain this but the flow meter shows no change. Could bubbles be affecting fluorescence? Whatever the cause, the underway fluorescence appears to be more accurate for this cruise than that while the ship was stopped.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Sampling notes from the chief scientist noted a number of issues. Based on these notes a few corrections were made in the raw files.

Nutrients, extracted chlorophyll, dissolved oxygen, DMS and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked. They had all been used on 3 previous cruises.

The configuration changed stayed the same throughout the cruise.

All parameters in the configuration files were correct.

NOTE: After the bottle comparison described in section 5 the configuration file was changed; factor E in the dissolved oxygen sensor calibration was changed from 0.036 to 0.038 to improve the hysteresis correction.

##### Conversion of Full Files from Raw Data

All files were converted using 2017-06-ctd.xmlcon.

The hysteresis and Tau functions were selected. Depth was included in the conversion.

A few casts were examined and all expected channels are present.

As usual the temperature and salinity channels track fairly well during downcasts but are noisier on upcasts. There are spikes in both primary and secondary channels.

PAR, Fluorescence, Dissolved Oxygen and Altimetry look normal.

Transmissivity looks bad below about 250m. A similar problem was noted during the previous cruise, 2017-05, although the issue did not seem to arise until the CTD reached about 350m. The upcasts look like they are mostly bad, as seen in 2017-05. At the depth where the profiles look normal, the signal is “steppy” due to sampling at less than the usual 24Hz; this was something set at the factory.

Acquisition for a few casts was started right at the surface after the 10m soak. According to the notes from the chief scientist this was done to obtain PAR values out of the water and just below the surface. The rosette was said to be “half-way out of the water”. This would suggest pressures from the CTD would be about 0.8db, though it is a rough description.

During cast #9 during the first 15s the pressure was between +0.6db and +1.1db. The PAR signal goes down very rapidly once pressure become >1.2db. At no point does conductivity suggest that the CTD itself was out of water.

During cast #57 the CTD pressure was about +0.5db for a minute with a few records very briefly about -0.3db. The PAR values are high when the pressure is <1.2db, but only when the pressure is <0 does the salinity drop somewhat. The CTD was probably near the surface for too short a time to get very low salinity though it does get as low as 20psu. At the end of that cast the CTD acquired data as it went through the surface as evidenced by a sudden drop in salinity about 30s after the pressure reached negative values.

These results suggest that the pressure is reasonably accurate.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2017-06-ctd.xmlcon. Depth was included.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. CTDEDIT was used to clean a few points in the primary salinity channel of casts #62, 70 and 92.

A preliminary header check and no problems were found. CTD fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs. A few alterations were needed:

- Cast #51 had 15 bottles fired according to the rosette log, but 16 bottles closed according to the rosette file.  Niskin #3 closed at the surface and was the 11th bottle fired – there is a note on the log to indicate it was fired accidentally but the firing order entries were not corrected. No samples were taken from that Niskin so the line was removed from the ADDSAMP list.

- Niskin #14 was mistakenly fired during event #79; it was not sampled and has no sample number. That line was removed from the ADDSAMP file.

- Event #104 had 24 bottles fired but only 3 sampled, so the lines were removed for bottles 4-24.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The unwanted line for Niskin bottle 3 was dropped from file 2016-06-0051.SAMAVG.

The unwanted lines for Niskin bottles 4-24 were dropped from file 2016-06-0104.SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2017-06-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2017-06chl*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2017-06chl.csv. Loop data were saved in a separate file. The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2017-06oxy.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and loop data were moved to a separate file. The rosette file was saved as 2017-06oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in file 2017-06SAL.xlsx which included a precision study. The analyses were done 5 to 18 days after collection. Precision is lower than expected given the timely analysis.  The nylon inserts used were those that were found to provide poor seals, not the inserts we normally use.

The files were simplified and saved as 2017-06sal.csv. File 2017-06sal.csv was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2017-06_nutrients*.xlsx. This includes a precision study. The file was simplified and the saved as 2017-06-nuts.csv. The file was converted to individual NUT files.

DMS

DMS data were obtained in file DMS summary (2017-06).xls. Values given as < were changed to 0 and those given as – were replaced with pad values; the comments that will go into the header will explain that 0 means below minimum detectable level. A flag and comment were added based on a comment entered below the table. DMS: was entered before the comment. The file was then saved as 2017-06DMS.csv and converted to individual DMS files. There was a separate report on analysis techniques and problems.

The SAL, CHL, OXY, NUT and DMS files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. This step is important to avoid missing some samples in the creation of CHE files.

Some problems were found:

- Event #12 –The DMS sample said to be from Sample #65/Niskin 13 at 0m, but that Niskin was fired at 5m and Sample #66/Niskin 14 was fired at the surface. The rosette sampling log shows the sample being #66 from Niskin 14 at ~0m. The analyst checked the labels and found that the sample # should have been 66, not 65. This was corrected in the spreadsheet and in the bottle files for cast #12.

- Event 51 – Sample 357 is missing from final sheet, but is in the RAW data. The values were added to the CHL file; the analyst confirmed this and fixed the QF spreadsheet.

- There are some NH4 samples entered in the sampling log but the data were gathered for and analyzed by university researchers.

- One nutrient sample (# 624) was reported as being from CTD cast #115, but there was no CTD cast at event 115. It was a loop event. The sample is from the GO-FLO at event 116.

- There is also a salinity sample for #624 said to be from cast #115 but it was actually a Go-Flo sample.

The following lines were removed from the MRG files because bottles were fired but not needed and were not assigned sample numbers:

- Cast 51, Niskin #3

- Cast 79, Niskin #14

- Cast 104 Niskins #4-24

CLEAN was run to add quality flag 0 to empty entries and to reset the header limits.

There are loop data in the salinity, nutrient, chlorophyll and oxygen spreadsheets, so those data were moved to file 2017-06 Loop_Data.xlsx.

The chief scientist provided a log of loop sampling. There were loops taken while underway and others were taken at the end of CTD casts while the 5m rosette was being fired.

Times were added to the loop file based on the end time of CTD casts or the loop time if there was no CTD cast at the time.

##### Compare

Salinity

Compare was run with pressure as reference channel.

Two major outliers were found.

- Sample 213 is reported as being from cast #36 but there was no such sample for that cast. Somehow in the merging process it got added to the data for sample #217 but there was no salinity sampling for that bottle. This was reported to the analyst and removed from the bottle file for cast #36.

- Sample 296 for cast #46 is way out of line in profile and in comparison to the CTD. This was reported to the analyst who added a 3 flag and comment.

A few other outliers were from the near-surface and associated with large standard deviations in the CTD salinity, so the bottle values may be ok.

If we exclude data with standard deviation in the CTD Salinity >0.0008, Pressure<1000db and one outlier, we find:

|  | Average Difference | Standard Deviation | Fit against Pressure |
| --- | --- | --- | --- |
| Sal0 – Sal Bottle | -0.0012 | 0.0006 | Diff = -2E-07*Press - 0.0009 |
| Sal1 – Sal Bottle | +0.0032 | 0.0009 | Diff = -6E-07*Press + 0.0043 |

The two salinity channels are growing closer as pressure increases with the larger drift being in the secondary. Looking at the drift with time the secondary is also drifting lower more than the primary. When outliers are gradually removed to make the fit against pressure quite flat, the time variations also flatten somewhat, though for both, the secondary is larger. The temporal variation could be due to there being deeper casts later in the cruise, but when bottles are restricted to particular depth ranges (such as 2900-3100db) there still seems to be some time-dependence.

On the whole the bottle comparison is not as tight as seen when these sensors were last used despite quick analysis of bottles and excellent precision reported for the analysis. So there does seem to be real drift in at least the secondary sensors. During 2016-05 there was very little drift in the secondary salinity and the 2 salinity channels were further apart. That cruise only sampled to about 2000db.

When only data above 1000 to 1500db were included in the comparison the secondary fit looked flat with pressure and many of the 2000db values also looked fine. While there may be some pressure dependence in the secondary sensors that was not noted during the earlier cruise due to shallower sampling, but it does not appear to be significant. The temporal change looks like it is due to the secondary salinity decreasing and getting closer to the primary.

The study of 22 bottles fired at 2000db during cast #83 found a standard deviations of 0.0008 in the differences between bottles and both salinity channels. However, that included a bottle that is an outlier in profile and in the overall comparison. Removing that outlier produced standard deviations of 0.0006. The primary salinity was low by an average of 0.0014psu and the secondary was high by an average of 0.0025psu. This is a good result with the variability in line with the analysis precision study and that suggests there was consistent flushing of bottles.

For full details for the COMPARE run see file 2017-06-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There was one significant outlier, sample 30 from cast #4 at 30db. At that depth there was a very large DO gradient, with DO changing by about 0.8mL/L during the stop, so incomplete flushing of the Niskin bottle likely accounts for it reading lower than the CTD.

A check was made for hysteresis by excluding all points from below 2000m and there was clear evidence of hysteresis as the deep values are distinct from the shallow.

A preliminary check was made of one of the deepest casts by changing the value of E in the dissolved oxygen calibration from the default value of 0.036. When 0.038 was used the effect looked about right, so all the files with DO sampling were converted using that value, new MRG and SAM files were prepared and COMPARE was rerun. This time the deep values did not stand out.

So the configuration file was updated with E=0.038 and steps repeated using the new E value.

The hysteresis check shows good correspondence between comparisons using deep and shallower bottles.

When outliers were removed based on residuals, the fit was:

CTD DO Corrected = CTD DO * 1.0399 + 0.0645 (R2 = 0.9631)

During 2017-05 when the sensor was last used the fit was:

CTD DO Corrected = CTD DO * 1.0448 + 0.0527 (R2 = 0.9475)

For that cruise the E factor had not been corrected but there were no bottles below 2000m and most casts were quite shallow. The hysteresis check done then did not show a significant effect.

For more details see file 2017-06-DO-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor. Plots are distorted by the fact that all but 2 values have CHL<0.7ug/L, with the other two being 1.4 and 12.7ug/L. The data show the usual pattern for this type of fluorometer with the ratio FL/CHL being high at very low CHL values. At CHL=1.4ug/L the fluorescence was ~70% of the CHL and at 12.7ug/L it was 55% of the CHL.   When the 2 high CHL values are excluded the fluorescence is roughly 3 times the CHL.

For full details of the comparison see file 2017-06-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run during 2017-05 when the same equipment was in use and the best results were found using a setting of +2.5s. ALIGNCTD was run on all casts using +2.5s. A few casts were examined after this step and the results look reasonably good, though noisy upcasts make such a judgment very subjective.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice. One cast was checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. Some results from other cruises that used these sensors are included for comparison; those entries are shaded.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2017-63-0074 | 400 | -0.0001 | +0.00037 | +0.0040 | F.High, Steady |
| 2017-63-0083 | 400 | 0 | +0.00038 | +0.0039 | F.High, Steady |
| 2017-63-0107 | 320 | -0.0006 N | +0.00040 N | +0.0045 N | F.High, Steady |
| 2017-63-0129 | 300 | -0.00045 N | +0.00048 N | +0.0050 N | F.High, F.Steady |
| 2017-03-0027 | 90 | ~0.0001 | ~0.0005 | ~0.0054 | High, Steady |
| 2017-05-0038 | 1000 | +0.0001 | +0.0004 | +0.0046 | High, X Noisy |
|  | 1900 | +0.0003 | +0.0004 | +0.0040 |  |
| 2017-05-0101 | 1000 | +0.0006 | +0.0005 | +0.0055 | High, Moderate |
|  | 1900 | +0.0004 | +0.0005 | +0.0052 | “ |
| 2017-05-0139 | 1000 | +0.0002 | +0.0005 | +0.0051 | High, X Noisy |
|  | 1900 | +0.0004 | +0.0005 | +0.0055 | “ |
| 2017-06-0036 | 1000 | +0.0002 | +0.0004 | +0.0047 | High, X.Noisy |
|  | 1900 | +0.0004 | +0.0004 | +0.0045 | “ |
|  | 2900 | +0.0003 | +0.0004 | +0.0042 | “ |
| 2017-06-0064 | 1000 | ~0 | +0.0004 | +0.0047 | High, Noisy |
|  | 1900 | +0.0003 | +0.0004 | +0.0042 | “ |
|  | 2900 | +0.0004 | +0.0003 | +0.0038 | “ |
| 2017-06-0085 | 1000 | +0.0002 | +0.0004 | +0.0043 | High, V Noisy |
|  | 1900 | +0.0002 | +0.0003 | +0.0040 | “ |
|  | 2900 | +0.0002 | +0.0003 | +0.0036 | “ |

The temperature differences are small and similar to other recent cruises. The conductivity and salinity differences are smaller than during 2017-05, and salinity differences appear to be getting smaller with time. The 2 earlier cruises that used these sensors did not go deep enough to provide useful information.

The results of the bottle comparison discussed in section 5 are consistent with the slowly reducing differences between salinity channels. The secondary salinity is likely responsible for most of the drift.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number. Transmissivity values below 250m were replaced with pad values.

##### Checking Headers

A cross-reference list was checked against the log book. No errors were found.

The cruise track was plotted and added to the end of this report.

A surface check was run and shows an average surface pressure for the cruise was 2.1db. This is a little lower than usual for the Tully, but the associated salinity values are very low. Cast #57 was examined as it had some negative pressures at the end of the cast with the pumps still on. After about 1.5 seconds CTD values suggest it was out of water, though likely close enough to the surface to have some water splashing around it, the pressure was about -0.4db. The earlier study of PAR values at the beginning of casts suggested the pressure was reasonably accurate.

The header check showed that there were spikes in pressure for at least one cast. No other problems were noted and fluorescence did not go off-scale.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. Water depths were compared with the log book entries. There are some discrepancies but it is believed that most were adjusted by the chief scientist and they do look like sounder depths may have been wrong, or that drift occurred before the cast actually started. A calculation was made of water depth - maximum sampling depth + altimetry header. Where that value is >4m the entry was investigated. A few changes are needed to the headers:

- The depth entry for cast #14 is 784 in the log and the header. But in January the depth at that site was 815m. The altimetry looks fine and implies that the bottom depth is about 815m. That value was entered in the header.

- The depth entered in the file header for event #16 at P4 is 1320 while in the log it is 1299. The log entries for that site vary somewhat.  The altimetry looks good for event #16 and that implies a depth of about 1299 makes more sense. So the header entry was changed to match the log.

- For cast #85 the chief scientist indicated that the file header depth was wrong so it was changed from 4250 to 4225, but the altimetry reading shows that 4250 is more appropriate and the depths entered in January 2017 for this site was about 4250. The value 4250 was entered for #85 and also for casts #85, 87 and 92.

- The altimetry headers in casts 21, 22, 54 and 62 were caused by spikes; the CTD did not get near the bottom so the headers were removed.

- For the bottle files, the same changes were made as above plus header entries were removed from the following casts that did get near the bottom having only near-surface sampling: 36, 54, 62.

After correcting the SAM files they were bin averaged again and the final merge step was run for the bottle files; the CLEAN step was repeated as well.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 3 casts to see if the shifts applied to conductivity during the previous cruise when the same equipment was in use were appropriate and they were. As found during 2017-05 there is a lot of noise in the data and no alignment will remove it all.

SHIFT was run twice on all casts using -0.6 records for the primary conductivity and -0.6 records the secondary. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warnings came from the upcast of one file, so are of no concern.

##### Other Comparisons

Previous experience with these sensors –

All sensors were used during 3 earlier cruises in 2017. The first was a cruise in the Strait of Georgia and Juan de Fuca Strait. The bottles from Juan de Fuca were trusted more than those from the Strait of Georgia because conditions were better for flushing Niskin bottles; they indicated that the primary salinity was close to bottles while the secondary was high by about 0.0045. Since flushing may not have been complete even in Juan de Fuca, both salinity channels could have been reading a little higher than those results suggest. The dissolved oxygen comparison from 2017-63 was also based on Juan de Fuca bottles.

Cruise 2017-03 was in Baynes Sound and the Strait of Georgia, with shallow sampling and poor flushing of bottles, so the results were not considered reliable enough to use for recalibration.

Cruise 2017-05 was mostly off the west coast of Vancouver Island where flushing was good at depth. The comparison between the bottles and CTD salinity was quite tight; the primary salinity was found to be very close to bottles and the secondary was high by 0.0054psu. The dissolved oxygen comparison was also quite tight and led to a calibration of slope 1.0448 and offset +0.0527.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Salinity was slightly high just below the base of the mixed layer at some, but not all casts at station P4, and slightly high at P5 and P21. All other salinity values were well within the historic range. The low salinity values look like cases of a deeper halocline rather than calibration problems.  All temperature data fell within the climatology.

Repeat Casts – There were repeat casts. Casts #85 and #91 occurred about 7 hours apart at P26. Differences along lines of constant σt were almost indistinguishable before bin-averaging and after binning they were <0.002º for the primary temperature and <0.0002psu for the primary salinity at σt= 27.52 (near 1450m). This shows excellent repeatability.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

Both CTD salinity channels had a lot of spikes especially near the surface. There was excellent precision in the salinity analysis and it was done very soon after the cruise. The primary temperature/salinity pair was selected for archiving because the primary salinity showed little variation with pressure or time. The primary salinity was also found to be closest to bottles, reading low by ~0.001psu.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. Some bad salinity points were removed from a few files. All files required some editing. The descent rate was generally very noisy for the offshore casts and often had some complete reversals of direction so there was a lot of corruption by shed wakes.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts; no further editing was done.

##### Recalibration

There is insufficient evidence to justify recalibration of the pressure. The sensor calibration may be drifting to values that are slightly low, but that is not clear. It appears to be well within ±0.4db.

The primary salinity was found to be lower than the bottle salinity by an average of 0.0012psu. Since slight inefficiency in Niskin flushing could account for bottle values being slightly high, this is not considered a significant difference. Salinity will not be recalibrated.

Dissolved oxygen recalibration was discussed in section 5.

File 2017-06-recal1.ccf was prepared to apply the following correction to channel Oxygen:Dissolved:SBE:

CTD DO Corrected = CTD DO * 1.0399 + 0.0645

This correction was first applied to the SAM and MRGCLN2 files. COMPARE was rerun for dissolved oxygen using roughly the same points as in the fit used for recalibration. The average of differences in the DO fit was -0.0002mL/L and the standard deviation was 0.017mL/L. The fit against CTD DO was flatter than usual. (See file 2017-06-DO-comp2.xlsx for details.)

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When outliers were removed based on residuals the CTD DO was higher than the titrated samples by an average of ~0.012mL/L (standard deviation of 0.035mL/L). Looking at the differences versus pressure shows that the sensor DO tends to be on the high side above 600m, which is likely due to incomplete flushing of Niskin bottles combined with higher gradients.

No further recalibration is justified. See 2017-06-dox-comp3.xlsx for details.

##### Fluorescence Processing and special files for Dr. Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases. The SAM files were put through REMOVE and HEADEDIT. Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. The transmissivity profile frequently looks unrealistic below 250m. CLEAN was run on all casts to pad that channel below 250m. Even above that the values must be considered nominal.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

PAR was removed from casts 32, 46, 51, 64, 85.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Fluorescence and PAR data are nominal and unedited except that some records were

removed in editing temperature and salinity.

The Transmissivity data are nominal. Values were clearly bad below about 250m so

data below that level were replaced with pad values. Similar results were

found when this sensor was used during other recent cruises.

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.3 mL/L from 0 to 100db

±0.2 mL/L from 100db to 250db

±0.1 mL/L from 250db to 500db

±0.04 mL/L from 500db to 2500db

±0.02 mL/L below 2500db

For details on the processing see document: 2017-06_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found.

A cross-reference list was produced.

The sensor history was updated.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The values ranged from 75% to 85% in Juan de Fuca Strait, about 90% in Haro Strait and from P2 to P26 they are all between 102% and 108% with those furthest offshore near the lower end of that range. The results are consistent with good dissolved oxygen calibration since we usually see values ~103% offshore.

##### Final Bottle Files

The MRGCOR2 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

PAR was removed from casts 36, 46, 51, 64 and 85.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2017-06-bottles-final.xlsx. The entries were compared with the rosette log sheets to ensure no samples had been missed. Only one problem was found and corrected:

- Event 51: The rosette file indicates that a CHL sample was taken from Niskin 17, sample #332, but no sample was analyzed. The label for the intended sample was found on the rosette sheet which indicates it was not taken. This was corrected in the MRG, MRGCLN2, MRGREM, MRGREO and CHE files.

Standards check and a header check were run on all files. No problems were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files.

##### Thermosalinograph Data

There were 3 hex files.

a.) Checking calibrations

The configuration file did not change through the cruise, but contained an error in the fluorometer parameters. One file was renamed as 2017-06-tsg.xmlcon and the fluorometer calibration parameters were fixed.

b.) Conversion of Files

3 files were converted to CNV files using configuration file 2017-06-tsg.con.

Those CNV files were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

A time-series plot showed that the salinity signal was full of one-sided spikes likely to be due to bubbles, and unlikely to be amenable to editing. The flow rate was very steady, mostly around 1.1, but up to 1.2 for a few hours at the beginning. However, the flow signal went to zero during file #3. Since the temperature trace looks fine, this appears to be a flow meter failure rather than a real stop in the flow.

A more serious problem is that the GPS feed was lost on June 6 from about 0329 to 1451 UTC. At that point the file was closed and a new one opened with good positions. For at least 10 hours of that gap the ship was at station P4. A Fugawi navigation file with times and positions from when the ship was moving was provided by the chief scientist but the entries are not evenly spaced in time, with some parts heavily sampled and some gaps; also the format was awkward, but it was edited to enable saving as a CSV file.

The first ATC file was opened in EXCEL.

After some manipulation the Fugawi file (2017-06-time-positions.txt) was opened in EXCEL and sorted so that times ending close to 25s and 55s were selected to enable reasonable matches to the TSG file. Those data were added to the main file and aligned to match the times in the main file which was saved as 2017-06-0001fixed.csv. Where there are gaps it was assumed that the ship moved very little so positions were entered as identical to those preceding or following the gap. Thus the positions are less certain than usual, but considered close enough for most purposes.

Most of the header was removed leaving just the line with column names, and that file was converted to an IOS header file. The headers are missing, so those were added using a text editor and copying the headers from the originally converted file, 2017-06-0001.ios. This process took a few attempts as errors were found and resolved. The converted file was called 2017-06-0001.fix. It was put through CLEAN and ADD TIME CHANNEL.

The temperature differences were very noisy for file # with a median value -0.17C° when all data are included and-0.16C° for a quiet section late in the cruise.  As usual, the heating in the loop is higher when the intake temperature is lower. For the other two files the median values for quiet sections were -0.19 C° and -0.18 C°. The variations are in line with the median intake temperatures which were lowest for the 2nd file.

The TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, lab temperature, salinity and fluorescence and the files were reduced to the times of CTD files. There were 47 cases of overlap, but the last one was at a time when the flow meter was not working, so that was removed from the comparison, though the flow may have been fine. However, the difference was extremely noisy for that case, so it is not a great loss. Similarly, TSG data were found to match the loop samples.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude and longitude were all ≤0.00035° and the median differences were 0.0000° for both. This shows both the times and positions are reliable for both systems.

d.) Comparison of T, S and Fl from Loop & Rosette Samples and TSG and CTD data

- T1 vs T2 The intake temperature sensor worked throughout the cruise. The differences decrease slightly as temperatures increase as is expected as the intake temperature gets closer to the temperature of the ship. There were 3 significant outliers associated with large standard deviations in the intake temperature.

- Flow Rate The flow rate ranged from 1.0 to 1.2 while the flow was turned on; the median value was 1.13.

- TSG vs CTD The spreadsheet comparing CTD and TSG files was then examined to find the differences between the salinity, fluorescence and temperature from the CTD and the TSG.

| Using all data and median values over 2 minutes for TSG: |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | (TSG intake -CTD) Temp | (TSG lab -CTD) Temp |  | (TSG - CTD) Sal |  | FLtsg/FLctd |
| median | 0.0043 | 0.1988 |  | -0.0283 |  | 1.133 |
| stdev | 0.2581 | 0.1944 |  | 0.0436 |  | 0.272 |
| avg | 0.0687 | 0.2429 |  | -0.0315 |  | 1.235 |
| min | -0.0831 | 0.0722 |  | -0.2297 |  | 0.909 |
| max | 1.6257 | 1.3516 |  | 0.1384 |  | 1.986 |
| Using the 10 cases with lowest standard deviation: |  |  |  |  |  |  |
| (TSG intake -CTD) Temp | (TSG lab -CTD) Temp |  | (TSG - CTD) Sal |  |  |  |
| 0.0040 | 0.1915 |  | -0.0226 |  |  |  |

- The flow rate had no noticeable effect on the differences between lab and intake temperatures.

- The intake temperature is higher than the CTD temperature by a median value of ~0.0043C° with only a slight reduction to 0.0040C° when only the 10 cases with the lowest standard deviation are included.

- The TSG salinity is lower than the CTD salinity by a median value of 0.028psu with a standard deviation of 0.043psu. When only the 10 cases with the lowest standard deviation are included the difference is reduced to 0.023. To see if the spikes in the record could account for this difference, a few randomly chosen sections were examined. A rough estimate suggests that about 5% of values are outliers by at least 0.3psu. In one section 141 out of 150 salinity values were close to 31psu. The other 9 values were lower by an average of 0.7psu (ranging from 0.3psu to 1.3psu). The average for the 150 values was 30.946. So including those 9 values produces salinity that is lower by 0.05psu than what looks to be the appropriate value. A second section had smaller spikes that led to a value that was lower by an average of ~0.02psu than the values without spikes. This analysis only includes what are clearly spikes; there are small drops in values that may well be due to bubbles as well and there may be effects that are not noticeable.

- The ratio of the TSG fluorescence to CTD fluorescence has a median value of ~1.1 and standard deviation of ~0.3. The ratio is higher at the low end of the range, settling to close to 1.

See 2017-06-ctd-tsg-comp.xls.

- Loop vs Rosette

Differences were calculated between samples from the loop and from the rosette samples closest to 5m wherever possible. There were a few significant outliers:

- The nitrate from rosettes for stations P2, P4 and P8 had values 0.0 umol/L down to at least 10m, but the loop nitrate for station 16 at P4 had a value 0.74 umol/L. The loop samples from P2 and P8 also had values of 0.0 umol/L.

- The loop salinity from event 16 at P16 is much lower than the rosette sample.

- The loop silicate from event 30 at P8 looks a little out of line, but the silicate profile from the rosette looks complex, so this could be a depth mismatch rather than a bad value.

When those outliers were removed from the comparison, the following differences were found:

|  | Rosettte Samples - Loop Samples excluding 3 outliers |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | SAL | CHL | OXY | NITRATE | SILICATE | PHOSPHATE |
| Median Diff | 0.0075 | 0.0575 | -0.0912 | -0.1057 | 0.0486 | 0.0060 |
| Min | -0.0020 | -0.0100 | -0.2130 | -0.4800 | -0.4200 | -0.0410 |
| Max | 0.0469 | 0.2200 | 0.1000 | 0.1200 | 0.4100 | 0.0390 |
| Std Dev | 0.0193 | 0.0765 | 0.1019 | 0.2259 | 0.2874 | 0.0274 |

Extracted Chlorophyll, Dissolved Oxygen and Phosphate comparisons look good.

The salinity comparison is not as good as usual, but if one more outlier were excluded, it would be excellent. 5 out of 7 differences were ≤0.002 and 4 out of 7 <0.001psu.

The nitrate and silicate differences are larger than seen in January 2017, but the range of values is much larger so that may not be significant.

Loop Bottle - TSG Comparisons

There were 19 loop samples each for salinity and extracted chlorophyll. The TSG salinity was lower than loop samples by a median of 0.054psu with a standard deviation of ~0.3psu. Loops taken while the ship was stopped are low by a median of 0.016 (std dev ~ 0.1) and while moving they are low by a median of -0.27 (std dev ~0.3).

Extracted chlorophyll values ranged from 0.2 to 5.1ug/L. The median ratio of TSG fluorescence to CHL varies from 0.4 to 7.3 with a median of 1.7. There is little difference between the median CHL values while stopped and while underway, but the TSG fluorescence is higher than the loop chlorophyll samples by a factor of ~1.8 while stopped and by ~0.9 while underway. While there are not enough loop samples to say this with confidence, examination of TSG plots during some long stops does show somewhat elevated values. As the ship starts moving values fall. This may mean that the TSG fluorescence is more useful when moving than when stopped. From the salinity comparison we think there are fewer bubbles while stopped, but that would not seem to explain higher TSG fluorescence. Perhaps the flow in the loop is not as high while stopped, though the flow meter does not suggest that. There are too many variables such as high sea states and track direction to make much sense of this, but for at least this cruise.

A final check was to see how the TSG values compared with the rosette values. The comparison with the CTD done earlier matches downcast CTD data with the TSG. In this case we are matching upcast CTD data to TSG data gathered when the loops were taken. The median value shows the TSG salinity to be reading lower by ~0.039psu and the fluorometer to be reading high by a median factor of 1.24. This compares with values -0.028 and 1.13 for the downcast. Given we have fewer comparisons for the upcast and the TSG salinity spikes are random, this looks reasonably consistent for salinity and very close for fluorescence.

(See 2017-06-loops-tsg-rosette-comp.xls.)

- Calibration History

The temperature and conductivity sensors were recalibrated in March 2017 and this was the second use since then.

2017-05: The temperature in the lab was higher than CTD temperatures by about 0.18Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of ~0.005Cº. The TSG Salinity read lower than the CTD salinity and that appeared to be due to bubbles. The TSG fluorescence was higher than the CTD fluorescence when the latter is <1ug/L and about 77% of the CTD fluorescence for higher values >1ug/L. Comparisons with rosette chlorophyll samples were limited but suggest that the TSG fluorescence was about 70% of chlorophyll values. No recalibration was applied to the TSG data.

Conclusions

1. The TSG clock worked well.

2. The TSG flow rate was at the expected level and generally steady, but the flow meter failed during file #3.

3. The temperature increases in the loop by roughly 0.18Cº and increases as intake temperature decreases, as expected.

4. The TSG intake temperature was higher than the CTD temperature by a median of about 0.004Cº.  This is likely as close as we can achieve given the variability at 4m depth and the fact that the TSG may draw temperature from a little higher in the water column and that there could possibly be some heating right at the intake. No correction will be applied.

5. The TSG salinity is lower than the CTD salinity by 0.028psu or 0.023psu if only cases with low standard deviation in the TSG data are used. The TSG is lower than loop samples by a median of either 0.03 or 0.05psu depending on whether large differences were excluded or not. The TSG salinity is lower than rosette salinity by a median of 0.039psu but there were only 7 bottles in that comparison.

6. There are many single-point one-sided spikes towards lower salinity values and these are likely due to bubbles. While the comparison between CTD and TSG is likely affected by these spikes, they would not affect all the data, yet in almost all cases the TSG salinity is lower than the CTD. This could mean that there is calibration drift, but a large drift is unlikely given a recent factory calibration. More likely there are more bubbles than are obvious and/or there are slight mismatches in the depths from which the TSG draws water and that from which the CTD data was gathered.

7. The TSG salinity was lower than the loop salinity by 0.015 when the ship was stopped and by 0.027psu while underway. There were fewer TSG salinity spikes when stopped.

8. No recalibration of salinity is justified since the comparisons are not consistent enough to lead to a reliable adjustment. A warning will be put in the header that values are generally low.

9. The TSG fluorescence was fairly close to CHL loop samples when the ship was underway (factor ~0.9) and somewhat higher when stopped (factor of ~1.8), even though the median loop CHL values are close. Varied flow in the loop might explain this but the flow meter shows no change. Could bubbles be affecting fluorescence? The rising of fluorescence during a ship stop sometimes seems visible in the fluorescence traces, though noise and local variability complicate interpretation. Whatever the cause, the underway fluorescence appears to be more accurate than the stopped values for this cruise.

10. The TSG fluorescence values are fairly close to the CTD fluorescence and rosette samples, though they are larger at high values of CHL. The range of TSG fluorescence is small.

f.) Editing

Each file was opened in CTDEDIT but it was clear that editing salinity was not practical with so many spikes and frequent gradual drops in salinity followed by sudden increases that make it unclear which values are poor. No editing was applied.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number,

Temperature:Difference, Conductivity:Primary, Flag and Position:New channels.

File #1 needed slightly different treatment because in patching it together, different channel names had been inserted.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to standardize channel names and formats.

Those files were saved as TOB files.

The Standards Check and Header Check were run; no problems were found.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

##### Loop File

The Chief Scientist provided file 2017-06 Loop log.xlsx which included event numbers, sample numbers and what was sampled. Times and dates were added to this file based on the log entries, selecting the end time of samples taken during rosette casts.

The loop data provided by analysts were combined with the information from the loop log and saved as 2017-06 Loop_Data.xlsx. There were a few problems lining up the data samples.

- One salinity sample, identified as loop sample #456, is said to be from P20, but when it was taken is unknown. It could be from event #70 where there are CHL samples, but those don’t use that sample #. There is no record of any sampling with that sample #.

- There were 2 loop DO samples with sample #6008, one at P1, the other after leaving P1, probably soon afterwards. These data are not to be archived, so this is not critical.

- The salinity and nutrient loop samples #115 were misidentified as being rosette samples. There was no CTD event 115. There was a loop but only gases were collected there. The analysts were informed. The nutrients and salinity were moved to 116 which was a go-flo cast.

- There are a number of CHL samples listed on the Loop log but missing from the CHL spreadsheet (casts 36, 46, 64 and 85). In each case there was another cast at the same site with CHL sampling so there was probably a decision to not gather the sample.

- For event 82 there are salinity and CHL samples reported that are not in the loop log.

The CHE files were put through program DERIVE to obtain sigma-t.

Data from those files were exported to file 2017-06-che-surface.csv.

Because there are no real loop DO samples (only some from the transducer room), the DO channel in mass units was included in the spreadsheet.

Data from below 7m were removed.

A sample method column was added. ROS was entered for the method.

The Start Time was copied into a second column and the first was formatted for date and the second for time. Columns were rearranged to fit a model 6-line header.

Times were corrected for rosette samples to match the end of casts.

The data were sorted on event number, then pressure and added to the 6-line header file.

Loop data had been prepared earlier for use in comparisons with the TSG data. That included adding date/time based on log entries. The data were added to the spreadsheet and saved as 2017-06-surface-6linehdr.csv.

The loop data were aligned with the rosette columns.

The sampling method column was entered as USW.

The test version was sorted on event/date/time.

The file break column was filled with value 1 so all data will be in a single file when converted.

The file was sorted on date, time, sample method and pressure.

(Note: For future reference on another occasion this step failed several times and no cause could be found. Closing and reopening IOS SHELL and doing a reboot did not help. Closing all EXCEL files and opening and closing some convert spreadsheet routines in other projects but not actually running them eventually worked – the program conversion program seemed to remember an input file that had since been fixed.)

CONVERT was run to produce an IOS Header file.

CLEAN was run to get start and stop times and positions and to add flag 0 to empty flag cells.

A comment file was prepared which was essentially the same as the one used in preparing CHE files but including a description of the loop system and comments on the CTD data processing.

Header Edit was used to correct channel names and formats and to add comments. The final file was renamed as 2017-06-surface.loop. The track plots looks reasonable and a plot of temperature and salinity versus longitude looks reasonable.

Particulars – from log book and Chief Scientist’s sampling notes

PAR OFF: 36, 46, 51, 64 and 85

Out of Order firing: 1, 4, 12, 21, 22, 34, 36, 41, 44, 51, 54, 62, 70, 87 and 92.

Casts with no bottles closed: 2, 96, 130 and 134

Deployment Method – Rosette brought to surface, pumps turned on, rosette taken to 10m and back up. Start archiving and wait 30s, then cast run.

“PAR” casts include initial data archived with the rosette half out of water to record the PAR value just out of water and just below the surface. This includes casts 9, 45, 57, 87 and perhaps 8, 12, 41 and 56.

CTD:

8. Wrong file name. This was corrected in RAW files.

12. Spikes in transmissometer both down and up

16. Wrong depth in header. This was corrected in RAW files.

20. File started at 0913 but archiving started at 10:01, so time needs to be adjusted. The START time and latitude and longitude were adjusted in the IOS and ROS files to match log entries..

20. PAR spiky (fuzzy) ~350 to bottom, fluorescence a little spiky too. Secondary T-C pair a little weird on way down in first 200db.

24. Wrong depth in header – This was corrected in RAW files.

42. BL and XMLCON files but no actual cast. This was a DRF event.

51. Niskin 3 closed by mistake at 38m, not needed. All other samples taken from the next Niskin.

74. PAR very fuzzy at depth.

85, 87, 92 – wrong depth in files – This was corrected in RAW files.

100. No bottle file needed. Not processed beyond ROS stage

104. Niskins #4 to 24 not needed.

TSG:

File 1. Start in Haro Strait. Adjusted the flow down ( it was at ~1.25) a little after the start of the file.

The server for the TSG lost GPS signal at 0327, 6 June 2017. The position will have to be matched by time from 0327 to 1451.

File 2. Started after a reboot of the server.

File 3. The flow meter stopped sometime on the 16th of June. The other variables do not indicate any problem with the flow itself.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 22Feb2017 | Factory |  |  |
| Conductivity | 1766 | 30Mar2017 | Factory |  |  |
| Secondary Temp. | 4484 | 22Feb2017 | Factory |  |  |
| Secondary Cond. | 3531 | 30Mar2017 | Factory |  |  |
| Transmissometer | 1201DR | 27Apr2016 | Factory |  |  |
| SBE 43 DO sensor | 1119 | 28Mar2017 | Factory |  |  |
| PAR | 70613 | 21Mar2016 | Factory |  |  |
| SeaPoint Fluor. | 3685 | 13Mar2017 |  |  |  |
| Pressure Sensor | 0550 | 24Feb2017 | Factory |  |  |
| Altimeter | 62355 | n/a | Factory |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3411       Cruise ID#:	2017-06

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 4Mar17 | Factory |  |  |
| Conductivity | 2488 | 4Mar17 | Factory |  |  |
| WetLabs Fluorometer | Ws3s-953p | May2017 |  |  |  |
| Temperature:Secondary | ? | 0842 |  |  |  |
| Flow meter | ? | n/a |  |  |  |