## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 20 March 2025 | ### Updated channel names & formats in TOB & loop files. GG & SH |
| ### 25 July 2024 | ### Added Cesium data to 10 casts. SH |
| ### 1 June 2023 | ### Corrected flags/comments in CHE files casts #23,53,61,67,69,76. |
| ### 6 Dec 2021 | ### Corrected Salinity:Bottle precision that had been lost during HPLC addition. Re-ran addition of total nitrates to files 33, 67, 71 & 74. Made sure DMSP changes from 1 April 2021 were made to 19, 43, 61, 80, 97. S.H. |
| 14 May 2021 | Added DOC/TOC/TDN/TN to 10 CHE files. S.H. |
| ### 1 April 2021 | ### Clean up to DMSP – files had been corrected after first addition, not all values are padded. Changes to flags and comments. SH |
| ### 2 Sept 2020 | ### Added HPLC data. S.H. |
| 6 Feb 2020 | Changes to flags/comments for DMSP G.G. |
| 27 Nov 2019 | Added DMSP – all values padded due to problem with blanks G.G. |
| 6 Feb 2017 | Corrected header comments in underway TOB files G.G. |

## PROCESSING NOTES

Cruise: 2017-08

Agency: OSD

Location: North-East Pacific

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: 15 August 2017 – 30 August 2017

Processed by: Germaine Gatien

Date of Processing: 11 December 2017 – 31 January 2018

Number of original HEX files:  	49	Number of CTD files: 49

Number of bottle files: 		49	Number of bottle casts processed: 48

Number of original TSG files: 	3	Number of processed TSG files:	 3

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1201DR), a SBE 43 DO sensor (#1119) on the primary pump, a SeaPoint Fluorometer (#3685) on the secondary pump, a Biospherical QSP-400 PAR sensor (#70613), and an altimeter (#62355).

A thermosalinograph (Seacat 21 S/N 2488) was mounted with a Wet Labs WETstar fluorometer (S/N ws-3s 953p), remote temperature sensor and a flow meter.

Seasave version 7.26.7.107 was used for acquisition.

The data logging computer was the Tully CTD Laptop.

The deck unit was a Seabird model 11+, serial number 0425.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial #58879.

The oxygen kit was SIO (nsB-6009) model 2.36a kit #2.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

Note that as of January 2017 a change has been made to the threshold levels for quality flags 3and 4 for Extracted Chlorophyll. Flag=3 will now be 10%CV (formerly 15%) and Flag=4 will now be 30% (formerly 50%).

The Daily Science Log, rosette log sheets and analysis logs were in good order with detailed comments on problems. Notes from the Chief scientist were very helpful as was the log of loop sampling.

Both CTD salinity channels had spikes. The primary temperature/salinity pair was selected for archiving because the primary salinity showed little variation with pressure or time. The secondary salinity appears to have some pressure dependence and temporal drift.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples, especially when far from shore. When the data are divided geographically it is seen that for CHL values <2ug/L, the CTD fluorescence is fairly close to the CHL when nearer to shore but much higher well offshore.

The dark value of the fluorometer started a little higher than usual and gradually drifted to values of about 0.2ug/L. During the cruise that followed this drift continued until the fluorometer tubing and the secondary pump were cleaned. A time-dependent correction was applied based on comparisons of dark values at 4000m, with a small adjustment (+0.06ug/L) to ensure there are no negative fluorescence values.

The transmissivity profiles and values look reasonable but as noted for other recent uses of that sensor the values get stuck for a half-second or so, then register a few new values with the usual sampling rate then get stuck again.

The CTD dissolved oxygen correction based on comparison to bottles is higher than found during cruise 2017-06 in June; that cruise had extensive sampling in waters where Niskin bottles are likely to have flushed quite well.  However, cruise 2017-23 that was run between the two Line P cruises also showed a similar correction. 2017-23 included a variety of conditions including some deep casts where flushing should be good.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5 mL/L from 0 to 200db

±0.15 mL/L from 200db to 500db

±0.04 mL/L below 2500db

Overall the TSG performed well with intake temperature available throughout the cruise, but the salinity data were plagued with spikes. Channel Salinity:T0:C0 is given with only 3 decimal places. While comparisons with CTD data and rosette samples indicate that TSG salinity values were low by between 0.3psu and 0.5psu, recalibration is not justified due to great variability in the comparisons and in the spikiness of the traces.

The TSG fluorometer readings were generally close to the CTD fluorometer. TSG fluorescence read higher than chlorophyll samples from the loop and rosette for low chlorophyll, but as CHL rose the ratio gradually fell.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Sampling notes from the chief scientist noted a number of issues. Based on these notes a few corrections were made in the headers of the files before conversion.

Nutrients, extracted chlorophyll, dissolved oxygen, DMS and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked. They had all been used on 7 previous cruises.

The configuration stayed the same throughout the cruise except for corrections made to the PAR parameters after event #2.

Two changes were made to the configuration file for cast #1 and the file was saved as 2017-08-ctd.xmlcon:

- The calibration constant and offset were wrong for the PAR sensor for the first few casts; the correct values were entered.

- Factor E in the dissolved oxygen sensor calibration was changed from 0.036 to 0.038 to improve the hysteresis correction based on the tests done using data from cruise 2017-06. (Note: This was revised after running COMPARE and value 0.0375 was used instead.)

##### Conversion of Full Files from Raw Data

All files were converted using 2017-08-ctd.xmlcon.

The hysteresis and Tau functions were selected. Depth was included in the conversion.

A few casts were examined and all expected channels are present.

- As usual the temperature and conductivity channels track fairly well during downcasts but are noisier on upcasts. There are spikes in both primary and secondary channels. Differences are notable in high gradients but not elsewhere, except where there are notable spikes.

- PAR, Fluorescence, Dissolved Oxygen and Altimetry have normal profiles.

- As seen in other recent cruises using this transmissometer there are odd profiles that look as though values got stuck for a half-second or so, then registered a few new values with the usual sampling rate then got stuck again. The profiles look reasonable.

Acquisition for a few casts was started right at the surface before the 10m soak. During casts #2 and 19 the pressure was about 0.2db at the beginning of the file and for cast #42 acquisition started at about 0.4db. This was likely just as the CTD part of the package entered the water, so we would expect PAR to go down when it is at about 1.5db.  PAR goes down rapidly when the CTD reaches about 1.3db. So the surface reading is likely within 0.2db. The pumps were off at that point.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2017-08-ctd.xmlcon. Depth was included.

Cast #1 had no sampling so will not be processed further.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. Outliers were found in the primary salinity in 3 casts:

- Cast #16 - CTDEDIT was used to remove all points in channel Salinity:T0:C0 from Niskins #12 to 14 at 6db.

- Cast #88 - CTDEDIT was used to clean Salinity:T0:C0 for a few points from Niskin #6.

- Cast #94 - CTDEDIT was used to clean Salinity:T0:C0 for a few points from Niskin #11. Other parts of this cast have suspicious primary salinity, but it is not obvious what if any editing should be applied.

The output files from CTDEDIT were copied to the BOT files.

A preliminary header check and no problems were found.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The unwanted line for Niskin bottle 1 was dropped from file 2016-06-0072.SAMAVG.

The unwanted line for Niskin bottle 3 was dropped from file 2016-06-0080.SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2017-08-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2017-08chl*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2017-08chl.csv. Loop data were saved in a separate file. The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2017-08oxy.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and loop data were moved to a separate file. The rosette file was saved as 2017-08oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in file 2017-08SAL.xlsx which included a precision study. The analyses were carried out in a temperature-controlled lab 14 to 28 days after collection.

The files were simplified and saved as 2017-08sal.csv. File 2017-08sal.csv was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2017-08_nutrients*.xlsx. This includes a precision study. The file was simplified and the saved as 2017-08-nuts.csv. The file was converted to individual NUT files.

DMS

DMS data were obtained in file DMS summary (2017-08).xls. Values given as < were changed to 0 and those given as – were replaced with pad values; the comments that will go into the header will explain that 0 means below minimum detectable level which is 0.1nM. The file was then saved as 2017-08DMS.csv and converted to individual DMS files. There was a separate report on analysis techniques and problems.

The SAL, CHL, OXY, NUT and DMS files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. This step is important to avoid missing some samples in the creation of CHE files.

The only problem found was in the extracted CHL with some flags and comments, not getting transferred from the CV% page to the QF CHL profile page.

There are loop data in the salinity, nutrient, chlorophyll and oxygen spreadsheets, so those data were moved to file 2017-08 Loop_Data.xlsx.

The chief scientist provided a log of loop sampling. There were loops taken while underway and others were taken at the end of CTD casts while the 5m rosette was being fired.

Times were added to the loop file based on the end time of CTD casts or the loop time if there was no CTD cast at the time.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There was 1 major outlier. Sample #29 from cast #11 looks like the sample was either taken from Niskin 1 instead of the plan to take it from Niskin #3, or there was a problem in analysis. This was also an outlier in dissolved oxygen but the DO sample does not look like it came from anywhere near the bottom of the cast.  So, the problem doesn’t appear to be due to a premature closure. The analyst added a 3 flag and comment. The merge process was repeated for cast #11.

The primary salinity is low by an average of 0.0002 when bottles above 500m are excluded.

The secondary salinity is high by an average of 0.0004psu using the same points.

Excluding points with a standard deviation in the CTD salinity of >0.0008 removed no further points.

Plots against time show a drift from beginning to end that has the primary dropping by 0.0006psu and the secondary by 0.0012psu. The drift in the secondary brings it closer to the primary. The last really good calibration for these sensors showed the primary to be low by 0.0012psu and the secondary to be high by 0.0032psu. This may suggest that the drift in the primary is not significant, but that in the secondary probably is.

In summary:

|  | Average Difference | Standard Deviation | Fit against Pressure | Fit against Time |
| --- | --- | --- | --- | --- |
| Sal0 – Sal Bottle | -0.0002 | 0.0015 | Diff = -2E-08*Press - 0.0002 | Diff = -2E-05*File Pair # - 0.0002 |
| Sal1 – Sal Bottle | +0.0004 | 0.0015 | Diff = -5E-07*Press + 0.0014 | Diff = -5E-07*File Pair # + 0.0013 |

File Pair # is the consecutive # of rosette casts and ranged from 1 to 35 for this cruise.

The study of 23 bottles fired at 2000db during cast #69 found a standard deviations of 0.0013psu in the differences between bottles and both salinity channels which is a little higher than usual. During 2017-06 it was 0.0008psu. There is no obvious outlier to remove, just a general scatter. The primary was found to be high by 0.0004psu while the secondary was high by 0.0010psu.

For full details for the COMPARE run see file 2017-08-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

Tests were done during 2017-06 to see if the default value E=0.036 in the dissolved oxygen sensor was the best setting. A setting of E=0.038 was found to work better. While the ideal E value is not expected to vary significantly between factory calibrations,  a check was made for hysteresis by excluding all points from below 2000m and it looks as though E=0.038 is not the best choice since the deep values stand out somewhat from the shallow in a plot of differences versus DO concentration. Tests were run using a variety of settings. E=0.036 proved poor, but using 0.0375 did a good job of removing hysteresis.

The steps described in section 4 were repeated using the updated configuration file.

COMPARE was rerun.

First, outliers were considered. A number of samples had been flagged by the analyst, of which 2 were flagged 4:

- Sample 753 from event 99 had been flagged 4 and that looks appropriate – it is out of line but is not an extreme outlier.

- Sample 178 from event 33 and is slightly out of line, so leaving the 4 flag looks appropriate.

It is noted that there are many outliers that are as bad as or worse than the 2 above.

They are:

- Event 2, sample 13, 5m – significant outlier and CTD data not noisy. Fairly high local gradient, so bottle may contain water of lower DO due to poor flushing. Analysis is likely ok.

- Event 7, sample 21, 40m - significant outlier and CTD data only slightly noisy, though CTD DO did rise just before firing and the local vertical gradient is high, so analysis likely ok.

- Event 11, sample 29, 5m – near surface - CTD data standard deviation is high. There is a subsurface peak in DO and a very high vertical DO gradient. Poor flushing of the Niskin bottle likely explains this. The analysis result may be fine, representing what was in the bottle. The analyst added a 3 flag and comment.

- Event 47, sample 278, 2000m – significant outlier reading lower than expected – below DO min so poor flushing should lead to reading higher. CTD std dev ok. Other bottles and draw temperature look ok in profile. Local vertical gradient not especially high. The analyst decided to pad the value, add a 4 flag and comment.

- Event 47, sample 287, 200m – significant outlier reading higher than expected – above DO min so expect if poor flushing would read too low. CTD std dev ok. No obvious explanation – slight DO reversal on downcast at this level, but not large enough to explain this and the local gradient is not particularly large. Other bottle values and draw temperature are not out of line in profile. The analyst added a 4 flag and comment.

There are a number of other bottles flagged from cast #47 but not these two. The flagged samples are not particularly out of line. After corrections the merge process was repeated.

When outliers were removed based on residuals, the fit was:

CTD DO Corrected = CTD DO * 1.0514 + 0.0533 (R2 = 0.96)

This correction is much higher than found during cruise 2017-06 in June; that cruise had extensive sampling in waters where Niskin bottles are likely to have flushed quite well.

CTD DO Corrected = CTD DO * 1.0399 + 0.0645 (R2 = 0.96) 2017-06

Cruise 2017-23 in July had fewer deep water samples. The correction applied was;

CTD DO Corrected = CTD DO * 1.054 + 0.048 (R2 = 0.97) 2017-23

While the sampling from the June cruise was more extensive, there was enough deep sampling during 2017-23 to offer confidence that the calibration has drifted.

The other cruises that have used this equipment in 2017 were mostly in shallow, protected waters where incomplete flushing is likely to limit the reliability of a CTD-bottle comparison.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. An error was found in the MRG file for cast #72 – the bottle data had been linked with the wrong Niskin bottle.

The ratio of FL/CHL was calculated and there was one very high ratio. Investigation showed that the dark values for the fluorometer were quite high, starting at ~0.08ug/L and gradually drifting to >0.2ug/L by the end of the cruise. If a value of 0.15 is subtracted from all the fluorometer readings, there is one value <0. If  0.1ug/L is subtracted and used to look at the FL/CHL ratio, there are fewer very high ratios, but the general shape is the one we usually see, with the fluorometer reading higher than CHL at low CHL, close to CHL when CHL~1 and lower for CHL >1. There are only a few CHL readings above 5ug/L. The scatter is larger than usual but the usual pattern is seen.

In the fit of CTD fluorescence versus CHL there appear to be 2 different fits for CHL<2ug/L. When the data are divided geographically it is obvious that the CTD fluorescence is closer to the CHL values when  closer to shore, but much lower well offshore even when CHL values are similar.

For full details of the comparison see file 2017-08-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run during 2017-05 when the same equipment was in use and the best results were found using a setting of +2.5s. For the next 4 uses of this sensor the same setting looked appropriate. A few casts were examined from this cruise; there were many stops for bottles and both temperature channels were noisy so the tests were not easy to interpret, but overall +2.5s looks like a good choice.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice. One cast was checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. Some results from other cruises that used these sensors are included for comparison; those entries are shaded.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2017-05-0038 | 1000 | +0.0001 | +0.0004 | +0.0046 | High, X Noisy |
|  | 1900 | +0.0003 | +0.0004 | +0.0040 |  |
| 2017-05-0101 | 1000 | +0.0006 | +0.0005 | +0.0055 | High, Moderate |
|  | 1900 | +0.0004 | +0.0005 | +0.0052 | “ |
| 2017-05-0139 | 1000 | +0.0002 | +0.0005 | +0.0051 | High, X Noisy |
|  | 1900 | +0.0004 | +0.0005 | +0.0055 | “ |
| 2017-06-0036 | 1000 | +0.0002 | +0.0004 | +0.0047 | High, X.Noisy |
|  | 1900 | +0.0004 | +0.0004 | +0.0045 | “ |
|  | 2900 | +0.0003 | +0.0004 | +0.0042 | “ |
| 2017-06-0064 | 1000 | ~0 | +0.0004 | +0.0047 | High, Noisy |
|  | 1900 | +0.0003 | +0.0004 | +0.0042 | “ |
|  | 2900 | +0.0004 | +0.0003 | +0.0038 | “ |
| 2017-06-0085 | 1000 | +0.0002 | +0.0004 | +0.0043 | High, V Noisy |
|  | 1900 | +0.0002 | +0.0003 | +0.0040 | “ |
|  | 2900 | +0.0002 | +0.0003 | +0.0036 | “ |
| 2017-08-0047 | 1000 | +0.0002 | +0.00007 | +0.0009 | High, V Noisy |
|  | 1900 | +0.0002 | +0.00005 | +0.0005 | “ |
|  | 2900 | +0.0004 | +0.00004 | -0.0001 | “ |
| 2017-08-0059 | 1000 | -0.0002 | +0.00007 | +0.0009 | High, V Noisy |
|  | 1900 | +0.0000 | +0.00005 | +0.0004 | “ |
|  | 2900 | +0.0003 | +0.00002 | -0.0001 | “ |
| 2017-08-0074 | 1000 | -0.0002 | +0.00008 | +0.0011 | High, V Noisy |
|  | 1900 | +0.0000 | +0.00005 | +0.0005 | “ |
|  | 2900 | +0.0003 | +0.00002 | +0.0001 | “ |
| 2017-08-0082 | 1000 | -0.0001 | +0.00009 | +0.0011 | High, V Noisy |
|  | 1900 | +0.0002 | +0.00005 | +0.0006 | “ |
|  | 2900 | +0.0002 | +0.00001 | +0.0000 | “ |

The temperature differences are small and similar to other recent cruises, but show more depth dependence than noted during 2017-06. Conductivity and salinity are closer than during previous cruises and show similar depth dependence but there is some temporal drift.  The salinity differences are in agreement with the results of the bottle comparison which found that secondary salinity was higher than primary by +0.00006psu. In the comparison the pressure dependence was clearly in the secondary sensors. There is a larger time-dependence in the secondary salinity channel than the primary so that while CTD salinity is dropping in both channels they are getting closer together.

It is unfortunate that the secondary channel looks to be drifting since the primary salinity is so noisy.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

CLIP was used to remove initial records acquired during the 10m soak from casts 2, 19 and 43.

The CLIP files were copied to CLN.

##### Checking Headers

A cross-reference list was checked against the log book. No errors were found.

The cruise track was plotted and added to the end of this report.

A surface check was run and shows an average surface pressure for the cruise was 2.6db. This is fairly typical for the Tully when offshore, but there is a wide range of values. Cast #17 started acquisition at ~6db. There is no note in the log of a problem. A few had very low surface pressures but those are casts where acquisition was started before the initial 10db soak.

The header check showed that there were spikes in pressure for at least one cast but that proved to be at the end of the cast. No other problems were noted and fluorescence did not go off-scale.

The altimeter and water depth readings from the headers of the CLN and SAMAVG files were exported to a spreadsheet. Water depths were compared with the log book entries. A few changes were made to the headers:

- The depth entries were adjusted for casts #23, 59 and 66.

- The altimetry headers in casts 26, 38 and 80 were caused by spikes; the CTD did not get near the bottom so the headers were removed.

- For the bottle files, the same changes were made as above except that only casts # 59 and #66 had no altimetry header so no change was needed.

After correcting the SAM files they were bin averaged again and the final merge step was run for the bottle files; the CLEAN step was repeated as well.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 3 casts to see if the shifts applied to conductivity during the previous cruise when the same equipment was in use were appropriate and they were. As found during 2017-05 there is a lot of noise in the data and no alignment will remove it all.

SHIFT was run twice on all casts using -0.6 records for the primary conductivity and -0.6 records the secondary. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were two warnings. One came from the upcast of file #7 is of no concern. The other is more significant with a large spiky section early in the cast which made DELETE skip the top 5m. The first 600 records were removed from the SHFC1 file and DELETE was rerun with no warnings.

##### Other Comparisons

Previous experience with these sensors –

All sensors were used during 7 earlier cruises in 2017.

2017-63 in early April was a cruise in the Strait of Georgia and Juan de Fuca Strait. The bottles from Juan de Fuca were trusted more than those from the Strait of Georgia because conditions were better for flushing Niskin bottles; they indicated that the primary salinity was close to bottles while the secondary was high by about 0.0045. The dissolved oxygen comparison from 2017-63 was also based on Juan de Fuca bottles.

Cruise 2017-03 was in Baynes Sound and the Strait of Georgia, with shallow sampling and poor flushing of bottles, so the results were not considered reliable enough to use for recalibration.

Cruise 2017-05 was mostly off the west coast of Vancouver Island where flushing was good at depth. The comparison between the bottles and CTD salinity was quite tight; the primary salinity was found to be very close to bottles and the secondary was high by 0.0054psu. The dissolved oxygen comparison was also quite tight and led to a calibration of slope 1.0448 and offset +0.0527.
Cruise 2017-06 was mostly well offshore and showed the primary salinity to be low by ~0.0012 and the secondary high by ~0.0032. The dissolved oxygen was recalibrated using slope 1.0399 and offset 0.0645.

Cruise 2017-64 showed primary and secondary salinity low by 0.0012psu and high by 0.0034psu  in Juan de Fuca Strait. The dissolved oxygen comparison in Juan de Fuca was very close to that of 2017-06, so those results were used for 2017-64.

Cruise 2017-56 included only near-shore shallow casts so 2017-06 results were used for recalibration.

Cruise 2017-23 comparisons were not trusted as much as the 2017-06 results, so the latter were used for recalibration.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Salinity was above the historic maximum at P4 between 60 and 80db (during both outward and inward visits) and slightly high at P20 and P21 at about 130db. All other salinity values were well within the historic range.  All temperature data fell within the climatology..

Repeat Casts – There were repeat casts. Casts #82 and #89 occurred about 10 hours apart at P26. Differences along lines of constant σt were <0.002º for the both temperature channels and <0.0002psu for both salinity channels at σt= 27.51 (near 1400m). This shows excellent repeatability.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

Both CTD salinity channels had a lot of spikes and near-surface unstable salinity excursions that appear to be due to poor alignment of T and C in high gradients. Noisy descent rates account for much of this corruption due to either shed wakes or variability in alignment. For most casts the primary temperature/salinity pair was selected for archiving because the primary salinity showed little variation with pressure or time. However, for cast #76 the primary channels were extremely noisy. While most unstable features were clearly due to shed wakes that affected the secondary channels as well, the effects were less severe in the secondary than in the primary. So the secondary channels were chosen for archiving and editing.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. Some bad salinity points were removed from a few files. All files required some editing. The descent rate was generally noisy for the offshore casts and for some casts there were many complete reversals of direction so there was a lot of corruption by shed wakes.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts; further editing was applied to 2 files.

##### Recalibration

Pressure appears to be within ±0.2db based on observations during casts where acquisition started before the 10m soak. So recalibration of pressure is not required.

The primary salinity was found to be lower than the bottle salinity by an average of 0.0012psu. Since slight inefficiency in Niskin flushing could account for bottle values being slightly high, this is not considered a significant difference and that channel will not be recalibrated.

The secondary salinity was used for one cast. To bring it into correspondence with the primary salinity 0.0006psu will be subtracted from that channel. This does not account for the differences in pressure dependence, but that error is small.

The fluorescence dark values are larger than expected and increase through the cruise. Values found in June using the same fluorometer were ~0.04ug/L at 4000m at station P26. For this cruise the values at the same site were ~0.20ug/L. The drift in values is obvious from about 250m down but is hard to detect above that due to natural variability and internal waves. A time-dependent correction may help, but Julian time is needed to do that, so the ADD TIME channel was used to add that.

Plots were made of dark values versus cast start times to see if the drift was approximately linear. Looking at 2000m data showed that the drift was somewhat faster earlier in the cruise than later, though the minimum step in these values is ~0.013ug/L, so it is easy to over interpret variability. There are no 4000m values until cast #71 onwards. The fit based on 4000m data was:

Fluorescence Corrected = Fluorescence + 1.442 - 0.0068* Cast Time (Julian Days)

The fits will lead to some negative fluorescence values which are not reasonable. The minimum value found in the edited files is 0.06ug/L. So if we add 0.06 to the fit we should get no negative values and results that look reasonable, though dark values will be slightly higher than normally expected.

Dissolved oxygen recalibration was discussed in section 5.

File 2017-08-recal1.ccf was prepared to subtract 0.0006psu from the secondary salinity, to apply the following corrections to channels Oxygen:Dissolved:SBE and Fluorescence:URU:Seapoint:

CTD DO Corrected = CTD DO * 1.0514 + 0.0533

Fluorescence Corrected = Fluorescence + 1.502 - 0.0068* Cast Time (Julian Days)

This correction was first applied to the SAMATC and MRGCLN2 files.

COMPARE was rerun for salinity and dissolved oxygen using roughly the same points as in the fit used for recalibration to ensure that corrections were appropriate and they were. The secondary salinity was found to be low by an average of 0.0002, which is reasonably close to the primary.

See file 2017-08-sal-comp2.xlsx for details.

The average of differences in the DO fit was +0.0002mL/L and the standard deviation was 0.021mL/L. See file 2017-08-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files using the same recalibration file.

The fluorescence profiles confirm that the dark values are more appropriate after this step. For casts #23 and #111 (both at station P4), the values at 1300m differ by about 0.11ug/L before the correction and by about 0.03ug/L after recalibration. The near-surface maximum values are also closer by ~0.08ug/L.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When outliers were removed based on residuals the CTD DO was higher than the titrated samples by an average of ~0.0035mL/L (standard deviation of 0.006mL/L). Looking at the differences versus pressure shows that the sensor DO tends to be slightly on the high side above 500m which is likely due to bottles reading a little low due to incomplete flushing. The sensor also appears to be slightly high below 3500m, which may be a sign of a little hysteresis in the sensor data, though with only a few points of comparison and differences <0.02mL/L, this does not look significant.

No further recalibration is justified. See 2017-08-dox-comp3.xlsx for details.

##### Fluorescence Processing and special files for Dr. Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases. The SAM files were put through REMOVE and HEADEDIT. Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts except #76 REMOVE was run to remove the following channels:

Time:Day_Of_Year, Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

For cast #76 REMOVE was run to remove the following channels:

Time:Day_Of_Year, Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

PAR was removed from casts 47, 50, 52, 53, 56, 57, 59, 67, 69, 71, 74, 82, 99.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

ADD TIME CHANNEL was run to subtract 7 hours that were accidentally added when Julian time channel was added.

HEADER EDIT was used to fix formats and channel names, to remove the header entry TIME UNITS, fix the vessel name and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and unedited except that some

records were removed in editing temperature and salinity.

The fluorometer values drifted during the cruise starting with slightly high dark

values early in the cruise and rising to about 0.20 ug/L by the end. A

time-dependent correction was applied based on values at 4000m with an

adjustment to ensure no negative values were produced. This is an

estimate and produces dark values at 2000m ranging from 0.02ug/L to 0.09ug/L.

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5 mL/L from 0 to 200db

±0.15 mL/L from 200db to 500db

±0.04 mL/L below 500db

For details on the processing see document: 2017-08_Processing_Report.doc.

±0.5 mL/L from 0 to 200db

±0.15 mL/L from 200db to 500db

±0.04 mL/L below 2500db

For details on the processing see document: 2017-08_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found.

A cross-reference list was produced.

The sensor history was updated.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The values ranged from ~75% to 115% with the highest value in Saanich Inlet and the lowest in Haro Strait. All values from P7 to P26 were between 100% and 104%. The results are consistent with good dissolved oxygen calibration since we usually see values ~103% offshore.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For all casts except #76 REMOVE was run to remove the following channels:

Time:Day_Of_Year, Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

For cast #76 REMOVE was run to remove the following channels:

Time:Day_Of_Year, Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

PAR was removed from casts 47, 50, 52, 53, 56, 57, 59, 67, 69, 71, 74, 82, 99.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

ADD TIME CHANNEL was run to subtract 7 hours that were accidentally added when Julian time channel was added.

HEADER EDIT was run to ensure formats and units are correct, correct the vessel name, remove the header entry TIME UNITS, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2017-08-bottles-final.xlsx. The entries were compared with the rosette log sheets to ensure no samples had been missed. Only one problem was found and corrected:

- Event 61: The DMS value and comment for sample #442 appeared with sample #441 in the merged files because there was no sample 442 in the SAMAVG file. This was corrected in the all files beyond the MRG4 stage. A comment was left in the files to say that the Niskin bottle did not close for sample #442.

Standards check and a header check were run on all files. No problems were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files.

##### Thermosalinograph Data

There were 3 hex files.

a.) Checking calibrations

The configuration file did not change through the cruise, but contained an error in the fluorometer parameters. One file was renamed as 2017-08-tsg.xmlcon and the fluorometer calibration parameters were fixed.

b.) Conversion of Files

3 files were converted to CNV files using configuration file 2017-08-tsg.con.

Those CNV files were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots showed that the salinity signal had many one-sided spikes likely to be due to bubbles, and unlikely to be amenable to editing. The most heavily corrupted data come from the latter half of file #2. The flow rate was mostly steady, ranging between 1 and 1.2 except for a few patches with lower values that could affect data quality. At the beginning of the file #1 there was no flow for the first few records; judging by the differences in intake and lab temperatures, the flow was well established after scan #13, so the temperature, salinity and fluorescence should be removed from the first 13 scans.

The temperature differences were very noisy at the beginning and end of the cruise as expected for nearshore observations. Differences were about -0.15 C° nearshore and about -0.20 C° offshore where temperatures were lower.

The TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, lab temperature, salinity and fluorescence and the files were reduced to the times of CTD files. There were 46 cases of overlap.

Similarly, TSG data were found to match the loop samples.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude and longitude were all ≤0.0003° and the median differences were 0.0000° for both. This shows both the times and positions are reliable for both systems.

d.) Comparison of T, S and Fuor from Loop & Rosette Samples and TSG and CTD data

- T1 vs T2 The intake temperature sensor worked throughout the cruise. The differences decrease slightly as temperatures increase as is expected as the intake temperature gets closer to the temperature of the ship.

- Flow Rate The flow rate ranged from 1.0 to 1.2 while the ship was stopped for CTD casts; the median value was 1.08.

- TSG vs CTD The spreadsheet comparing CTD and TSG files was then examined to find the differences between the salinity, fluorescence and temperature from the CTD and the TSG.

| Using all data and median values over 2 minutes for TSG: |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | (TSG intake -CTD) Temp | (TSG lab -CTD) Temp |  | (TSG - CTD) Sal |  | FLtsg/FLctd |
| median | 0.0038 | 0.1789 |  | -0.0481 |  | 0.982 |
| stdev | 0.0908 | 0.1052 |  | 0.0547 |  | 0.190 |
| avg | 0.0276 | 0.2044 |  | -0.0581 |  | 1.029 |
| min | -0.0244 | 0.0695 |  | -0.3661 |  | 0.658 |
| max | 0.4496 | 0.6282 |  | 0.0155 |  | 1.637 |
| Using the 10 cases with lowest standard deviation: |  |  |  |  |  |  |
| (TSG intake -CTD) Temp | (TSG lab -CTD) Temp |  | (TSG - CTD) Sal |  |  |  |
| 0.0034 | 0.1768 |  | -0.0268 |  |  |  |

- The flow rate had no noticeable effect on the differences between lab and intake temperatures.

- The intake temperature is higher than the CTD temperature by a median value of ~0.0038C° with only a slight reduction to 0.0034C° when only the 10 cases with the lowest standard deviation are included.

- The TSG salinity is lower than the CTD salinity by a median value of 0.048psu with a standard deviation of 0.055psu. When only the 10 cases with the lowest standard deviation are included the difference is reduced to 0.027.

- The ratio of the TSG fluorescence to CTD fluorescence has a median value of 0.98 and standard deviation of ~0.19. When only the 10 cases with the lowest standard deviations are included, the median ratio was 0.94.

For more details see 2017-08-ctd4-tsg-comp.xls.

- Loop vs Rosette

Differences were calculated between samples from the loop and from the rosette samples closest to 5m wherever possible.

|  | Loop Samples - Rosette Samples |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | SAL | CHL | OXY | NITRATE | SILICATE | PHOSPHATE |
| Median Diff | -0.0002 | 0.0078 | 0.211 | -0.10 | 0.13 | 0.007 |
| Min | -0.0039 | -0.0166 | 0.191 | -0.09 | -1.80 | -0.010 |
| Max | 0.0009 | 0.0412 | 0.250 | 0.48 | 0.63 | 0.008 |
| Std Dev | 0.0021 | 0.0264 | 0.026 | 0.23 | 0.92 | 0.038 |

There are only 4 points of comparison for salinity but 3 out of the 4 show the two differing by <0.001psu. The 4th loop sample has a difference of 0.004. The 4th was from station P4 where variability tends to be high. The analysis of both samples appears to have been of good quality.

The dissolved oxygen sample values are consistently higher from the loop than from the rosette.

There were only 6 cases where nutrients could be compared; the median differences are small, but there is a lot of variability.

- Loop Bottle - TSG Comparisons

There were 14 salinity loop samples and 15 for extracted chlorophyll.

The TSG salinity was lower than loop samples by a median of 0.035psu with a standard deviation of ~0.12psu. Loops taken while the ship was stopped are low by a median of 0.048 (std dev ~ 0.019) and while moving they are low by a median of -0.021 (std dev ~0.144).  There were 2 significant outliers in salinity, both taken while the ship was moving. If those are excluded the TSG salinity is low by 0.04psu overall and by 0.03psu while underway. The standard deviations are high for the underway sampling.

Extracted chlorophyll values ranged from 0.09 to 4.4 ug/L. The median ratio of TSG fluorescence to CHL varies from 0.5 to 9.5 with a median of 1.9. There is little difference between the ratio of TSG FL to Loop CHL while stopped and while underway. As usual the ratio is largest for low CHL.

- Rosette Bottle – TSG data

A final check was to see how the TSG values compared with the rosette values – only cases with loop samples were examined as the TSG data were readily available from other comparisons. There were only 6 points of comparison for salinity and 7 for fluorescence. The median value shows the TSG salinity to be reading lower by ~0.042psu and the fluorometer to be reading high by a median factor of 1.76.

(See 2017-08-loops-tsg-rosette-comp.xls.)

- Calibration History

The temperature and conductivity sensors were recalibrated in March 2017 and this was the fourth use since then.

2017-05: The temperature in the lab was higher than CTD temperatures by about 0.18Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of ~0.005Cº. The TSG Salinity read lower than the CTD salinity and that appeared to be due to bubbles. The TSG fluorescence was higher than the CTD fluorescence when the latter is <1ug/L and about 77% of the CTD fluorescence for higher values >1ug/L. Comparisons with rosette chlorophyll samples were limited but suggest that the TSG fluorescence was about 70% of chlorophyll values. No recalibration was applied to the TSG data.

2017-06: The temperature in the lab was higher than CTD temperatures by from 0.15Cº to 0.20Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of 0.0038Cº. The TSG Salinity read lower than the CTD salinity and that appeared to be due to bubbles. The TSG fluorescence was very close to the CTD fluorescence. Comparisons with rosette chlorophyll samples were limited but suggest that the TSG fluorescence was about 1.8 times the loop and rosette chlorophyll values. No recalibration was applied to the TSG data.

2017-23: The temperature in the lab was higher than CTD temperatures by from 0.16Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of 0.004Cº. The TSG Salinity read lower than the CTD salinity by ~0.29psu and that appeared to be due to bubbles. The TSG fluorescence was very close to the CTD fluorescence. There were no loop samples. No recalibration was applied to the TSG data.

Conclusions

1. The TSG clock worked well.

2. The TSG flow rate was at the expected level and generally steady.

3. The temperature increases in the loop by a median value of 0.17Cº and increases as intake temperature decreases, as expected.

4. The TSG intake temperature was higher than the CTD temperature by a median of about 0.004Cº.  This is likely as close as we can achieve given the variability at 4m depth and the fact that the TSG may draw temperature from a little higher in the water column and that there could possibly be slight heating right at the intake. No correction will be applied.

5. The TSG salinity is lower than the CTD salinity by a median of 0.048psu or 0.027psu if only cases with low standard deviation in the TSG data are used. The TSG is lower than loop samples by a median of 0.035psu. The TSG salinity is lower than rosette salinity by a median of 0.042psu but there were only 6 bottles in that comparison.

6. There are many single-point one-sided spikes towards lower salinity values and these are likely due to bubbles. While the comparison between CTD and TSG is likely affected by these spikes, they would not affect all the data, yet in all but 1 of 46 cases the TSG salinity is lower than the CTD. This could mean that there is calibration drift, but a large drift is unlikely given a recent factory calibration. More likely there are more bubbles than are obvious and/or there are slight mismatches in the depths from which the TSG draws water and that from which the CTD data was gathered.

7. The TSG salinity was lower than the loop salinity by 0.048 when the ship was stopped and by 0.021psu while underway. This differs from observations from other cruises and from our expectation that there are fewer TSG salinity spikes were seen when stopped. However, there are few data points available and when one outlier is removed the underway difference becomes 0.03psu.  To check if there were really more spikes while stopped than while underway time-series plots were examined and there is no obvious difference between stopped and underway.

8. No recalibration of salinity is justified since the low values are likely due to bubbles, and the comparisons are not consistent enough to lead to a reliable adjustment to correct for that. A warning will be put in the header that values are generally low.

9. The TSG fluorescence data were very close to the CTD fluorescence. They were about 1.9 times the loop chlorophyll while stopped and 1.8 times while underway and 1.8 times the rosette chlorophyll samples. For CHL<1ug/L the fluorescence is higher than CHL. Above that level it is close to or < CHL. This fits the usual pattern of fluorometers reading higher than chlorophyll for low chlorophyll values.

f.) Editing

Each file was opened in CTDEDIT but it was clear that editing salinity was not practical with so many spikes and frequent gradual drops in salinity followed by sudden increases that make it unclear which values are poor.

The salinity, temperature and fluorescence data were removed from the first 13 scans of file #1.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number,

Temperature:Difference, Conductivity:Primary, Flag and Position:New channels.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to standardize channel names and formats. A note was added to warn that salinity values are believed to be mostly reading low. The format for salinity was reduced to 3 decimal places to reflect the poor quality.

Those files were saved as TOB files.

The Standards Check and Header Check were run; no problems were found.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

##### Loop File

The Chief Scientist provided file 2017-08 Loop log.xlsx which included event numbers, sample numbers and what was sampled. Times and dates, latitude and longitude were added to this file based on the log entries, selecting the end time of samples taken during rosette casts.

The loop data provided by analysts were combined with the information from the loop log and saved as 2017-08 Loop_Data.xlsx. There were a few problems lining up the data samples.

- DO sample #5018 does not exist in the loop log.

- DO sample #5041 is found in the loop log for a cesium sample only.

- Loop DO samples came from the sink in the lab and there are draw temperatures to go with them.

The CHE files were put through program DERIVE to obtain sigma-t.

Data from those files were exported to file 2017-08-che-surface.csv.

Because there are loop DO samples the Oxygen:Dissolved channel in mass units was not included in the spreadsheet as that will be calculated later for both bottle and CTD. The Oxygen:Dissolved:SBE channel in mass units was included.

Data from below 7m were removed.

A sample method column was added. ROS was entered for the method.

The Start Time was copied into a second column and the first was formatted for date and the second for time. Columns were rearranged to fit a model 6-line header.

Times were corrected for rosette samples to match the end of casts.

The data were sorted on event number, then pressure and added to the 6-line header file.

Loop data had been prepared earlier for use in comparisons with the TSG data. That included adding date/time based on log entries. The data were added to the spreadsheet and saved as 2017-08-surface-6linehdr.csv.

The loop data were aligned with the rosette columns.

The sampling method column was entered as USW.

The test version was sorted on event/date/time.

A few changes were made to sample numbers and flags after consulting with analysts and chief scientist.

The file break column was filled with value 1 so all data will be in a single file when converted.

The file was sorted on date, time, sample method and pressure.

(Note For future reference: On another occasion this step failed several times and no cause could be found. Closing and reopening IOS SHELL and doing a reboot did not help. Closing all EXCEL files and opening and closing some convert spreadsheet routines in other projects but not actually running them eventually worked – the program conversion program seemed to remember an input file that had since been fixed.)

CONVERT was run to produce an IOS Header file.

CLEAN was run to get start and stop times and positions and to add flag 0 to empty flag cells.

A comment file was prepared which was essentially the same as the one used in preparing CHE files but including a description of the loop system and comments on the CTD data processing.

DERIVED QUANTITIES was run to add sigma-t to the files.

Change Units was run to add Oxygen:Dissolved in mass units. CTD salinity values are available for more bottles than Salinity:Bottle values, so they were selected for the conversion.

Reorder was run to get channel names in the usual order.

Header Edit was used to correct channel names and formats and to add comments. The final file was renamed as 2017-08-surface.loop. The track plots looks reasonable and a plot of temperature and salinity versus longitude looks reasonable.

Particulars – from log book and Chief Scientist’s sampling notes

PAR OFF: 47, 50, 52, 53, 56, 57, 59, 67, 69, 71, 74, 82 and 99.

Out of Order firing: 80.

Deployment Method – For all but 3 casts: Rosette brought to surface, pumps turned on, rosette taken to 10m and back up. Start archiving and wait 30s, then cast run.  For casts 2, 19, 42 archiving started before the 10m soak.

CTD:

1. SI – test cast only- NO FILE NEEDED.

2. Archive started before soak. Small negative pressure spike before soak.

11. Use secondary channels for bottle file. Downcast ok.

19. Archive started before soak. Negative pressure spike before 10m soak. Went up to 270, then returned to 300 to close Niskin.

42/43. – Archive started before soak. Originally called event 42, should be called 43- file names corrected before conversion.

61. Niskin 17 at surface didn’t get closed. No DMS sample there so no s/n 442.

71. Niskin 3 closed at wrong depth 2500 instead of 3000.

72. Niskin 1 closed by mistake at 2005, not needed in CHE file.

76 – At surface long time, many negative pressure spikes during that time. Used secondary sensors for CTD and CHE since primary too spiky.

80.  Niskin 3 not needed in CHE file – closed by mistake. Bottles out of order. Used Niskin 22 for 200db sample. Remove line for Niskin 3.

91 – Station name should be PA-011, not PA-11; corrected before conversion.

111. Salinity samples 778-780 not collected

TSG: Files quite spiky, mainly in salinity. Piece of copper tubing still on TSG just before fluorometer.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 16Dec2017 | Factory |  |  |
| Conductivity | 1766 | 19Dec2017 | Factory |  |  |
| Secondary Temp. | 4484 | 22Feb2017 | Factory |  |  |
| Secondary Cond. | 3531 | 30Mar2017 | Factory |  |  |
| Transmissometer | 1396DR | 27Apr2016 | Factory |  |  |
| SBE 43 DO sensor | 1119 | 28Mar2017 | Factory |  |  |
| PAR | 70613 | 21Mar2016 | Factory |  |  |
| SeaPoint Fluor. | 3685 | 13Mar2017 |  |  |  |
| Pressure Sensor | 0550 | 24Feb2017 | Factory |  |  |
| Altimeter | 62355 | n/a | Factory |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/2488       Cruise ID#:	2017-08

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 4Mar17 | Factory |  |  |
| Conductivity | 2488 | 4Mar17 | Factory |  |  |
| WetLabs Fluorometer | Ws3s-953p | May2017 |  |  |  |
| Temperature:Secondary | ? | 0842 |  |  |  |
| Flow meter | ? | n/a |  |  |  |