## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 20 March 2025 | Updated channel names & formats in TOB files. G.G. |
| ### 26 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| ### 1 Sept 2020 | ### Added HPLC data. S.H. |
| 14March2017 | Correction to CHL sample #288 |

## PROCESSING NOTES

Cruise: 2017-09

Agency: OSD

Location: WCVI

Project: La Perouse

Party Chief: Robert M.

Platform: John P. Tully

Date: 31 August 2017 – 12 September 2017

Processed by: Germaine Gatien

Date of Processing: 9 January December 2018 – 8 February 2018

Number of original HEX files:  115	Number of CTD files: 115

Number of bottle files: 	69		Number of bottle casts processed: 69

Number of original TSG files: 2		Number of processed TSG files:	 2

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#1119) on the primary pump, a SeaPoint Fluorometer (#3685) on the secondary pump, a Biospherical QSP-400 PAR sensor (#70613), and an altimeter (#62355).

A thermosalinograph (Seacat 21 S/N 2488) was mounted with a Wet Labs WETstar fluorometer (S/N ws-3s 953p), remote temperature sensor and a flow meter.

Seasave version 7.26.7.107 was used for acquisition.

The data logging computer was the Tully CTD Laptop.

The deck unit was a Seabird model 11+, serial number 0425.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial #58879.

The oxygen kit was SIO (nsB-6009) model 2.36a kit #2.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

Note that as of January 2017 a change has been made to the threshold levels for quality flags 3and 4 for Extracted Chlorophyll. Flag=3 will now be 10%CV (formerly 15%) and Flag=4 will now be 30% (formerly 50%).

The Daily Science Log, rosette log sheets and analysis logs were generally in good order with detailed comments on problems; however, most loop sampling was not noted in the log. Notes from the Chief scientist were very helpful.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples, especially when far from shore. Users are advised to use extracted chlorophyll values, where available.

As was noted in the previous cruise using this equipment, 2017-08, the fluorescence channel dark values were higher than usual. They started at about 0.2ug/L and gradually drifted to values of about 0.4ug/L Before cast #180 the tubing and pump were cleaned with a bleach solution and from that point on the dark values were in the normal range. For casts #1-177 a time-dependent lienar correction was applied based on comparisons of dark values at 2000m, with a small adjustment (+0.02ug/L) to ensure there are no negative fluorescence values.

The CTD dissolved oxygen correction based on comparison to bottles is much larger than found during cruise 2017-06 in June and a little larger than for 2017-08 in August. This may be due to calibration drift, but different conditions and sampling patterns may also be a factor. There were many reversals in dissolved oxygen gradients in the top 75m for many of the casts off the west coast of Vancouver Island, with some having complex profiles. For this reason the fit was based primarily on bottles from below 75m. Small sub-surface DO maxima are also seen in the 2017-06 data and larger ones in the 2017-08 data, though those profiles are generally simpler than for 2017-09.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5 mL/L from 0 to 125db

±0.1 mL/L from 125db to 500db

±0.05 mL/L from 500db to 1000db

±0.02 mL/L below 1000db

Overall the TSG performed well with intake temperature data available throughout the cruise, but the salinity data were plagued with spikes not seen in the temperature or fluorescence. The spikes are typically of size 0.5psu and often larger and the frequency varies greatly. Channel Salinity:T0:C0 is given with only 3 decimal places. While comparisons with CTD data and rosette samples indicate that TSG salinity values were low by between 0.3psu and 0.5psu, recalibration is not justified due to great variability in the comparisons and in the spikiness of the traces.

The TSG fluorescence data were fairly close to the CTD fluorescence early in the cruise but much higher by the end. Fluorescence data were removed from 14:00 on September 7th to the end, and a warning was put in the headers that after 21:30 on September 3rd TSG fluorescence values may be too high.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Sampling notes from the chief scientist noted a number of issues. Based on these notes a few corrections were made to file names and station names in the raw files.

Nutrients, extracted chlorophyll, dissolved oxygen, NH4 and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked. They had all been used on 8 previous cruises.

The configuration changed stayed the same throughout the cruise.

One change was made to the configuration file and the file was saved as 2017-09-ctd.xmlcon:

- Factor E in the dissolved oxygen sensor calibration was changed from 0.0375 to improve the hysteresis correction based on the tests done using data from cruises 2017-06 and 2017-08.

##### Conversion of Full Files from Raw Data

All files were converted using 2017-09-ctd.xmlcon.

The hysteresis and Tau functions were selected. Depth was included in the conversion.

A few casts were examined and all expected channels are present.

- The temperature and conductivity channels track well during downcasts but are a little noisier on upcasts. There are spikes in both primary and secondary channels.

- PAR, Dissolved Oxygen and Altimetry look normal.

- As noted during 2017-08 the fluorescence dark value is higher than expected and drifting higher with time. However, towards the end of the cruise cleaning was done and the dark value returned to normal values.

- The descent rate is generally noisy, but high. Some casts from well offshore have very noisy descent rates with some complete reversals of direction.

- As seen in other recent cruises using this transmissometer there are odd profiles that look as though values got stuck for a half-second or so, then registered a few new values with the usual sampling rate then got stuck again. The profiles look reasonable.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2017-09-ctd.xmlcon. Depth was included.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. Outliers were found in the primary salinity in 5 casts:

- Casts #55, 69, 90 and 121 - CTDEDIT was used to clean a few points in Salinity:T0:C0 around 10m.

- Cast #207 - CTDEDIT was used to clean a few points in Salinity:T0:C0 around 40m.

The output files from CTDEDIT were copied to the BOT files.

A preliminary header check and no problems were found.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs. A few bottles were removed from the file because there was no sample number assigned.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The following changes were made to SAMAVG files:

- The unwanted line for Niskin bottle 2 was dropped from file 2017-09-0078.SAMAVG.

- File 2017-09-0124.SAMAVG was removed as there was no sampling and no sample numbers were assigned.

- The unwanted line for Niskin bottle 1 was dropped from file 2017-09-0171.SAMAVG.

- The unwanted line for Niskin bottle 3 was dropped from file 2017-09-0226.SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2017-09-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2017-09chl*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2017-09chl.csv. Loop data were saved in a separate file. The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2017-09oxy*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and loop data were moved to a separate file. The rosette file was saved as 2017-09oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in file 2017-09SAL*.xlsx which included a precision study. The analyses were carried out in a temperature-controlled lab 7 to 19 days after collection.

The files were simplified and saved as 2017-09sal.csv. File 2017-09sal.csv was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet NUTS_QF2017-09*.xlsx. This includes a precision study. The file was simplified and the saved as 2017-09-nuts.csv. The file was converted to individual NUT files.

NH4

NH4 data were obtained in file 2017-09_NH4.xls which included a report on precision. The file was simplified and saved as 2017-09NH4.csv and converted to individual NH4 files.

The SAL, CHL, OXY, NUT and NH4 files were merged with CST files in 5 steps.

After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. This step is important to avoid missing some samples in the creation of CHE files.

There was one case of NH4 coming from a different bottle than planned, but it was from the same depth, so that is not a problem. There were also some nutrients taken but not marked on the sampling log. Salinity sample #278 was to be taken at 5m but the sample was marked as #276 which would be at 20m.

This will be checked when COMPARE is run. No further problems were noted.

There are loop data in the salinity, nutrient, chlorophyll and oxygen spreadsheets, so those data were moved to file 2017-09 Loop_Data.xlsx.

The chief scientist provided a log of loop sampling. All loops were taken at the end of CTD casts while the 5m rosette was being fired.  Times were added to the loop file based on the end time of CTD casts.

##### Compare

Salinity

Compare was run with pressure as reference channel. There were 3 very large outliers (differences >0.4pus):

- For sample #276 (event #88) at 20m the bottle value was much lower than that from the CTD. The analyst noted that while the rosette sheet indicated that the sample should have been taken at 5m rather than 20m and flagged it 2. But it looks as though the error was in the sample number and it should be #278 as shown on the rosette log. It would then be within 0.0025 of the CTD.

- The other 2 large outliers are from 5m for casts with large salinity gradients. The standard deviations in CTD salinity were high for those 2 firings. So the bottle contents might be expected to be significantly saltier than the CTD reading due to vertical separation and/or incomplete flushing.

Most other outliers are from surface samples and to be expected, but 3 samples from deeper water look like they may be significant outliers:

- Cast #2 at 200m – This was near the bottom of the cast and the standard deviation in the CTD data was high. The error is in the direction expected from poor flushing of bottles when it is a bottle taken at the bottom of a cast, though not what is expected a little above the bottom. However, in this case the CTD went above the target pressure and was then lowered, so it would behave like a bottom bottle. There is likely no problem with the sample analysis or collection.

- Cast #37 at 500m (sample 96) – The CTD data is a little noisy but the profile shows water as fresh as the bottle sample would be from around 475m. Both CTD channels agree. The DO sample is also off with 475m looking like a reasonable source for the sample. So this suggests a misfire or leaky bottle. A flag of 5 is suggested with pad values. The analysts agreed to pad the value and add a flag of 5 though it is noted that the nutrients are not obviously out of line.

- Cast #136 – The CTD profile near the 74m bottle is odd, with both temperature and conductivity having a reversal during the upcast that was not seen in the downcast. The standard deviation in the CTD data is fairly high. This is likely not a problem with analysis or sampling.

When points are excluded with salinity differences >0.01psu and/or pressure < 350m the primary salinity is found to be lower than bottles by an average of 0.0029psu (standard deviation 0.0010). Using the same points the secondary salinity is found to be low by an average of 0.0006psu (standard deviation 0.0011).

Looking at the last 4 cruises with good salinity sampling the following patterns emerge:

- The primary salinity started very close to bottles in May, was lower by 0.001 in June, very close again in August and low by 0.003 for this cruise. Given some differences in sampling levels and flushing efficiency and tightness of fits, there was little drift through the first 3 cruises, but the September difference may show drift is increasing.

- The secondary salinity has been falling relative to bottles with time. It was high by about 0.005 in May, by 0.003 in June, by 0.0004 in August and is low by 0.0006 for this cruise. So there is steady drift that may be slowing down recently.

There is very little pressure dependence though the scatter in the data makes such a judgment weak. The pressure dependence in the secondary channel has been consistent through the 4 cruises. The primary pressure dependence varies more but is lower. The time-dependence is of the same magnitude for both channels, though of opposite sign. Over the 4 cruises the primary temporal drift is low and consistent, while that off the secondary started high and is dropping

Excluding points with a standard deviation in the CTD salinity of >0.0008 removed no further points.

In summary:

|  | Average Difference | Standard Deviation | Fit against Pressure | Fit against Time |
| --- | --- | --- | --- | --- |
| Sal0 – Sal Bottle | -0.0029 | 0.0010 | Diff = +5E-07*Press - 0.0034 | Diff = +2E-05*File Pair # - 0.0031 |
| Sal1 – Sal Bottle | -0.0006 | 0.0011 | Diff = -5E-07*Press - 0.00006 | Diff = -2E-07*File Pair # - 0.0004 |

File Pair # is the consecutive # of rosette casts and ranged from 1 to 25 for this cruise.

For full details for the COMPARE run see file 2017-09-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

When outliers were removed based on residuals, the fit was:

CTD DO Corrected = CTD DO * 1.0694 + 0.0482 (R2 = 0.97)	(Fit 1)

This correction is larger than the previous cruise and there is a lot of scatter.

If only DO<5mL/L are included and outliers are removed based on residuals, the fit was:

CTD DO Corrected = CTD DO * 1.0620 + 0.0606 (R2 = 0.97)	(Fit 2)

It is surprising that dropping many of the near-surface values leads to a smaller correction. Incomplete flushing usually leads to a smaller correction because bottle values are lower than they should be, and this is most noticeable near the surface where vertical gradients are larger. Dropping such points should thus lead to a larger correction. However, an examination of individual casts shows that for many casts off the west coast of Vancouver Island there were large DO reversals in the top 50 to 75m. Reversals would lead to larger corrections since some bottle values might be too high rather than too low. This would also explain the large scatter in the fit. So, it is best to remove at least some of the shallow data. When records were excluded that appear to have been affected by reversals and then further data removed based on residuals the fit was:

CTD DO Corrected = CTD DO * 1.0621 + 0.0589 (R2 = 0.92)	(Fit 3)

Fit 3 removes some data not removed in Fit 2 as some reversals had DO<5mL/L; there is little difference in the fits and Fit 2 is a little tighter.

During 2 previous cruises adjustments were made to the hysteresis parameter E in the configuration files. A check of hysteresis for this cruise is limited because there are only 7 bottles from 1500m downwards and 4 of those were fired at the bottom of casts where results tend to be anomalous. The deep bottles do not look out of line in a plot of differences versus CTD DO. When a fit was made of those 7 bottles the slope was lower at 1.0416 and offset much higher at 0.0747. When the offset was forced to match the full fit, the slope was similar at 1.069. This is weak evidence, but nothing indicates that there is significant hysteresis.

Significant outliers were investigated:

- Cast #31 at 100m – sample #58 – Out of line in comparison and in profile. Value looks too high. But this is a tricky one because nutrients and CTD profile show a reversal in at 75m. Nutrients and CTD DO do not show a reversal at 100m but the DO sample does and the draw temperature does. The T and S profiles are complex with average gradients quite low and some small reversals. There was likely active mixing in this region, so even a slight mismatch between bottle and CTD depths could explain some large differences. The nearest cast to the offshore side shows T and S reversals near 100m. No flag needed.

- Cast #37 at 500m – Sample #96 – As noted for salinity it looks like the bottle closed late or leaked. Should be flagged and padded.

- Cast #59 at 5m - Sample #181 – Complex gradient at bottle level. The standard deviation in the CTD DO data is high. Bottle likely ok. No flag needed.

- Cast #63 at 10m - Sample #187 – Very high DO gradient at this depth. The standard deviation in the CTD DO data is high. Bottle likely ok. No flag needed.

- Cast #66 at 10m - Sample #204 - High DO gradient near surface, but no sub-surface reversal. Poor flushing would explain bottle value higher than expected. No flag needed.

- Cast #69 at 10m - Sample #209 - High and complex DO gradient at bottle level. High standard deviation in CTD DO data. Bottle likely ok. No flag needed.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable.

The usual pattern of the fluorometer reading higher than the CHL samples when the latter have low values is seen in these data. At higher CHL the fluorescence seems a little closer to CHL than usual. But there are few CHL values >10ug/L

During 2017-08 this sensor was found to have high dark values and those values increased through the cruise reaching ~0.2ug/L by the end. During this cruise that drift continued until by cast #147 it had reached 0.39ug/L at 2000m. Before event #168 some tubing was cleaned and before event #180 a more thorough cleaning was done to the pump and tubing. Thereafter the dark value was noted to be much better. This method assumes that a simple offset is appropriate to the whole profile; this is impossible to prove especially because of the geographic variability. There were some very high values after the cleaning. Using the pre-cleaning data a fit was made that excluded CHL>4ug/L to judge the difference.

The most notable feature is the large offset in the fit before cleaning. Since the dark value drifted through the cruise, the difference between the two offsets suggests an average difference on the order of 0.4ug/L which is about the dark value noted late in the cruise. Being as close as that is likely accidental given that the offset grew through the cruise and the fluorometer behaves differently as CHL values vary. But it does suggest that the drift in fluorescence is not just a deep-water phenomenon but affects the whole water column.

For full details of the comparison see file 2017-09-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run during 2017-05 when the same equipment was in use and the best results were found using a setting of +2.5s. For the next 4 uses of this sensor the same setting looked appropriate. A few casts were examined from 2017-08; there were many stops for bottles and both temperature channels were noisy so the tests were not easy to interpret, but overall +2.5s looked like a good choice. For this cruise it looks like a higher setting is better.

ALIGNCTD was run on all casts using +3.5s.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice. One cast was checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. Some results from other cruises that used these sensors are included for comparison; those entries are shaded.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2017-05-0101 | 1000 | +0.0006 | +0.0005 | +0.0055 | High, Moderate |
|  | 1900 | +0.0004 | +0.0005 | +0.0052 | “ |
| 2017-05-0139 | 1000 | +0.0002 | +0.0005 | +0.0051 | High, X Noisy |
|  | 1900 | +0.0004 | +0.0005 | +0.0055 | “ |
| 2017-06-0036 | 1000 | +0.0002 | +0.0004 | +0.0047 | High, X.Noisy |
|  | 1900 | +0.0004 | +0.0004 | +0.0045 | “ |
|  | 2900 | +0.0003 | +0.0004 | +0.0042 | “ |
| 2017-06-0064 | 1000 | ~0 | +0.0004 | +0.0047 | High, Noisy |
|  | 1900 | +0.0003 | +0.0004 | +0.0042 | “ |
|  | 2900 | +0.0004 | +0.0003 | +0.0038 | “ |
| 2017-06-0085 | 1000 | +0.0002 | +0.0004 | +0.0043 | High, V Noisy |
|  | 1900 | +0.0002 | +0.0003 | +0.0040 | “ |
|  | 2900 | +0.0002 | +0.0003 | +0.0036 | “ |
| 2017-08-0047 | 1000 | +0.0002 | +0.00007 | +0.0009 | High, V Noisy |
|  | 1900 | +0.0002 | +0.00005 | +0.0005 | “ |
|  | 2900 | +0.0004 | +0.00004 | -0.0001 | “ |
| 2017-08-0074 | 1000 | -0.0002 | +0.00008 | +0.0011 | High, V Noisy |
|  | 1900 | +0.0000 | +0.00005 | +0.0005 | “ |
|  | 2900 | +0.0003 | +0.00002 | +0.0001 | “ |
| 2017-08-0082 | 1000 | -0.0001 | +0.00009 | +0.0011 | High, V Noisy |
|  | 1900 | +0.0002 | +0.00005 | +0.0006 | “ |
|  | 2900 | +0.0002 | +0.00001 | +0.0000 | “ |
| 2017-09-0041 | 1000 | -0.0005 | +0.00012 | +0.0020 | High, Noisy |
|  | 1900 | 0.0000 | +0.00011 | +0.0014 | “ |
| 2017-09-0085 | 1000 | -0.0005 | +0.00011 | +0.0017 | High, Noisy |
|  | 1900 | 0.0000 | +0.00010 | +0.0013 | “ |
| 2017-09-0114 | 1000 | -0.0006 | +0.00013 | +0.0019 | High, Noisy |
|  | 1900 | 0.0000 | +0.00011 | +0.0014 | “ |
| 2017-09-0177 | 1000 | -0.0004 | +0.00014 | +0.0022 | High, Noisy |
|  | 1900 | 0.0000 | +0.00011 | +0.0014 | “ |

The temperature differences are slightly higher than during the previous cruise, but are small. There is some depth dependence. Conductivity and salinity differences are slightly larger than during 2017-08 though smaller than 2017-05 and 2017-06. It is likely that drift in different directions for the two pairs leads to that odd pattern. The salinity differences are in agreement with the results of the bottle comparison.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers and Pressure

A cross-reference list was checked against the log book. No errors were found.

The cruise track was plotted and added to the end of this report.

A surface check was run and shows an average surface pressure for the cruise was 2.0db. This is a little low for the Tully when offshore, but there is a wide range of values.

During a few casts the CTD appears to have moved through the surface while data was being acquired: after the soak period for casts #36, and #110 and at the end of casts #104 and #119.

- During event #36 the CTD was seen to leave the water briefly as it came up from the 10m soak. The pressure varied from -0.01db to +0.09db over the 1s period when it appears to have been in air. The pumps were on and salinity was low ~2 to 5psu. At pressure 0.04 the CTD data looks like it might be right at the surface with 0 transmissivity and very low salinity. This suggests the pressure is good to ±0.1db.

- For event #110 the CTD was seen to leave the water briefly as it came up from the 10m soak. Acquisition had started and the pumps were on. The CTD was at -0.04db for about 10s during which the salinity was extremely low (~2psu) and falling during that period except for one spike to ~8psu. The pumps then turned off and the CTD returned to the water. This suggests the pressure is good to ±0.1db.

- During event 104 the CTD came out at the end of the cast. The pump turned off when the pressure was at ~1db. Transmissivity spikes to low values about 1s after the pressure became negative; this might be due to the CTD was going through a layer of debris at the surface, or it could just be a random spike. The pressure reads -1db which might suggest that the pressure is reading low by 1db. That is inconsistent with the observations during events #36 and #110.  PAR spikes to higher values (thought not very high since it was nighttime) at 0db.  Other variables are on pumped channels so are not helpful in assessing the pressure accuracy.

- During event #119 the PAR signal increases rapidly between 0.2 and 0.1db and then stays fairly steady when pressure moves from +0.1db to -0.2db. This suggests that the surface is at about 0.1db but this is weak evidence given PAR variability.

There is no consistent evidence to justify recalibration of pressure.

The header check showed that there were spikes in pressure for at least one cast but that proved to be during the upcast in a cast with no bottle sampling. No other problems were noted.

The altimetry checks were done later in processing than usual. The altimeter and water depth readings from the headers of the DEL and SAM files were exported to a spreadsheet. Water depths were compared with the log book entries. A few changes were made to the headers:

- The depth entries were adjusted for casts #14, 40, 44, 45, 73, 88, 103, 116, 131, 180 and 203.

- For the bottle files, the depth readings were made as above for casts #44 and 88. For casts #20 and 32 the altimetry header was removed because the only sampling was at the surface.

After correcting the SAM files they were bin averaged again and the final merge step was run for the bottle files.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 2 casts to see if the shifts applied to conductivity during the previous cruises when the same equipment was in use were appropriate and they were. As found during 2017-05 there is a lot of noise in the data and no alignment will remove it all.

SHIFT was run twice on all casts using -0.6 records for the primary conductivity and -0.6 records the secondary. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were many warnings from casts 110, 138, 145 and 153.

- For cast #110 these are from the time when the CTD left the water and a few spikes at the end of the cast. To ensure that the best data are selected by DELETE, some initial records were removed from the SHFC1 file and DELETE was rerun. Only the warnings from the end of the cast remained and those are of no concern for this file.

- For cast #138 there were a few bad pressure values but DELETE chose data appropriately.

- For casts #145 and 153 the warnings only pertain to upcast data, so are of no concern.

After this step ADD TIME CHANNEL was run to add Julian time channel to enable a time-dependent recalibration of fluorescence.

##### Other Comparisons

Previous experience with these sensors –

All sensors were used during 8 earlier cruises in 2017.

Cruise 2017-63 in early April was in the Strait of Georgia and Juan de Fuca Strait. The bottles from Juan de Fuca were trusted more than those from the Strait of Georgia because conditions were better for flushing Niskin bottles; they indicated that the primary salinity was close to bottles while the secondary was high by about 0.0045. The dissolved oxygen comparison from 2017-63 was also based on Juan de Fuca bottles.

Cruise 2017-03 was in Baynes Sound and the Strait of Georgia, with shallow sampling and poor flushing of bottles, so the results were not considered reliable enough to use for recalibration.

Cruise 2017-05 was mostly off the west coast of Vancouver Island where flushing was good at depth. The comparison between the bottles and CTD salinity was quite tight; the primary salinity was found to be very close to bottles and the secondary was high by 0.0054psu. The dissolved oxygen comparison was also quite tight and led to a calibration of slope 1.0448 and offset +0.0527.
Cruise 2017-06 was mostly well offshore and showed the primary salinity to be low by ~0.0012 and the secondary high by ~0.0032. The dissolved oxygen was recalibrated using slope 1.0399 and offset 0.0645.

Cruise 2017-64 showed primary and secondary salinity to be low by 0.0012psu and high by 0.0034psu in Juan de Fuca Strait. The dissolved oxygen comparison in Juan de Fuca was very close to that of 2017-06, so those results were used for 2017-64.

Cruise 2017-56 included only near-shore shallow casts so 2017-06 results were used for recalibration.

Cruise 2017-23 comparisons were not trusted, so the 2017-06 results were used for recalibration.

Cruise 2017-08 was mostly well offshore and showed the primary salinity to be low by 0.0002 and the secondary high by 0.0004. The dissolved oxygen was recalibrated using slope 1.0514 and offset 0.0533.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Salinity was low very close to the surface at a few casts in Queen Charlotte Sound and at a few deeper places in near-shore areas of the northern Strait of Georgia. Temperature data fell within the climatology except for slightly high values in the offshore area of the south-west section of the cruise and at one cast in the Gulf Islands. None of these excursions looks systematic or significant. All other salinity values were well within the historic range.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The primary T-S plots look slightly smoother than the secondary, but both have many unstable features. The comparison with bottles shows the secondary closest to bottles. The temporal and pressure dependence is very similar.

The ATC files were opened in CTDEDIT.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. Some bad salinity points were removed from a few files. All files required some editing.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts; no further editing was applied.

##### Recalibration

There is no evidence that the pressure requires recalibration.

The primary salinity was found to be lower than the bottle salinity by an average of 0.0029psu. Since slight inefficiency in Niskin flushing could account for bottle values being slightly high, recalibration by adding 0.002 looks reasonable and brings it into line with the secondary salinity.

During 2017-08 the fluorescence dark values were larger than expected and increased through most of the cruise. A time-dependent correction was applied that required addition of a Julian time channel. The same problem occurred through most of this cruise.

For 2017-08 there were data from 4000m, but for this cruise in shallower waters, the deepest sampling available is 2100m. Dark values were found for 10 casts with maximum pressures of at least 1900db.

Plots were made of dark values versus cast start times to see if the drift was approximately linear. As found for 2017-09 a linear fit was reasonable. We don’t expect a perfect fit because usage is likely more important than time. The fit based on 2000m data was:

Fluorescence Corrected = Fluorescence + 8.2244 - 0.0343* Cast Time (Julian Days)

An initial recalibration produced some small negative fluorescence values, so 0.02ug/L was added to the offset so the correction applied was:

Fluorescence Corrected = Fluorescence + 8.2444 - 0.0343* Cast Time (Julian Days)

Dissolved oxygen recalibration was discussed in section 5. There are 2 fits that might be appropriate. A first attempt used the larger correction (Fit 1) and that produced results that looked poor when the downcast CTD data were compared with the upcast bottles (§19). This may be because a large number of near-surface bottles in large gradients skewed the result so the fit did not suit most of the profile. So Fit 2 was used and that proved more successful.

File 2017-09-recal1.ccf was prepared to add 0.002 to the primary salinity and, to apply the following corrections to channels Oxygen:Dissolved:SBE and Fluorescence:URU:Seapoint:

CTD DO Corrected = CTD DO * 1.0620 -0.0606

For casts #1 to #177 only fluorescence was corrected by applying:

Fluorescence Corrected = Fluorescence + 8.2444 - 0.0343* Cast Time (Julian Days)

This correction was first applied to the SAMATC and MRGCLN2 files.

COMPARE was rerun for salinity and dissolved oxygen using roughly the same points as in the fit used for recalibration to ensure that corrections were appropriate and they were. The primary salinity was found to be low by an average of 0.001psu. This is as expected since incomplete flushing is likely responsible for part of the difference between bottle and CTD salinity.

See file 2017-09-sal-comp2.xlsx for details.

The average of differences in the DO fit was +0.0001mL/L and the standard deviation was 0.0221mL/L. See file 2017-09-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files using the same recalibration file.

The fluorescence profiles confirm that the dark values are more appropriate after this step and there are no negative values.

##### Fluorescence Processing and special files for Dr. Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases. The SAM files were put through REMOVE and HEADEDIT. Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems.

The fluorescence looks bad between 250 and 367m during cast #103. A text editor was used to replace the bad values with pad values. A note was added to the header of that file.

T-S plots look ok.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

After the first attempt at recalibration (using Fit 1) the differences were much higher than we usually see, so a different recalibration of dissolved oxygen was applied (Fit 2) and the results were better.

When outliers were removed based on residuals the CTD DO was higher than the titrated samples by an average of ~0.015mL/L (standard deviation of 0.028mL/L). This difference is still a little larger than usual. Between 100m and 400m this is likely due to incomplete flushing of Niskin bottles in the presence of fairly high gradients. Below 400m the differences are small. Above 100m the flushing effect will have variable signs due to many reversals. Since this is a rough comparison and conditions are so variable, no further fine-tuning of calibration is justified.

See 2017-09-dox-comp3.xlsx for details.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels:

Time:Day_Of_Year, Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to remove the header entry TIME UNITS, fix the vessel name and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence and PAR data are nominal and unedited except that some

records were removed in editing temperature and salinity.

The fluorometer values drifted during the previous cruise that used this equipment.

That drift continued starting with dark values early in the cruise of about

0.2ug/L and rising to about 0.4 ug/L by cast #177. At that point the tubing

and pump were cleaned and dark values then returned to the normal range.

A time-dependent correction was applied to casts #2 to #177 based on values at

2000m with a small adjustment to ensure there would be no negative values.

The correction was an offset that was applied to the whole profile.

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.5 mL/L from 0 to 125db

±0.1 mL/L from 125db to 500db

±0.05 mL/L from 500db to 1000db

±0.02 mL/L below 1000db

For details on the processing see document: 2017-09_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found.

A cross-reference list was produced.

The sensor history was updated.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The range of values was extremely large, from ~55% to 170%, with the majority of values between 100 and 110%. There is too much variability to conclude anything about the calibration of the dissolved oxygen sensor.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For all casts REMOVE was run to remove the following channels:

Time:Day_Of_Year, Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, correct the vessel name, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Data were exported from the CHE files to file 2017-09-bottles-final.xlsx. The entries were compared with the rosette log sheets to ensure no samples had been missed. A problem was found in cast #171 – it had out-of-order firing and an error was made in the building of the file. This was corrected.

Standards check and a header check were run on all files. Two casts (114 and 177) had some slightly negative fluorescence values; they were replaced with 0 values.

The track plot looks ok.

Plots of each file were examined. A few cases with reversals in nutrients were investigated but the CTD dissolved oxygen showed similar features so the bottle data are likely fine.

A cross-reference listing and header check were produced for the CHE files.

##### Thermosalinograph Data

There were 2 hex files.

a.) Checking calibrations

The configuration file did not change through the cruise, but contained an error in the fluorometer parameters. One file was renamed as 2017-09-tsg.xmlcon and the fluorometer calibration parameters were fixed.

b.) Conversion of Files

2 files were converted to CNV files using configuration file 2017-09-tsg.con.

Those CNV files were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots showed that the salinity signal had many one-sided spikes. These are likely to be due to bubbles and unlikely to be amenable to editing. The fluorescence was climbing quite steadily towards the end of the cruise, though not as smoothly as would be expected if the flow had stopped. The ship had moved into the northern part of the Strait of Georgia so a change in fluorescence is reasonable, but the extracted chlorophyll from rosette samples does not support a significant rise. The flow rate looks reasonably steady throughout the cruise.

The temperature differences were very noisy.

The TSG files were opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, lab temperature, salinity and fluorescence and the files were reduced to the times of CTD files. There were 111 cases of overlap.

Similarly, TSG data were found to match the loop samples.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude and longitude were all ≤0.0008° and the median differences were 0.0001° for the latitude and 0.0000° for the longitude. This shows both the times and positions are reliable for both systems.

d.) Comparison of T, S and Fuor from Loop & Rosette Samples and TSG and CTD data

- T1 vs T2 The intake temperature sensor worked throughout the cruise. The differences decrease slightly as temperatures increase as is expected as the intake temperature gets closer to the temperature of the ship.

- Flow Rate The flow rate ranged from 0.9 to 1.2 while the ship was stopped for CTD casts; the median value was 1.04. There were a few abrupt changes though not terribly large ones.

- TSG vs CTD The spreadsheet comparing CTD and TSG files was then examined to find the differences between the salinity, fluorescence and temperature from the CTD and the TSG.

| Using all data and median values over 2 minutes for TSG: |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | (TSG intake -CTD) Temp | (TSG lab -CTD) Temp | (TSG - CTD) Sal | FLtsg/FLctd |
| median | 0.0439 | 0.2762 | -0.0644 | 1.331 |
| Std Deviation | 0.3619 | 0.3351 | 0.4977 | 2.911 |
| avg | 0.1647 | 0.3663 | -0.1553 | 2.400 |
| max | 2.0397 | 2.0675 | 0.0704 | 16.982 |
| min | -1.0946 | -0.1530 | -4.5563 | 0.197 |

In an effort to simplify the very noisy comparison the 15 cases with the lowest standard deviation was selected and then the 5 of those with the largest differences between TSG and CTD values were excluded. The results below show that this produces a reasonable difference between the TSG temperature and the CTD and especially makes the heating in the loop look similar to what would be estimated based on the history of this system:

|  | (TSG intake –CTD) Temp | (TSG lab –CTD) Temp | (TSG – CTD) Sal |
| --- | --- | --- | --- |
| Median | 0.0024 | 0.1826 | -0.0245 |
| Std Deviation | 0.0086 | 0.0380 | 0.0241 |

The fluorescence data are generally very noisy where values are highest, so that approach is not reasonable for studying the ratio of the two.

The flow rate had no noticeable effect on either the differences between lab and intake temperatures or between TSG temperature and salinity and CTD temperature and salinity.

The TSG salinity is lower than the CTD salinity by a median value of 0.064psu using all data and by 0.024psu using the reduced set. There were sections of the record that had few problems with severe spiking, but it is possible that even where spikes are not large, there are small ones that are not noticeable on the time-series plots. The differences found here are unlikely to be primarily due to calibration drift.

The ratio of the TSG fluorescence to CTD fluorescence has a median value of 1.33 and standard deviation of 2.91, but there is an enormous variability with very high values in the Strait of Georgia that are not seen in the CTD fluorescence. The ratio varied from 0.2 to 17. In contrast during 2017-05 in the same area and with a wide range in extracted CHL values, there was a much tighter fit with TSG fluorescence varying from 0.2 to 1.9 times the CTD fluorescence.

This will be explored further in the next part of this section.

For more details see 2017-09-ctd4-tsg-comp.xls.

- TSG Fluorescence versus Extracted CHL and CTD Fluorescence

For casts #2 to #78 the two fluorometers have similar values with the TSG instrument a little lower, and no obvious dependence on fluorescence levels which vary greatly. From #82 to 150 the TSG fluorometer is reading a little higher than the CTD. After that the TSG rises relative to the CTD until it is much higher in the Strait of Georgia. There are no extracted CHL samples from the Strait of Georgia. There is a local peak in CHL during cast 156 (~4.6ug/L) after which values are mostly <2ug/L. And CTD fluorescence is highly variable late in the cruise but never reaches values as high as those measured on the west coast of Vancouver Island. It is quite possible that the TSG sampled waters with much higher chlorophyll waters in Queen Charlotte and/or Johnstone Straits. That might explain the rise in values through those areas, but not in the Strait of Georgia where we have CTD values to compare with; those don’t support the high TSG fluorescence. However, if high values were encountered earlier, there could be a problem with either the loop or the instrument being affected.

There is no problem with the flow and there is detail in the trace, but there could be a build-up of biological material that is gradually raising values.

(See 2017-09-loops-tsg-rosette-comp.xls.)

Loop vs Rosette

Differences were calculated between samples from the loop and from the rosette samples closest to 5m where available. There were only 3 points of comparison for nutrients, and the variability within those 3 was large especially for nitrates which were much higher in the loop for 2 of the 3. Silicate and Phosphate were closer with 1 of the 3 having loop values higher by about 8% for the 1 sample each. The other two silicate samples differed by <1% and the phosphates by <3%. All Nitrate and Phosphate samples were higher from the loop, but one Silicate sample was lower and 1 was very close.

There are only 9 points of comparison for salinity with the loop reading lower than the rosette by a median value of 0.008psu, but the loop readings ranged from being lower by 0.07 to higher by 0.02. Most cases had the loop value lower but there were 2 cases of loop values being higher and the standard deviation is 0.139psu. The rosette samples are likely from slightly deeper water.

Dissolved oxygen values were consistently higher in the loop than from the rosette with a median difference of 0.15mL/L. This could be at least partly due to rosette water coming from lower in the water column.

(See 2017-09-loops-tsg-rosette-comp.xls.)

- Loop Bottle - TSG Comparisons

There were 10 salinity loop samples and 7 for extracted chlorophyll. All were taken while the ship was stopped for a CTD cast. Most were not recorded in the log book.

The TSG salinity was lower than loop samples by a median of 0.055psu but the standard deviation of ~0.14psu.

Extracted chlorophyll values ranged from 0.4 to 8.8ug/L. The median ratio of TSG fluorescence to CHL varies from 1.0 to 15.5 with a median of 2.0. As usual the ratio is highest for low CHL It is ~1 for the 3 samples with CHL>1.5ug/L.

(See 2017-09-loops-tsg-rosette-comp.xls.)

- Rosette Bottle – TSG data

A final check was to see how the TSG values compared with the rosette values. The median value shows the TSG salinity to be reading lower by ~0.067psu (standard deviation of 0.07) and the fluorometer to be reading high by a median factor of 1.82, but the comparison was very noisy with a standard deviation of 5.1 in the ratio.

(See 2017-09-loops-tsg-rosette-comp.xls.)

- Calibration History

The temperature and conductivity sensors were recalibrated in March 2017 and this was the fourth use since then.

2017-05: The temperature in the lab was higher than CTD temperatures by about 0.18Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of ~0.005Cº. The TSG Salinity read lower than the CTD salinity and that appeared to be due to bubbles. The TSG fluorescence was higher than the CTD fluorescence when the latter is <1ug/L and about 77% of the CTD fluorescence for higher values >1ug/L. Comparisons with rosette chlorophyll samples were limited but suggest that the TSG fluorescence was about 70% of chlorophyll values. No recalibration was applied to the TSG data.

2017-06: The temperature in the lab was higher than CTD temperatures by from 0.15Cº to 0.20Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of 0.0038Cº. The TSG Salinity read lower than the CTD salinity and that appeared to be due to bubbles. The TSG fluorescence was very close to the CTD fluorescence. Comparisons with rosette chlorophyll samples were limited but suggest that the TSG fluorescence was about 1.8 times the loop and rosette chlorophyll values. No recalibration was applied to the TSG data.

2017-23: The temperature in the lab was higher than CTD temperatures by from 0.16Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of 0.004Cº. The TSG Salinity read lower than the CTD salinity by ~0.29psu and that appeared to be due to bubbles. The TSG fluorescence was very close to the CTD fluorescence. There were no loop samples. No recalibration was applied to the TSG data.

2017-08: The temperature in the lab was higher than CTD temperatures by a median of about 0.018Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of 0.0038Cº. The TSG Salinity read lower than the CTD salinity and that appeared to be due to bubbles. The TSG fluorescence was very close to the CTD fluorescence. Comparisons with rosette chlorophyll samples were limited but suggest that the TSG fluorescence was about 1.9 times the loop and rosette chlorophyll values and that there was noticeable difference between results while underway or stopped. No recalibration was applied to the TSG data.

Conclusions

1. The TSG clock worked well.

2. The TSG flow rate was at the expected level and generally steady.

3. The temperature increases in the loop by a median value of 0.18Cº and increases as intake temperature decreases, as expected.

4. The TSG intake temperature was higher than the CTD temperature by a median of about 0.002Cº.  This is as close as we can achieve given the variability at 4m depth and the fact that the TSG may draw temperature from a little higher in the water column and that there could possibly be some heating right at the intake. No correction will be applied.

5. The TSG salinity is lower than the CTD salinity by a median of 0.64psu or 0.024psu if only 10 cases with low standard deviation in the TSG data are used. The TSG salinity is lower than loop samples by a median of 0.055psu, but the very large standard deviation makes that comparison untrustworthy. The TSG salinity is lower than rosette salinity by a median of 0.067psu, but again the standard deviation is large and part of that difference could be due to rosette water coming from lower in the water column.

6. There are many single-point one-sided spikes towards lower salinity values and these are likely due to bubbles. While the comparison between CTD and TSG is likely affected by these spikes, they would not affect all the data, yet in all but 1 of 111 cases the TSG salinity is lower than the CTD. A similar result was found during cruise 2017-08.  This could mean that there is calibration drift, but a large drift is unlikely given a fairly recent factory calibration and similar results during cruises immediately after that calibration. There are likely more bubbles than are obvious and/or there are slight mismatches between the depths from which the TSG draws water and that from which the CTD data were gathered.

7. No recalibration of salinity is justified since the low values are likely due to bubbles and the comparisons are not consistent enough to lead to a reliable adjustment to correct for that. A warning will be put in the header that values are generally low and only 3 decimal places will be used in the final files as a further warning.

9. The TSG fluorescence data were fairly close to the CTD fluorescence until 21:30 on September 3rd and are likely a little high until 14:00 on September 7th. Thereafter they look unreliable and should be removed.

f.) Editing

Each file was opened in CTDEDIT but it was clear that editing salinity was not practical with so many spikes and frequent gradual drops in salinity followed by sudden increases that make it unclear which values are poor. Fluorescence data were removed from 14:00 September 7th onwards.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number,

Temperature:Difference, Conductivity:Primary, Flag and Position:New channels.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to standardize channel names and formats. A note was added to warn that salinity values are believed to be mostly reading low. A note was added explaining why some fluorescence data were removed and warning that from about 21:30 on September 3rd values may be too high. The format for salinity was reduced to 3 decimal places to reflect the poor quality.

Those files were saved as TOB files.

The Standards Check and Header Check were run; no problems were found.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

Particulars – from log book and Chief Scientist’s sampling notes

PAR OFF: None.

Out of Order firing: 78, 171.

Loop Sampling: 5009, 5037, 5044, 5055, 5072, 5090, 5121 (some samples mistakenly called 5120), 5156, 5177, 5188 (all with CHL, OXY, SAL – last 6 have NUT)

Deployment Method – Rosette brought to surface, pumps turned on, rosette taken to 10m and back up. Start archiving and wait 30s, then cast run.

TSG: 2 files – Fluorescence kept rising at end of file 2

CTD:

36 – LB13: looks like CTD came out of the water after the “down 10 and back”.

103 – Bad fluorescence data between 250db and 367db.

104 – LJ06: still archiving at end when out of the water.

110 – LBP8: CTD came out of the water after the “down 10 and back”, pressure spikes.

116 – LBP6 – depth varying through cast

145 – CS04: pressure spikes on the way up during the 50 db 30-sec wait.

153 – CS1B: pressure spikes at the bottom and a few on the upcast.

168 – SS4 – Washed Tubing between conductivity and fluorescence instruments to remove growth, fluorometer seemed to have a positive offset relative to normal performance.

181 – Washed through 0.1% bleach solution on secondary pump, then flushed with warm water took off and cleaned tubing leading to and from the fluorometer; offset looks closer to expected norms during cast.

- Niskins & Samples at CTD/Rosette casts:

- 44 – LC11: Rosette sheet has note saying that Niskin 19 (s/n 132) ran out of water so no salinity sample, but a sample was found and the value looks reasonable.

- 51 – LC08: Niskin 2 closed at 175 instead of 200 because shallower than expected.

- 64 – LC01: Niskin 4 (s/n 193) closed at 40 instead of 30.

- 78 – LD11: Niskin 2 closed at wrong depth; use Niskin 21 for 1500 db sample (#237).  Niskin 2 not needed.

- 124 – LBP2: Niskin 13 closed just for test, not needed, no file needed.

- 171 – SS3: Niskin 1 closed at wrong depth (281), Niskin 4 used for 20 db sample instead.  Niskin 1 not needed in file.

- 226 – 41: Niskin 3 closed by mistake at 200 db.  All other depths are off by one Niskin (so 0 db is Niskin 12 instead of 11).  Niskin 3 not needed in the CHE file.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 16Dec2017 | Factory |  |  |
| Conductivity | 1766 | 19Dec2017 | Factory |  |  |
| Secondary Temp. | 4484 | 22Feb2017 | Factory |  |  |
| Secondary Cond. | 3531 | 30Mar2017 | Factory |  |  |
| Transmissometer | 1396DR | 27Apr2016 | Factory |  |  |
| SBE 43 DO sensor | 1119 | 28Mar2017 | Factory |  |  |
| PAR | 70613 | 21Mar2016 | Factory |  |  |
| SeaPoint Fluor. | 3685 | 13Mar2017 |  |  |  |
| Pressure Sensor | 0550 | 24Feb2017 | Factory |  |  |
| Altimeter | 62355 | n/a | Factory |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/2488       Cruise ID#:	2017-09

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 4Mar17 | Factory |  |  |
| Conductivity | 2488 | 4Mar17 | Factory |  |  |
| WetLabs Fluorometer | Ws3s-953p | May2017 |  |  |  |
| Temperature:Secondary | ? | 0842 |  |  |  |
| Flow meter | ? | n/a |  |  |  |