## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 20 March 2025 | Updated channel names & formats in TOB files. G.G. |

## PROCESSING NOTES

Cruise: 2017-23

Agency: OSD

Location: B.C. Coast

Project: OSD Mooring Cruise

Party Chief: Spear D.

Platform: John P. Tully

Date: 4 July 2017 –18 July 2017

Processed by: Germaine Gatien

Date of Processing: 25 October 2017 – 10 November 2017

Number of original HEX files:  	62	Number of CTD files: 62

Number of bottle files: 	28		Number of bottle casts processed: 24

Number of original TSG files: 	2	Number of processed TSG files:	 2

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1396DR), a SBE 43 DO sensor (#1119) on the primary pump, a SeaPoint Fluorometer (#3685) on the secondary pump, a Biospherical QSP-400 PAR sensor (#4565), and an altimeter (#62355).

A thermosalinograph (Seacat 21 S/N 2488) was mounted with a Wet Labs WETstar fluorometer (S/N ws-3s 953p), remote temperature sensor and a flow meter.

Seasave version 7.26.1.8 was used for acquisition for casts #1 to 68 and V 7.26.2.13 from cast #75 to 106.

All casts were run with the LARS mid-ship station.

The salinometer used at IOS was a Guildline model 8400B, serial #68572.

The oxygen kit was SIO (nsB-6009) model 2.36a kit #2.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

Note that as of January 2017 a change has been made to the threshold levels for quality flags 3and 4 for Extracted Chlorophyll. Flag=3 will now be 10%CV (formerly 15%) and Flag=4 will now be 30% (formerly 50%).

Serious network problems occurred while processing these data, some of which resulted in truncated files. Checks were made to ensure no data were lost, but users are urged to report any odd features as they might be due to interruptions in routines.

The Daily Science Log contained no list of equipment or science crew. There were useful comments about problems encountered during the cruise and the deployment method used. There were some problems with sample numbers and event number entries in the rosette log sheets.

There was corruption of data saved for 3 casts. Cast #61 has readable data for only part of the downcast; this was not a rosette cast. Casts #34 and #67 have full downcasts but corruption begins part-way though the upcast. Acquisition appeared to be working normally at sea, and bottles were fired successfully, but the corrupted records could not be fixed. Because pressure is not available for the majority of bottles from those two casts, Depth:Nominal was used with values taken from the rosette sheets instead of the usual channel Depth which is computed from pressure. For the bottles with no CTD data available the sample values were compared with nearby casts and they look appropriate. The profiles also look ok.

The usual IOS CTD deployment method is to soak at 10m with pumps on before returning to the surface to run the full caste. This method was only used for the first 2 casts. The researchers wanted to observe undisturbed surface waters, so for all other casts the CTD was lowered to about 2m with pumps off. The pumps were then turned on and the CTD was soaked at 2m for at least 40s. In most cases the 2 conductivity channels were in reasonable agreement after the soak, but occasionally it looks like a longer wait was needed and in general the correspondence does look better below 10m.

The general advice to take salinity samples from Niskin #2 is to ensure there are samples that are deep but at least 10m off the bottom. During this cruise Niskin #2 was often fired near the bottom, so this is not the best choice for CTD calibration purposes. Adding or substituting a bottle from higher in the water column is recommended to avoid the complications found in bottom samples.

The transmissivity data are considered of lower quality than usual due to what appear to be “stuck values” that last for about half a second. The profiles and the correspondence between downcasts and upcasts look reasonable.

A PAR sensor was used on cast #1 only.

The primary temperature/salinity pair were selected for processing as the data were a little less noisy than the secondary pair and the primary salinity was found to be close to bottles during the last cruise with lots of deep calibration sampling.

While CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples, especially when far from shore. The data show the usual pattern for this type of fluorometer with the ratio FL/CHL being high at very low CHL values and dropping as CHL rises.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.2 mL/L from 0 to 100db

±0.1 mL/L from 100db to 250db

±0.04 mL/L from 250db to 400db

±0.02 mL/L below 400db

The TSG performed well overall, with intake temperature available throughout the cruise. While there are many large 1-sided spikes in salinity, they are not as frequent as have often been seen over the past few years. Many of the larger single-point spikes were removed with a graphical editor. The TSG fluorometer readings were very close to those from the CTD sensor. No loop samples were taken.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The history of the pressure sensor, conductivity and DO sensors were checked. They had all been used on 6 previous cruises, the data from one of which had not been processed yet.

The configuration changed after the first cast and stayed the same through the rest of the cruise. There were no errors in the configuration files, but based on a bottle comparison during 2017-06 the DO configuration was updated. Factor E in the dissolved oxygen sensor calibration was changed from 0.036 to 0.038 to improve the hysteresis correction.

##### Conversion of Full Files from Raw Data

All files were converted using 2017-23-ctd1.xmlcon for cast #1 and 2017-23-ctd2.xmlcon for casts #2 – 106 The hysteresis and Tau functions were selected. Depth was included in the conversion.

A few casts were examined and all expected channels are present.

As usual the temperature and salinity channels track fairly well during downcasts but are noisier on upcasts, except during casts 45 and 46. For those casts the differences between the conductivity channels was larger. Just before cast #48 a jellyfish was flushed out of the plumbing and differences thereafter return to values similar to before cast #45. There are spikes in both primary and secondary channels.

Fluorescence and Dissolved Oxygen look normal.

Altimetry is sometimes very noisy right at the bottom though examination of plots would enable reasonable estimates.

As has been noted before for transmissometer #1396 has an odd profile with values seeming to get stuck for about a half-second. The up- and down-traces track reasonably well.

The PAR sensor was removed and the transmissometer was moved to where the PAR had been to see if that would improve the signal, but that did not help.

The PAR data looked good for the one cast for which it was mounted.

The descent rate was generally high and fairly steady except in the offshore areas where it was much noisier.

There was corruption of the data saved for 3 casts. Cast #61 has readable data for only part of the downcast. Casts #34 and #67 have full downcasts but corruption begins part-way though the upcasts. Acquisition appeared to be working normally at sea, and bottles were fired successfully, but the saved files were corrupted and could not be fixed.

Acquisition for the first two casts followed the usual pattern of soaking at 10m before running a full cast. For the rest of the cruise the 10m-soak was not used. The CTD was lowered to about 2m where it was soaked for at least 1 minute. Usually the wait looks long enough for the two conductivity channels to look similar, but occasionally it looks like a longer wait was needed.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2017-23-ctd1.xmlcon for cast #1 and 2017-23-ctd2.xmlcon for all other casts. Depth was included.

Since there was no sampling from cast #66 the ROS file was deleted.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. Cast #43 has significant differences between the salinity channels starting at 50m. This looks like the beginning of the plumbing problem noted in the log book for cast #45.

A preliminary header check was run and no problems were found. CTD fluorescence did not go off-scale.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs. A few problems were discovered:

- Event #34 – There are only data for bottles #1 to 5. After that there is a mess of symbols. This does not look like a case of corruption by carriage returns. Two routines programs that sometimes help fix hex files, fixwp.exe and fixdat.wp, failed to produce any further data. The latter routine produced the error message that the scan length could not be determined, and the lines are full. When the SAMAVG files are ready, extra lines can be added so that the sample data can be added. Nominal depths can be added but no CTD data will be available.

- Event 67 – the problem is the same as for cast #34 with CTD data from bottles 1 to 3 only. So this will be another case that needs extra lines added to the SAMAVG file.

- Events 48 and 51 – sample numbers 231 and 232 were used in both events, change to 9231 and 9232 for 48 and leave 51 alone. The first instance of #232 was used for SPC sampling only.

- Event 62 – there are 20 sample numbers but 22 bottles closed – 2 were not sampled, so -99 will be entered for sample numbers. The lines should be removed from the SAM and SAMAVG files when they are ready.

- Event 97 – the sample numbers were originally entered as 314-327 but those were repeats so they were changed on the log sheet to 914-927 but the salinity spreadsheet has 9314-9327. The latter is the usual way to handle this sort of problem. Since there was only salinity sampling for that cast the numbers 9314 to 9327 will be used. This was noted on the rosette sheet. There was also a sample 999 but with no sampling that will go into the bottle files. The line for sample 999 was removed as it is not needed.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The missing lines for casts #34 and #67 were added to the SAMAVG files.

The unwanted lines for Niskin bottles 4-24 were dropped from file 2017-23-0062.SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2017-23-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2017-23chl*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2017-23chl.csv. The csv file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2017-23oxy.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and loop data were moved to a separate file. The rosette file was saved as 2017-23oxy.csv. That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in file 2017-23SAL.xlsx which included a precision study. The analyses were done 5 to 18 days after collection. Precision is lower than expected given the timely analysis.  The nylon inserts used were those that were found to provide poor seals, not the inserts we normally use.

The files were simplified and saved as 2017-23sal.csv. File 2017-23sal.csv was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2017-23_nutrients*.xlsx. This includes a precision study. The file was simplified and the saved as 2017-23-nuts.csv. The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. This step is important to avoid missing some samples in the creation of CHE files.

Some problems were found:

- Event #37 –The salinity samples for this event were supposed to be 184, 187 & 194. Two samples are listed in the salinity spreadsheet and analysis log but are named 196 (station MS1, Niskin 2) and 199 (station MS1 Niskin 5). The analysis values make sense for those Niskin bottles, so it looks like 199 should be 187 and 196 should be 184. No entry was found for Niskin 13 at station MS1. The sample numbers were changed and flag 2 and comments added. Note that the same problem applies to DEL18 samples, but not to DO, CHL or Nutrients.

- Event #55 – The only sample #243 is for instrument calibration, so no bottle file is required.

The following lines were removed from the MRG files because bottles were fired but not needed and were not assigned sample numbers:

- Cast 62, Niskin #21 and 22

CLEAN was run to add quality flag 0 to empty entries and to reset the header limits.

##### Compare

Salinity

Incomplete flushing of Niskin bottles complicates interpretation of CTD salinity comparisons with bottles. During upcasts the water in the bottles likely comes from a little deeper in the water column than the level at which the bottle closed. This effect is most significant in quiet waters where flushing is less efficient and in high salinity gradients where incomplete flushing has a larger effect. Bottles fired at the bottom complicate interpretation further because poor flushing there will result in bottles containing water of lower salinity than found in ambient waters and there may be complex effects due to shed wakes bouncing off the bottom or the presence of bottom currents. So we usually recommend that salinity samples taken for calibration purposes be taken at least 10m above the bottom. Niskin #2 is usually recommended to get the deepest samples that are at least 10m above the bottom. For this project this is an inappropriate choice because there are often 2 bottles fired at the bottom. If bottom salinity samples are needed for other purposes, then it is recommended that there also be samples taken above that level.

Compare was run with pressure as reference channel. One extreme outlier was found in cast #67.

Sample #281 is lower than the CTD salinity by about 2.4psu. It looks like it is from about 50db, not the intended 2200db.  There is no sample missing, so it either came from the wrong Niskin bottle, or the bottle closed at the wrong level or there was a serious problem in analysis.

There were many cases of the bottle salinity looking high relative to CTD salinity in the top 200m which is to be expected because of incomplete flushing of Niskin bottles in the region of higher salinity gradient. There are also some cases of outliers with bottle values being low; most come from near the bottom of casts which is also likely to be due to poor flushing since the effect is of opposite sign for bottom samples. One exception was the 75m sample from cast #97; the standard deviation in the CTD salinity was very high, so flagging the sample is not appropriate.

Excluding the bottles above 200m, the primary salinity was low by an average of 0.0032 and the secondary salinity was high by 0.0026; the standard deviations were 0.0022 and 0.0026, respectively. The primary fit is flatter than the secondary.  The variations with time are similar (the primary is slightly flatter) but there is too much scatter to be significant. When bottom bottles are excluded as well, the primary salinity was low by 0.0045psu and the secondary high by 0.0014psu and the secondary looks slightly flatter. Unfortunately, that leaves only 10 bottles in the fit. Using only the 8 bottom bottles below 200db, the primary salinity is low by an average of 0.0012psu and the secondary is high by 0.0034psu.

Cast #97 was examined carefully to see if small flushing problems seemed likely to account for the differences in the comparisons. At 500m the salinity measured by the CTD at the beginning of the stop was higher by ~0.004psu at the beginning of the stop than when the CTD data had clearly equilibrated. The CTD showed little vertical motion during the stop so it is quite likely that the bottle would contain water with salinity higher than ambient values. The bottle comparison showed bottle values higher than CTD by 0.0027psu. At 200m the values were higher by 0.01 to 02 early in the stop, though there was more vertical motion which would improve the flushing somewhat. These results show that having the CTD look low by from 0.001psu to 0.005psu cannot be presumed to be evidence of calibration drift.

For full details for the COMPARE run see file 2017-23-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There were 2 significant outliers, both from cast #67. The bottle values were 9.7 and 11.4mL/L while the CTD values were 1.5 and 0.6mL/L. The CTD values look reasonable for the depths. Given both analysis concerns and completely unreasonable values for the depths, it is suggested that the flags be changed to 5 and the values be padded.

When those outliers were removed based on residuals, the fit was:

CTD DO Corrected = CTD DO * 1.054 + 0.048		 (R2 = 0.9668)

Most of the rejected points came from the top 200m. The correction is higher than during 2017-06 by about 1% and about 0.8% higher than that for 2017-64.

Because this cruise visited such a wide range of sites – narrow inlets, deep offshore sites and areas of hypoxic conditions, some casts with very low DO gradients, others having sections of extremely high gradients – it is difficult to say what points should be rejected as outliers. One fit that rejected some bottles from casts in hypoxic waters and active inlets looked similar to the fits found on earlier cruises, but there were many outliers. While some of the casts have conditions that may affect the DO sensor response, the fits for other casts that do not look unusual also lead to a higher correction.

A check was made for hysteresis by excluding all points from above 800m and there was no evidence of hysteresis.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further outliers were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor. The data show the usual pattern for this type of fluorometer with the ratio FL/CHL being high at very low CHL values and then falling as CHL increases.

The scatter in the fit of CHL versus CTD Fluorescence is very high for CHL>5ug/L.

For full details of the comparison see file 2017-23-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run during 2017-05 when the same equipment was in use and the best results were found using a setting of +2.5s. A few tests were run for this cruise with the same result although noisy upcasts made the judgment very subjective. ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice. One cast was checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. Some results from other cruises that used these sensors are included for comparison; those entries are shaded.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2017-05-0038 | 1000 | +0.0001 | +0.0004 | +0.0046 | High, X Noisy |
|  | 1900 | +0.0003 | +0.0004 | +0.0040 |  |
| 2017-05-0101 | 1000 | +0.0006 | +0.0005 | +0.0055 | High, Moderate |
|  | 1900 | +0.0004 | +0.0005 | +0.0052 | “ |
| 2017-05-0139 | 1000 | +0.0002 | +0.0005 | +0.0051 | High, X Noisy |
|  | 1900 | +0.0004 | +0.0005 | +0.0055 | “ |
| 2017-06-0036 | 1000 | +0.0002 | +0.0004 | +0.0047 | High, X.Noisy |
|  | 1900 | +0.0004 | +0.0004 | +0.0045 | “ |
|  | 2900 | +0.0003 | +0.0004 | +0.0042 | “ |
| 2017-06-0064 | 1000 | ~0 | +0.0004 | +0.0047 | High, Noisy |
|  | 1900 | +0.0003 | +0.0004 | +0.0042 | “ |
|  | 2900 | +0.0004 | +0.0003 | +0.0038 | “ |
| 2017-06-0085 | 1000 | +0.0002 | +0.0004 | +0.0043 | High, V Noisy |
|  | 1900 | +0.0002 | +0.0003 | +0.0040 | “ |
|  | 2900 | +0.0002 | +0.0003 | +0.0036 | “ |
| 2017-23-0026 | 1000 | +0.0003 | +0.0004 | +0.0050 | High, Mod |
|  | 1500 | +0.0003 | +0.0004 | +0.0044 | “ |
| 2017-23-0066 | 1000 | +0.0001 | +0.0004 | +0.0050 | High, V Noisy |
|  | 1500 | +0.0003 | +0.0004 | +0.0049 | “ |
|  | 1900 | +0.0004 | +0.0004 | +0.0046 | “ |
| 2017-23-0068 | 1000 | +0.0002 | +0.0004 | +0.0050 | High, Noisy |
|  | 1500 | +0.0004 | +0.0004 | +0.0048 | “ |

The temperature and conductivity differences are small and similar to other recent cruises. The salinity differences are higher than during 2017-06 but similar to 2017-05. The most recent cruise using these sensors had no sampling deep enough to produce a useful comparison.

The results of the bottle comparison discussed in section 5 are consistent with a difference of ~0.005psu.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. There were errors in the station names in 7 casts. Those were corrected in both the CLN and SAM files.

The cruise track was plotted and added to the end of this report.

A surface check was run and shows an average surface pressure for the cruise was 0.36db. This is much lower than usual for the Tully, but the casts were not run in the usual way; rather they started right at the surface. Plots indicate that oceanic values for conductivity start at pressures 0db to 0.5db, with most ~0.1db. Given there would be some delay in conductivity with the pumps off, this suggests that the pressure calibration is good.

The header check showed that there were spikes in many channels; these are likely all from near-surface data before the pumps were turned on.

The altimeter and water depth readings from the headers of the CLN and SAM files were exported to a spreadsheet.

A “check value” was calculated as Water Depth - Max Depth Sampled – bottom altimeter reading. When that value was >4m the water depth and altimetry entries were investigated.

In 16 cases the water depths were altered. For most the log entries for depth were used. In a few cases where the log entry also looks wrong, an estimate was made based on the maximum depth sampled plus the altimetry header value.

After correcting the SAM files they were bin averaged again and the final merge step was run for the bottle files; the CLEAN step was repeated as well.

A check was made of all CLN files to see if profiles are complete. Three are not:

Event 34 has a complete downcast and partial upcast (5 bottles up to 75m). There are samples for the full upcast.

Event 61 has a partial downcast. The log indicates it went to 710db but we have data only to 425db.

Event 67 has a complete downcast and partial upcast (3 bottles up to 1500m). There are samples for the full upcast but the samples from the first 3 bottles are all flagged with values that look like they are from the top 50m.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel,

Conductivity

Tests were run on 3 casts to see if the shifts applied to conductivity during the previous cruise when the same equipment was in use were appropriate and they were. As found during 2017-05 there is a lot of noise in the data and no alignment will remove it all.

SHIFT was run twice on all casts using -0.6 records for both the primary conductivity and the secondary. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were many warnings but most came from the beginning of files before the pumps came on or from the end of upcasts, so are of no concern. A few problems were found later:

- A problem occurred with cast #32. There were pressure spikes that led to DELETE removing all data above 40db. The SHFC1 file was put through CLIP to remove the first 3000scans and DELETE was rerun with good results. The data removed were collected with pumps off.

- There was also a problem with cast #84 which was repaired in the same way by removing 4000 records.

- There was also a problem with cast #91 which was repaired in the same way by removing 3500 records.

##### Other Comparisons

Previous experience with these sensors –

Pressure, temperature, conductivity and dissolved oxygen sensors were used during 6 earlier cruises in 2017 but one has not yet been processed.

- 2017-63 was a cruise in the Strait of Georgia and Juan de Fuca Strait. The bottles from Juan de Fuca were trusted more than those from the Strait of Georgia because conditions were better for flushing Niskin bottles; they indicated that the primary salinity was close to bottles while the secondary was high by about 0.0045. Since flushing may not have been complete even in Juan de Fuca, both salinity channels could have been reading a little higher than those results suggest. The dissolved oxygen comparison from 2017-63 was also based on Juan de Fuca bottles and had slope 1.0422 and offset 0.036mL/L.

- Cruise 2017-03 was in Baynes Sound and the Strait of Georgia, with shallow sampling and poor flushing of bottles, so the results were not considered reliable enough to use for recalibration.

- Cruise 2017-05 was mostly off the west coast of Vancouver Island where flushing was good at depth. The comparison between bottles & CTD salinity was quite tight; the primary salinity was found to be very close to bottles and the secondary was high by 0.0054psu. The dissolved oxygen comparison was also quite tight and led to a calibration of slope 1.0448 and offset +0.0527.

- Cruise 2017-06 was mostly well offshore; primary salinity was low by ~0.0012 and the secondary high by ~0.0032. The dissolved oxygen recalibration used slope 1.0399/offset 0.0645.

- Cruise 2017-64 was mostly in inland waters; the best-flushed bottles showed the primary salinity to be low by 0.0012psu and the secondary high by 0.0034psu. Dissolved oxygen was recalibrated using slope 1.0399 and offset 0.065mL/L.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Most temperature and salinity data fell within those limits. Temperatures were high at depth in some of the inlets, likely because those areas are not well represented in the climatology; salinity was low at depth in the same areas. There is no indication of calibration problems.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

Both CTD salinity channels have noisy sections with unstable features, but the primary seems less affected. The primary temperature/salinity pair showed less variation with pressure.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some records from near the top and bottom of the casts. Some bad salinity points were removed from one file. All files required some editing. There were network problems during the process.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts. Three casts had been truncated and had to be re-edited, presumably due to the network interruptions experienced during editing.

##### Recalibration

There is no evidence to suggest the pressure should be recalibrated.

The bottle comparison is too confusing to show how the salinity should be recalibrated. The primary salinity appears to be reading low, but the effects of incomplete flushing account for some of the difference. The best comparison available for these sensors comes from cruise 2017-06 when the primary salinity was found to be lower than the bottle salinity by an average of 0.0012psu. During 2017-64 it was found to be low by 0.0018 but the effects of incomplete flushing were likely more significant for that cruise. Salinity was not recalibrated for 2017-06 since the small difference was quite possibly an effect of incomplete flushing. During 2017-64 the primary salinity was not selected for archiving.  No recalibration will be applied to salinity.

Dissolved oxygen recalibration was discussed in section 5. The result for this cruise was:

CTD DO Corrected = CTD DO * 1.054 + 0.048

Since the salinity comparison showed evidence of incomplete flushing of Niskin bottles we must assume that DO is similarly affected. Calibration drift leads this type of sensor to read lower than it should. If the DO samples come from deeper waters, then their values would generally be lower than ambient waters, leading to smaller corrections except at the bottom of casts. However, the correction for this cruise is larger than for the 2 previous cruises using these sensors. This may be partly due to higher gradients and hypoxia may lead to some delay in sensor response, but even casts without such factors have fits that suggest a larger correction. However, it does appear that the casts that are expected to be most affected by poor flushing had a lot of bottles rejected from the general fit. The fit for those 5 casts after removing outliers was:

CTD DO Corrected = CTD DO * 1.0406 + 0.0783

So the larger correction in the general fit is likely due to many of the poorly flushed bottles being rejected as outliers. The general fit is the best measure we have of the CTD DO sensor calibration.

File 2017-23-recal1.ccf was prepared to apply the following correction to channel Oxygen:Dissolved:SBE:

CTD DO Corrected = CTD DO * 1.054 + 0.048

This correction was first applied to the SAM and MRGCLN2 files. COMPARE was rerun for dissolved oxygen; however, cast #67 was not included in the comparison because the DO values had all been padded based on the previous comparison. After outliers were removed based on residuals resulting in roughly the same number of records left in the comparison, the average of differences in the DO fit was +0.0008mL/L and the standard deviation was 0.0022mL/L. The fit against CTD DO was flatter than usual. (See file 2017-23-DO-comp2.xlsx for details.)

CALIBRATE was then run on the EDT files using the same recalibration file.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

When outliers were removed based on residuals the CTD DO was higher than the titrated samples by an average of ~0.033mL/L (standard deviation of 0.053mL/L). Looking at the same data, differences versus pressure shows that the sensor DO tends to be on the high side above 400m, which is likely due to incomplete flushing of Niskin bottles combined with higher gradients.

No further recalibration is justified. See 2017-23-dox-comp3.xlsx for details.

##### Fluorescence Processing and special files for Dr. Peña

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined and no problems were found.

Profile plots were examined. The transmissivity profiles for two casts had some 0 values.

- For cast #94 the low values are at 70db and the fluorescence trace shows a rise at that depth and the upcast also shows low values at that depth; so, the transmissivity is likely correct and it was left unchanged.

- For cast #46 the zero values are deeper at 170db, so fluorescence is not helpful, but the upcast transmissivity shows no low values at that depth, so the very low transmissivity values in the AVG file were padded.

No other problems were found in the profiles.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Fluorescence and PAR data are nominal and unedited except that some records were

removed in editing temperature and salinity.

The transmissometer data alternated between sampling at the normal rate and getting

stuck on a single value for about a half second. While these data are

considered of lower quality than usual, the general shape of profiles look

reasonable and downcasts and upcasts are in reasonable agreement.

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.2 mL/L from 0 to 100db

±0.1 mL/L from 100db to 250db

±0.04 mL/L from 250db to 400db

±0.02 mL/L below 400db

For details on the processing see document: 2017-23_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found.

A cross-reference list was produced.

The sensor history was updated.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The values ranged from 100% to 150% with most between 105% and 110%.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Casts #34 and 67 were processed differently to deal with the fact that we know nominal depths for many samples for which we do not have pressure.

- First, both files were sorted on bottles numbers in reverse order. This also orders them on depth and using bottle numbers makes it a little easier to compare with the rosette logs where more than one bottle is fired at a single depth. (Output MRGSORT2)

- Next, the channels were reordered with depth before pressure. The depth channel will be renamed as DEPTH:NOMINAL, so the depths for those cases where pressures are known were changed to match the rosette plan.

- For cast #67 the comments for samples 280 and 282 were simplified since all samples are out of line and the Niskin bottles must have closed at the wrong depths.

- Headedit was run with different settings to rename Depth as Depth:Nominal and to enter separate  header comments specific to each of these files.

The CHE file for cast #55 is not needed and will not be archived.

Data were exported from the CHE files to file 2017-23-bottles-final.xlsx. The entries were compared with the rosette log sheets to ensure no samples had been missed.

Extra comparisons were done between casts #34 and #67 with nearby bottle casts to ensure that the bottle analysis results are consistent with values expected at the nominal depths and they were. (See files “casts34-36-comp.xlsx” and “casts67-62-comp.xlsx”.)

Standards check and a header check were run on all files. No problems were found.

The track plot looks ok.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files.

##### Thermosalinograph Data

There were 2 hex files.

a.) Checking calibrations

The configuration file did not change through the cruise, but contained an error in the fluorometer parameters. One file was renamed as 2017-23-tsg.xmlcon and the fluorometer calibration parameters were fixed.

b.) Conversion of Files

The 2 files were converted to CNV files using configuration file 2017-23-tsg.con.

The CNV files were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

A time-series plot showed that the salinity signal had some one-sided spikes of size 0.5 to 1.0psu that are not associated with temperature spikes. These are likely due to bubbles, but the problem was less severe than during many other cruises over the past few years. The flow rate was quite steady, starting at about 1.2 for the first short file, and then ranging from ~1.05 to ~1.1 for the second file. Flow stopped at about 0810 on July 17th, so data from after that will be removed. The intake thermometer worked throughout the cruise. The temperature differences are extremely noisy.

There was no overlap between CTD files and the first TSG file. The second TSG file was opened in EXCEL, median and standard deviations (over 5 records) were calculated for intake temperature, lab temperature, salinity and fluorescence and the files were reduced to the times of CTD files. There were 61 casts that overlapped with TSG records. There were no loop samples.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The differences in latitude and longitude were all ≤0.0005° and the median differences were 0.0000° for both. This shows both the times and positions are reliable for both systems.

d.) Comparison of T, S and Fl from Rosette Samples and TSG and CTD data

- T1 vs T2 The intake temperature sensor worked throughout the cruise. The differences are highest in the inlets on the west coast of Vancouver Island. The differences decrease slightly as temperatures increase, as is expected as the intake temperature gets closer to the temperature of the ship. The median heating during file #2 was 0.17C°.

- Flow Rate The flow rate ranged from 1.05 to 1.13.

- TSG vs CTD The spreadsheet comparing CTD and TSG files was then examined to find the differences between the salinity, fluorescence and temperature from the CTD and the TSG.

|  | TSG Intake Temp – CTD Temp | TSG SAL-CTD SAL | TSG Fluor/ CTD Fluor |
| --- | --- | --- | --- |
| Maximum | 1.2885 | 0.8518 | 1.797 |
| Minimum | -2.4493 | -3.5004 | 0.384 |
| Average | -0.0059 | -0.1111 | 0.982 |
| Median | 0.0066 | -0.0292 | 0.963 |
| Std. Dev. | 0.4754 | 0.5377 | 0.241 |

The standard deviations in the temperature and salinity differences are very high. When the data were restricted to events 32 to 66 when differences look smoother, the results were:

|  | TSG Intake Temp – CTD Temp | TSG SAL-CTD SAL | TSG Fluor/ CTD Fluor |
| --- | --- | --- | --- |
| Median | 0.0064 | -0.0288 | 0.937 |
| Std. Dev. | 0.0866 | 0.0235 | 0.186 |

- The flow rate varied from 1.05 to 1.13. At the highest flow rate there are larger differences between CTD and TSG, but there are few instances of the higher rate and they all occurred in areas where the temperature data were very noisy.

- The intake temperature is higher than the CTD temperature by a median value of ~0.0066C° with only a very slight reduction when only the 24 stations in a section of quiet data are included. Using the 20 casts with the lowest standard deviation showed the intake temperature higher than the CTD by 0.004C° and standard deviation of 0.073C°.

- The TSG salinity is lower than the CTD salinity by a median value of 0.029psu with a standard deviation of 0.054psu. When only the 24 cases with used for the temperature study are used, the difference is the same though the standard deviation is reduced to 0.024psu. Using the 20 casts with the lowest standard deviation in TSG salinity indicates that TSG salinity is lower than CTD salinity by 0.021psu with a standard deviation of 0.008psu. From a previous cruise it is believed that even small spikes can lead to values that are low by ~0.02psu.

- The ratio of the TSG fluorescence to CTD fluorescence has a median value of 0.96 and standard deviation of ~0.24.  Using the 20 casts with the lowest standard deviation in the TSG fluorescence leads to a median difference of 1.02 and standard deviations 0.03.

- The lab temperature is higher than the intake temperature by 0.18 C° during stops, or by 0.16C° if only the 20 casts with the lowest standard deviation in the TSG intake temperature are used.

For details see 2017-23-ctd-tsg-comp.xls.

- Calibration History

The temperature and conductivity sensors were recalibrated in March 2017 and this was the second use since then.

2017-05: The temperature in the lab was higher than CTD temperatures by about 0.18Cº which looks reasonable based on expected warming in the loop. The TSG intake temperature was higher than the CTD temperature at 4m by a median value of ~0.005Cº. The TSG Salinity read much lower than the CTD salinity likely due to many bubbles. The TSG fluorescence was higher than the CTD fluorescence when the latter is <1ug/L and about 77% of the CTD fluorescence for higher values >1ug/L. Comparisons with rosette chlorophyll samples were limited but suggest that the TSG fluorescence was about 70% of chlorophyll values. No recalibration was applied to the TSG data.

2017-06: The intake temperature was higher than the CTD by about 0.004C and the salinity lower by ~0.023psu. The 2 fluorescence readings were close but values were mostly low. When underway the TSG fluorescence was close to CHL, but was higher when stopped. The range of FL values was small.

Conclusions

1. The TSG clock worked well.

2. The TSG flow rate was at the expected level and generally steady.

3. The temperature increases in the loop by roughly 0.16Cº.

4. The TSG intake temperature was higher than the CTD temperature by a median of about 0.004Cº.  This is likely as close as we can achieve given the variability at 4m depth and the fact that the TSG may draw temperature from a little higher in the water column and that there could possibly be some heating right at the intake. No correction will be applied.

5. The TSG salinity is lower than the CTD salinity by 0.029psu or 0.021psu if only cases with low standard deviation in the TSG data are used. No recalibration of salinity is justified since the comparisons are not consistent enough to lead to a reliable adjustment and the lower values are likely due primarily to bubbles which occur randomly.

6. There are many single-point one-sided spikes towards lower salinity values and these are likely due to bubbles.

7. There were no loop samples so we only have comparisons while stopped.

8. The TSG fluorescence values are very close to the CTD fluorescence readings.

9. The comparisons are generally consistent with the results of cruise 2017-06.

f.) Editing

File #1 required no editing.

File #2 was opened in CTDEDIT and used to remove some of the larger single-point spikes. Intake Temperature, Lab Temperature, Salinity and Fluorescence points were also removed from the end of the file when the flow was turned off; this took 2 passes through the editor.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number,

Temperature:Difference, Conductivity:Primary, Flag and Position:New channels.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to standardize channel names and formats.

Those files were saved as TOB files.

The Standards Check and Header Check were run; no problems were found.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

The cruise plot was added to the end of this report.

Particulars

PAR ON: Cast 1 only

CTD:

1. All bottles fired but no sampling.

2. PAR was removed and transmissometer moved from channel 2 to channel 6 to see if transmissivity signal improved.

3-end. The 10m soak was not used.

8. DO sampling cancelled.

28. Called event 27 on original rosette log sheet.

34. Hex file corrupted. Downcast ok but upcast contains only bottles 1 to 5.

37. Nutrients found on deck during cast 41 @13:38. More than 8 hours on deck.

45. Conductivity traces not lining up nicely. Salinity comparison inconclusive.

48. Jelly fish flushed out of plumbing just before the cast. Conductivity values look better.

52. Computer time adjusted to match NMEA.

55. Bit of a delay ~80m – wire leaning aft. Sample for instrument calibration - no CHE file required.

60. Winch issues @150m; went past 75m stop by accident but went back to it to fire bottle 9.

61. Partial downcast. The log indicates cast went to 710db but we have data only to 425db

66. Bottles 1, 2 and 3 fired to check for H2S but no samples – no CHE file requried.

Stopped ~1 min on downcast at ~380m to add heave compensator.

Bottle #1 lost top valve during cast

67. Hex file corrupted. Downcast ok but upcast contains only bottles 1 to 3

93. Out of order bottle firing

97. Sample numbers confusing as rosette record didn’t match labels. Sample #s 9314 to 9327 used.

106. Bottle 1 tripped for DOC-related instrument calibration – no CHE file requried.

CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 22Feb2017 | Factory |  |  |
| Conductivity | 1766 | 30Mar2017 | Factory |  |  |
| Secondary Temp. | 4484 | 22Feb2017 | Factory |  |  |
| Secondary Cond. | 3531 | 30Mar2017 | Factory |  |  |
| Transmissometer | 1201DR | 27Apr2016 | Factory |  |  |
| SBE 43 DO sensor | 1119 | 28Mar2017 | Factory |  |  |
| PAR | 70613 | 21Mar2016 | Factory |  |  |
| SeaPoint Fluor. | 3685 | 13Mar2017 |  |  |  |
| Pressure Sensor | 0550 | 24Feb2017 | Factory |  |  |
| Altimeter | 62355 | n/a | Factory |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3411       Cruise ID#:	2017-23

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 4Mar17 | Factory |  |  |
| Conductivity | 2488 | 4Mar17 | Factory |  |  |
| WetLabs Fluorometer | Ws3s-953p | May2017 |  |  |  |
| Temperature:Secondary | ? | ? |  |  |  |
| Flow meter | ? | n/a |  |  |  |