## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2017-24

Agency: PBS, Ecosystem Sciences Division, Nanaimo, B. C.

Party Chief: Neville C.		Location: Strait of Georgia

Project: Cowichan Bay PS – Juv. Salmon Strait of Georgia Project

Platform: Ocean Venture

Date: 9 May 2017 – 24 August 2017

Processed by: Germaine Gatien

Date of Processing: 13 March 2018 – 21 march 2018

Number of original HEX files: 	31 (+1 surface test only)

Number of CTD files: 31

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (s/n 0456) was used with temperature sensor #5130, conductivity sensor #3500, dissolved oxygen sensor #1592 and pressure sensor 0668.

# SUMMARY OF QUALITY AND CONCERNS

A scan of the log book and a spreadsheet giving station names, water depths, positions and times of all casts was provided.

The deployment method was to soak the CTD at the surface for 1.5 to 2.5 minutes. A wait of at least 2 to 3 minutes is suggested if you are doing a surface soak, especially if a dissolved oxygen sensor is being used. The pumps are typically turned on 45s after deployment and the DO sensor does not register a signal until about 10s later. The conductivity and DO sensors require some further time to equilibrate. In order to remove bubbles from the conductivity cell SeaBird recommend lowering the CTD to 10m, waiting a minute then returning it to near the surface. A further short wait at the surface is a good idea before running the full cast to allow turbulence created by CTD rise to dissipate. This process takes about 2minutes, so does not add to the soak time.

The choice of CTD #456 with pressure sensor #668 was not the best for this region as the pressure range was set to 2800db which is required for offshore work. Better resolution can be obtained in shallow water by using a CTD with a 1000db range. Pressure was smoothed in processing but poor resolution does reduce data quality somewhat.

For two casts there was evidence that the CTD hit bottom. Upcast data had spikes or near-constant values. Fortunately, the casts that followed looked normal so any mud must have got cleared from the conductivity cell during the upcast.

There was no salinity sampling but since the temperature and conductivity sensors had not been used since they were calibrated at the factory in late February the salinity data are likely accurate.

There was no dissolved oxygen calibration sampling, so no correction was possible. Based on previous experience with such sensors the CTD DO values are likely a little low, though not likely by more than 1 or 2% given this was the first use of the sensor since calibration in early March.

As is generally the case for small boat cruises, the headers do not contain position information or water depths and station names. This slows down processing of the data and the entry of the information is error-prone, though the spreadsheet provided was very helpful in that step. To enter such details in the headers in a way that gets converted properly into IOS Header files, use a format like:

* Latitude: 48 39.53 N

* Longitude: 123 29.96 W

* Station:ABCD

* Depth: 171

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea. The file names were standard.

2. Preliminary Steps

- The file names were in standard format.

- The hex files are missing for 2 casts listed in the log: 2017-24-0013.hex and 2017-24-0037.hex. Those casts were accidentally not logged.

- The original hex files were misnamed for 3 events; those were corrected.

- A scan of the Daily Log was obtained.

- The deployment method was to soak just under the surface at about 0.6db for at least 1.5 minutes and often longer than that. SeaBird recommend lowering the CTD to 10m, keeping it there for 1 minute, then returning to the surface. It is a good idea to wait a short time before the full cast to allow the turbulence from the CTD rise to dissipate. This process takes about 2 to 2.5 minutes. If you don’t do the 10m soak, then SeaBird advise a 3 to 5 minute wait at the surface.

- The pumps came on after 45s and the DO signal took another 10s to become non-zero. Equilibration is expected to take at least another minute.

- The cruise summary sheet was completed.

- A single configuration was file used at sea; the parameters were checked and were all correct. The file was saved as 2017-24-ctd.xmlcon.

## 3. Conversion of Raw Data

All files were converted using configuration file 2017-24-ctd.xmlcon.

Plots show that the channels all produced reasonable values. The pressures where conductivity is near zero vary from about -0.6db to +0.5db, so within the ±1db accuracy quoted by SeaBird for the pressure sensor.

The pressure resolution is low with steps on the order of 0.6db with an updated value every 6 scans. Similar results were found during 2016-58 when a different CTD was used. It is believed that this problem arose because the pressure sensor was set for a high pressure range.

4. WILDEDIT

Since there are no obvious spikes in the data below the surface, this step was skipped. Once the data are converted to IOS SHELL, plotting should be done to ensure that spikes are not a problem.

5. WFILTER

Tests were run on a few casts using cosine filters of sizes 3, 5, 7, 9 and 11 on pressure, depth, temperature and conductivity. The setting of 11 worked best on all channels, just removing unstable T-S features and the steps in pressure.

6. ALIGNCTD

Tests were run on a few casts to see what alignment made the offset between the upcast and downcast DO traces resemble that for the temperature traces. An advance of 2.5s produced the best results.

ALIGNCTD was run on all casts to advance the DO channel by 2.5s.

7. CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25 and it proved best when this equipment was last used. Tests proved hard to interpret, so the default choice was made.

CELLTM was run on all casts using the default setting: (α, 1/β) = (0.04, 8).

Casts were examined after this step and CELLTM worked well in bringing upcast and downcast T-S traces closer.

8. DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included).

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

The dissolved oxygen data are very noisy during the soak period. A first attempt at processing showed that the DELETE program chose some bad dissolved oxygen values from the soak depth. To avoid having to edit those value from the files CLIP was run to remove the first 735 records; this removed oxygen data have either zero values or are clearly not yet equilibrated. Plots were made to check that too many data were not removed and for casts #41 and 114 CLIP was rerun to remove only 690 records. This step removed all 0 values and very high DO values, but it is likely that some DO data remain where the sensor has not fully equilibrated.

CLEAN was run to add event numbers and replace pad values in the pressure channel with interpolated values.

Spreadsheet “2017-24 CTD-Bongo Data.xlsx” provided information missing from the headers, namely, positions, station names and water depth. The spreadsheet was simplified and formats adjusted. Then program “Merge CSV file to Headers” was used to add the information to the IOS Header files.

10. Checking Headers

The cross-reference check was run. The times differ from the log times by 7 hours, so are in PDT. The log times are in UTC. The start time for the 1st cast is about 20 minutes earlier than the log entry. Since it could be an error in the log or the header, no change was made to that cast. There were no other problems noted.

ADD TIME CHANNEL was run to add 7 hours from the header time.

Track plots look ok except for one station that appeared to be on land. However, a check on a map with better resolution of the coast shows that it was definitely in water and the position agrees with the log book. Track plots were produced and added to the end of this report with separate maps with event numbers and station names for each of the 4 months.

HEADER CHECK was run. Negative values are seen in many variables, but are likely from the surface. No other problems were noted.

The surface check shows an average of -0.38db but all cases with salinity > 1psu have positive surface pressure.

Plots indicate a problem with cast #27, with upcast values showing little variation. The temperature is almost constant for about 10m and then slowly begins to rise. The profile looks fine for the next cast which was run a few hours later. Cast #41 has a large salinity spike right at the bottom but it quickly recovers to expected values. For both those casts the maximum depth sampled is very close to the water depth recorded in the log, so the CTD might have hit bottom; mud may have entered the cell.

##### 11. SHIFT

Conductivity

Tests were run using values from +1 to -1 record to see which did best at removing noise from the salinity channel so that the T-S curves are just stable. No setting had any noticeable effect, so this step was skipped.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen; no further adjustment was found appropriate.

##### 12. DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.
Pressure was not filtered as it had been filtered earlier.

COMMENTS ON WARNINGS: There were no warnings.

13. DETAILED EDITING

CTDEDIT was used to remove records near the top and bottom of many casts and occasionally records corrupted by shed wakes. It was also used to clean salinity where unstable features looked likely to be caused by misalignment of T and C. All casts required light editing except for events 44, 106 and 114 which required no editing.

Notes of editing details were made in the headers.

T-S plots were examined and no significant problems were found; there are small unstable features but this is an area where they are expected.

Profile plots show that no non-zero of clearly bad near-surface dissolved oxygen data had been removed.

14. Other calibration checks

Sensor History – This was the 1st use of these sensors after factory calibration; there is no useful history.

Comparison of repeat casts –There were only 2 casts close in position and time but they are too shallow to expect good repeatability, though they show reasonable correspondence in T-S space.

Historic Ranges – The local climatology is not representative of many of the sites occupied during this study and the casts are all close to shore where a3-standard deviation climatology is not suitable. The salinity data all fell within the historic ranges. Temperatures tended to be higher than the climatology maximum in the top 15m in July and more generally in August.

There is nothing in these observations to suggest any problems with sensor calibration.

Post-cruise calibrations – None were available.

15. CALIBRATE

There was no dissolved oxygen or salinity sampling and we have no history for this instrument.

The pressure does not need recalibration.

No calibration was applied.

16. Bin Average and REMOVE

The files were bin averaged using 1db bins.

REMOVE was run to remove Scan_Number, Conductivity:Primary, Oxygen:Voltage, Descent Rate and Flag channels.

Dissolved Oxygen was derived in mass units and that was used to calculate DO saturation. Plots of near-surface saturation show values between 70% and 150%. This cruise covered a large area between May and August. The lowest values were found in August when saturation ranged from 65% to 100%. These data are not very useful in assessing the accuracy of the dissolved oxygen data in near-shore waters. For this cruise they are especially suspect because the sensor may not have fully equilibrated when the surface data were acquired. However, the values do not look unreasonable.

REORDER was used to get the 2 dissolved oxygen channels together.

17. HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and channel names and to add the following note to the headers:

Data Processing Notes:

----------------------

Dissolved Oxygen and Fluorescence are nominal and unedited, except

that some  records were removed in editing temperature and salinity.

There was no calibration information available. The sensors had been calibrated

at the factory recently, so drift is likely small.

The soak at the surface was shorter than recommended for good dissolved oxygen

equilibration, so the near-surface data may be of lower quality than usual.

For details on the processing see processing report: 2017-24-proc.doc.

A cross-reference listing was produced.

A header check was run on the CTD files and no further errors were found.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

Particulars

13. Not logged

17. Hex file originally named as 13.

20. Hex file originally named as 20.

24. Hex file originally named as 24.

37. Not logged

CRUISE SUMMARY

| Cruise ID#: 2017-24 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 5May2017 End: 24August 2017 |  |  |  |  |  |  |
| Location: Cowichan Bay |  |  |  |  |  |  |
| Party Chief: Neville C. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 25 | 0456 | No | Yes |  |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0456	Cruise ID#:		2017-24

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5130 | 23May2014 | Factory |  |  |
| Conductivity | 3500 | 23May2014 | Factory |  |  |
| SBE43 Oxygen | 1592 | 02Jun2014 | Factory |  |  |
| Pressure | 0668 | 23May2014 | Factory |  |  |