## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2017-25

Agency: PBS, Ecosystem Sciences Division, Nanaimo, B. C.

Party Chief: Neville C.		Location: Strait of Georgia

Project: Cowichan Bay PS – Juv. Salmon Strait of Georgia Project

Platform: Ocean Venture

Date: 9 May 2017 – 24 August 2017

Processed by: Germaine Gatien

Date of Processing: 12 March 2018 – 22 March 2018

Number of original HEX files: 	18 (+1 surface test only)

Number of CTD files: 18

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (s/n 0456) was used with temperature sensor #5130, conductivity sensor #3500, dissolved oxygen sensor #1592 and pressure sensor 0668.

# SUMMARY OF QUALITY AND CONCERNS

A scan of the log book and a spreadsheet giving station names, water depths, positions and times of all casts was provided.

The deployment method was to soak the CTD at the surface for 1.5 to 2.5 minutes. A wait of at least 2 to 3 minutes is suggested if you are doing a surface soak, especially if a dissolved oxygen sensor is being used. The pumps are typically turned on 45s after deployment and the DO sensor does not register a signal until about 10s later. The conductivity and DO sensors require some further time to equilibrate. In order to remove bubbles from the conductivity cell SeaBird recommend lowering the CTD to 10m, waiting a minute then returning it to near the surface. A further short wait at the surface is a good idea before running the full cast to allow turbulence created by CTD rise to dissipate. This process takes about 2minutes, so does not add to the soak time.

The choice of CTD #456 with pressure sensor #668 was not the best for this region as the pressure range was set to 2800db which is required for offshore work. Better resolution can be obtained in shallow water by using a CTD with a 1000db range. Pressure was smoothed in processing but poor resolution does reduce data quality somewhat.

There was no salinity sampling but since the temperature and conductivity sensors had been used on only 1 other cruise since they were calibrated at the factory in late February the salinity data are likely accurate.

There was no dissolved oxygen calibration sampling, so no correction was possible. Based on previous experience with such sensors the CTD DO values are likely a little low, though not likely by more than 1 or 2% given this was the first use of the sensor since calibration in early March.

As is generally the case for small boat cruises, the headers do not contain position information or water depths and station names. This slows down processing of the data and the entry of the information is error-prone, though the spreadsheet provided was very helpful in that step. To enter such details in the headers in a way that gets converted properly into IOS Header files, use a format like:

* Latitude: 48 39.53 N

* Longitude: 123 29.96 W

* Station:ABCD

* Depth: 171

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea. The file names were standard.

2. Preliminary Steps

- The file names were in standard format.

- A scan of the Daily Log was obtained.

- The deployment method was to soak just under the surface at about 0.6db for at least 1.5 minutes and usually longer than that. SeaBird recommend lowering the CTD to 10m, keeping it there for 1 minute, then returning to the surface. It is a good idea to wait a short time before the full cast to allow the turbulence from the CTD rise to dissipate. This process takes about 2 to 2.5 minutes. If you don’t do the 10m soak, then SeaBird advise a 3 to 5 minute wait at the surface.

- The pumps came on after 45s and the DO signal took another 10s to become non-zero. Equilibration is expected to take at least another minute.

- The cruise summary sheet was completed.

- A single configuration was file used at sea; the parameters were checked and were all correct. The file was saved as 2017-25-ctd.xmlcon.

## 3. Conversion of Raw Data

All files were converted using configuration file 2017-25-ctd.xmlcon.

Plots show that the channels all produced reasonable values except during the soak period. The pressures where conductivity is near zero vary from about -0.6db to +0.6db, so within the ±1db accuracy quoted by SeaBird for the pressure sensor.

The pressure resolution is low with steps on the order of 0.6db with an updated value every 6 scans. Similar results were found during 2016-58 when a different CTD was used and 2017-24 with the same CTD and pressure sensor. It is believed that this problem arose because the pressure sensor was one with a high pressure range.

4. WILDEDIT

Since there are no obvious spikes in the data below the surface, this step was skipped. Once the data are converted to IOS SHELL, plotting should be done to ensure that spikes are not a problem.

5. WFILTER

Tests were run on a few casts using cosine filters of sizes from width 9 and 11 on pressure, depth, temperature and conductivity. As found for 2017-24 the setting of 11 worked best on all channels, just removing unstable T-S features and the steps in pressure.

6. ALIGNCTD

Tests were run on 1 cast to see what alignment made the offset between the upcast and downcast DO traces resemble that for the temperature traces. An advance of 2.5s produced good results and was the setting found best for cruise 2017-24.

ALIGNCTD was run on all casts to advance the DO channel by 2.5s.

7. CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25 and it has proved appropriate for many data sets.

CELLTM was run on all casts using the default setting: (α, 1/β) = (0.04, 8).

Casts were examined after this step and CELLTM worked well in bringing upcast and downcast T-S traces closer.

8. DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included).

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

The dissolved oxygen data are very noisy during the soak period including some zero values and spikes to high values. In processing 2017-24 it was found useful to remove the first 735 records from all casts, except for a few for which the downcast had started early so that a smaller CLIP setting was needed. This step is intended to remove oxygen data that have zero values or spikes as the signal first starts.

CLIP was run to remove the first 735 records. Plots were made to check that too many data were not removed and cast #27 CLIP was rerun to remove only 675 records. This step removed all 0 values and all very high DO values, but it is likely that some DO data remain where the sensor has not fully equilibrated.

CLEAN was run to add event numbers and replace pad values in the pressure channel with interpolated values.

Spreadsheet “2017-25 CTD-Bongo Data.xlsx” provided information missing from the headers, namely, positions, station names and water depth. The spreadsheet was simplified and formats adjusted. The entries were compared with the log entries and 1 position entry and 2 water depth entries had small errors which were corrected. Then program “Merge CSV file to Headers” was used to add the information to the IOS Header files.

10. Checking Headers

The cross-reference check was run. The times differ from the log times by 7 hours, so are in PDT. The log times are in UTC. There were no other problems noted.

ADD TIME CHANNEL was run to add 7 hours to the header time.

Track plots were produced and look fine. They were added to the end of this report with separate maps with event numbers and station names for each of the 2 months.

HEADER CHECK was run. Negative values are seen in pressure, but those are likely from the surface. No other problems were noted.

The surface check shows an average of 0.8db. The only case with pressure <0 has a very low salinity.

##### 11. SHIFT

Conductivity

Based on tests run for cruise 2017024 no shift was applied.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen; no further adjustment was found appropriate.

##### 12. DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.
Pressure was not filtered as it had been filtered earlier.

COMMENTS ON WARNINGS: There were no warnings.

13. DETAILED EDITING

Plots show that editing will be needed near the surface.

CTDEDIT was used to remove records near the top and bottom of many casts and was also used to clean salinity for a few casts. Editing was required for all events except for #12 and #40.

Notes of editing details were made in the headers.

T-S plots were examined and no significant problems were found; there are small unstable features but this is an area where they are expected.

The casts from July fall into quite a tight group while the June casts show much more variability.

Profile plots show that all non-zero or clearly bad near-surface dissolved oxygen data had been removed.

14. Other calibration checks

Sensor History – This was the 2nd use of these sensors after factory calibration; there was no calibration sampling during the previous use.

Comparison of repeat casts –The only casts close in position and time are too shallow to expect good repeatability, though they show reasonable correspondence in T-S space.

Historic Ranges – The 3-standard-deviation climatology available is considered too severe for sites close to shore. The temperature data all fell within the climatology. Salinity values tended to be lower than the climatology maximum near the bottom of casts in the central part of the survey areas, especially in June.

There is nothing in these observations to suggest any problems with sensor calibration.

Post-cruise calibrations – None were available.

15. CALIBRATE

There was no dissolved oxygen or salinity sampling during this cruise or the previous one. We have no calibration history for this instrument since it was last serviced at the factory.

The pressure does not need recalibration.

No recalibration was applied.

16. Bin Average and REMOVE and DO saturation study

The files were bin averaged using 1db bins.

REMOVE was run to remove Scan_Number, Conductivity:Primary, Oxygen:Voltage, Descent Rate and Flag channels.

Dissolved Oxygen was derived in mass units and that was used to calculate DO saturation. Plots of near-surface saturation show values between 105% and 150%. The highest values were observed in June while on July the range was 105% to 130%. The values do not look unreasonable for this inshore region.

REORDER was used to get the 2 dissolved oxygen channels together.

17. HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and channel names and to add the following note to the headers:

Data Processing Notes:

----------------------

----------------------

Dissolved Oxygen data are nominal and unedited, except that some records were

removed in editing temperature and salinity.

There was no calibration information available. The sensors had been calibrated

at the factory fairly recently, so drift is likely small.

The soak at the surface was sometimes shorter than recommended for good dissolved

oxygen equilibration, so the near-surface data may be of lower quality than usual.

For details on the processing see processing report: 2017-25-Processing_Report.doc.

A cross-reference listing was produced.

A header check was run on the CTD files and no errors were found.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

CRUISE SUMMARY

| Cruise ID#: 2017-25 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 5 June2017 End: 18July 2017 |  |  |  |  |  |  |
| Location: Strait of Georgia |  |  |  |  |  |  |
| Party Chief: Neville C. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 25 | 0456 | No | Yes |  |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0456	Cruise ID#:		2017-25

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5130 | 23May2014 | Factory |  |  |
| Conductivity | 3500 | 23May2014 | Factory |  |  |
| SBE43 Oxygen | 1592 | 02Jun2014 | Factory |  |  |
| Pressure | 0668 | 23May2014 | Factory |  |  |