## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2017-42

Agency: PBS, Ecosystem Sciences Division, Nanaimo, B. C.

Party Chief: King J. / Boldt J.		Location: West Coast Vancouver Island

Project: Integrated Pelagics Ecosystem Survey

Platform: Sea Crest

Date: 8 July 2017 – 2 August 2017

Processed by: Germaine Gatien

Date of Processing: 5 April 2018 – 10 April 2018

Number of original HEX files: 43 (8 missing)		Number of CTD files: 43

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (# 0404) was used with temperature sensor #2449, conductivity sensor #1764, Wetlabs ECO Fluorometer #4185, dissolved oxygen sensor #1176, pH sensor #0852 and pressure sensor #0464.

# SUMMARY OF QUALITY AND CONCERNS

A scan of the log book and a spreadsheet giving positions, water depths and times of all casts were provided.  Station names were missing from the list, but were added based on log entries.

The deployment method varied. For casts #1001-1033 the CTD was soaked at about 5 or 6m for 1 to 2minutes, but the CTD was not then returned to the surface so a full cast was not obtained. The wait was also shorter than recommended and in at least one case the cast began before the pumps had turned on. For casts #1034 to 1045 the soak was at about 2m and the wait was also only 1 to 2minutes. The factory recommendation is that the CTD should be lowered to 10m to remove bubbles from the system, then returned to the surface. Then after a short wait the full cast is run. If there is a good reason to not do the deep soak, then a wait of 3 minutes is recommended. The pumps do not come on until the CTD has been in the water for at least 45s and equilibration takes some time after that. Even an unpumped sensor like the pH needs a few minutes to equilibrate.

As is generally the case for small boat cruises, the headers do not contain position information or water depths and station names. This slows down processing of the data and the entry of the information is error-prone, though the spreadsheet provided was helpful in that step. To enter such details in the headers in a way that gets converted properly into IOS Header files, use the following format when prompted by the acquisition routine:

Latitude: 49 3.345 N

Longitude: 123 22.372 W

Station: 41

Depth: 245 m

Having a template on the desktop so you can quickly update entries, then copy and paste avoids the problem of the acquisition system timing out.

There was no salinity or dissolved oxygen calibration sampling. Results of an inter-ship comparison during a previous cruise using this equipment suggests that the salinity calibration is good, but dissolved oxygen values are likely low by up to 4%.

The protective cover on the fluorometer was left on for most of the casts. There is a plan to avoid this error in future by putting some tape on the cap to make it obvious when it is left on.

Some cast files were lost from the end of the cruise. Whether they were never downloaded or have been mislaid is unknown.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea.

The event numbers were assigned in a non-standard way with a different set of numbers for different sorts of activities.

The file names were non-standard. While there are event numbers in the log the file names include the station names instead and those sometimes contain more than 4 digits.

##### Preliminary Steps

A scan of the Daily Log was obtained.

The cruise summary sheet was completed.

No configuration file was provided but the same equipment was used for several other cruises, so that was copied and renamed 2017-42-ctd.xmlcon.

The files were copied to a folder where they were renamed using the event numbers rather than station names (which were all numbers). The event numbers were chosen to match the log entries; they have a leading 1, ex. 1001, 1002 …

Some files were missing – events 44-51 from July 29, 30 and 31. Two extra files were found – one was a copy of the file for event 1027 and the other was a test cast from June 2017.

##### Conversion of Raw Data

All files were converted using configuration file 2017-42-ctd.xmlcon.

Plots were examined and all expected channels were present but a few problems were found:

- The fluorescence data look bad for most casts, likely because the protective cap was left on the sensor. Late in the cruise the data look ok. That channel will be removed later from casts 1-35.

- The CTD was soaked at ~5m for many casts but not returned to the surface to get a full profile.

- Later casts were soaked at the surface but the wait was also short there.

The deployment method was to soak the CTD at about 5 or 6m  or 2m for 1 to 2minutes. The wait was  shorter than recommended and in at least one case the cast began before the pumps had turned on. The factory recommendation is that the CTD should be lowered to 10m to remove bubbles from the system, returned to the surface and then the full cast is run. If there is a good reason to not do the deep soak, then a wait of 3 minutes is recommended. The pumps do not come on until the CTD has been in the water for at least 45s and equilibration takes some time after that. Even an unpumped sensor like the pH needs a few minutes to equilibrate.

The pressures where conductivity drops suddenly during upcasts ranged from -0.3 and +0.6db; ±1db is the accuracy quoted by SeaBird for the pressure sensor so this looks reasonable.

##### WILDEDIT

Since there are no obvious spikes in the data below the surface, this step was skipped. The near-surface spikes will be removed in the DELETE step or using CTDEDIT.

##### FILTER

WFILTER was run with cosine file width 5. That setting worked well on the pressure, depth, temperature and conductivity during past cruises using this equipment. There remain some pressure reversals but they are likely real given these casts were from the offshore area. The resulting salinity looks much improved.

##### ALIGNCTD

During 2016-40, 2017-41, 2017-52 and 2017-53 when this DO sensor was also used the DO channel was advanced by 2.5s. That setting looks appropriate for these data as well.

ALIGNCTD was run to apply a 2.5s advance.

##### CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25 and it has proved appropriate for many data sets.

CELLTM was run on all casts using the default setting: (α, 1/β) = (0.04, 8).

Casts were examined after this step and CELLTM worked well in bringing upcast and downcast T-S traces closer.

##### DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included).

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers. This was done in 2 groups so the chief scientist could be entered differently for the two legs of the cruise.

Spreadsheet 2017-42CTD_BridgeLog.xlsx was provided. This contained station locations and depths, but not station names. The spreadsheet was simplified and station names were added, the format of positions was adjusted and file names were added based on event numbers. The file was saved as 2017-42 merge-hdr.csv.

Program “Merge CSV file to Headers” was used to add the information to the IOS Header files.

CLEAN was not run since the event numbers have already been added and there are no pad values in the pressure channel.

##### Checking Headers

The cross-reference check was run. The times were wrong for 2 casts showing they had been mis-identified when they were renamed. After fixing the file names the check was rerun and no further errors were found. The times appear to all be in UTC.

CLIP was used to remove initial records from the soak period to ensure that data are only kept from after the pumps came on and the sensors were able to equilibrate. Because some of the casts had very short waits only 400 scans were removed to start with. This will leave some bad data near the surface. Where obvious, this will be removed at the CTDEDIT stage.

Track plots were produced using event numbers and station names. They were added to the end of this report.

HEADER CHECK was run. Negative values are seen in pressure and fluorescence but may be removed by steps to follow.

The surface check shows an average of 0.7db on files before CLIP was applied. The associated salinity values are mostly extremely low.

##### SHIFT

Conductivity

Tests were run on a few casts using a variety of shifts to align conductivity. A setting was found that improved a few unstable features, but it made many others worse. As has been found during other cruises using this equipment, there was no setting that improved most casts.

SHIFT was not applied to conductivity.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen; there is no evidence that further adjustment is needed but it is difficult to judge.

pH

Tests were run on the alignment of the pH traces with temperature. An advance of 10 or 15 records was found to improve the alignment on other 2017 cruises using this equipment. For this cruise +10 records looked best. SHIFT was run to advance the pH signal by 10 records.

##### DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.
Pressure was not filtered as it had been filtered earlier.

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

Plots show that editing will be needed near the surface.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records near the top and/or bottom of many casts and a few records corrupted by shed wakes. It was also used to clean salinity for some casts.

Dissolved Oxygen data were removed from the tops of casts where it was clear that the sensor had not yet equilibrated.

Editing was applied for all events except for #1004.

Notes of editing details were made in the headers.

The edited files were copied to *.EDT.

T-S plots were examined. There are some small unstable features but they are small and possibly real.

Profile plots showed the bad fluorescence data in casts 1001 to 1035. There are also some very slightly negative fluorescence values at depth in casts 1036-1043. That will be dealt with later.

The only other problem was bad near-surface dissolved oxygen data in cast #1024; it was edited again.

##### Other calibration checks

Sensor History – The sensors were used on 3 previous 2017 cruises. There was no calibration sampling during those cruises. There was one inter-ship comparison in June 2017 that suggested that the salinity accuracy was good and that dissolved oxygen was probably reading low.

Comparison of repeat casts –None available.

Historic Ranges – The temperature data were all within the local climatology except for a small section with slightly high values around 20m at the most southerly cast. Salinity values were low in the top 5 to 20m of the first 3 casts close to Johnstone Strait; all other values were within the climatology.

Post-cruise calibrations – None were available.

##### CALIBRATE

There was no dissolved oxygen or salinity sampling during this cruise.

An inter-ship comparison of 1 cast from 2017-52 with a cruise that included calibration sampling suggests the salinity values are good but that DO values are low. However, the results were not clear enough to justify recalibration.

The pressure does not need recalibration.

##### Filter

The fluorescence channel was filtered using a median filter, size 5. This removed large spikes but a few negative values remained at depth.

CALIBRATE was run using file 2017-42-recal1.ccf to replace negative fluorescence values with value 0; all other fluorescence values were left unchanged.

##### Bin Average and REMOVE and DO saturation study

The files were bin averaged using 1db bins.

REMOVE was run to remove Scan_Number, Conductivity:Primary, Oxygen:Voltage, Descent Rate and Flag channels. Channel Fluorescence:URU:Wetlabs was removed from casts #1001-1035.

Dissolved Oxygen was derived in mass units. There is no useful near-surface oxygen, so saturation rates were not examined.

REORDER was used to get the 2 dissolved oxygen channels together.

##### HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and channel names and to add the following note to the headers:

Data Processing Notes:

----------------------

The length of the surface soaks varied greatly, with some not being long enough

for the pumps to come on and equilibration of pumped channels to be achieved.

Dissolved Oxygen and pH:SBE data are nominal and unedited, except that some

records were removed in editing temperature and salinity and some bad

near-surface dissolved oxygen data were removed.

Channel Fluorescence:URU:Wetlabs was removed from casts #1001-1035 because

the protective cap had not been removed. The data that remain are nominal.

Some small negative values in fluorescence were set to 0 using program

Calibrate.

There was no calibration sampling available from this cruise. An inter-ship

comparison of 2 casts from a June 2017 cruise using these sensors with a cast

from a cruise that did include calibration sampling, suggested that the salinity

values from this CTD were accurate but that dissolved oxygen values were likely

low by up to 4%.

For details on the processing see processing report: 2017-42-Processing_Report.doc.

A cross-reference listing was produced.

A header check was run on the CTD files. No problems were found.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

##### Bottle File Preparation

There was nutrient and chlorophyll sampling at a nominal depth of 10m. Examination of some profiles show a depth of 11m as the most common level for stops. Assuming that the Niskin bottle was 1m above that makes 10m a reasonable estimate. Upcast CTD data are not very reliable given that the CTD would be just beginning its upward movement and would pass through the wake of the CTD. Generally downcast data provide a better comparison with bottle samples.

CTD data from 10m were extracted from bin-averaged downcast files to a spreadsheet; date, time, water depth, station names and location information were also extracted. The fluorescence data were not included because, as noted earlier, the values indicate that the protective cap was not removed.

No CTD data were found above 13m for event #25.

The data from nutrient and chlorophyll analysts were added to the spreadsheet.

A 6-line header was added and the files were converted to IOS header format.

CLEAN was run to add 0 flags for any cases with no quality flag and to create a start time.

Remove was used to remove Time and Date channels.

Header Edit was used to add header comments and to remove END TIME and TIME ZERO which were created during CLEAN are identical to the Start Time.

The final bottles are name 2017-42-****.BOT.

No comparison of fluorescence and chlorophyll was possible as there was no sampling from the casts with usable fluorescence.

Particulars

1001-1035. Fluorometer cap left on.

1025. CTD was at about 9.5m of the downcast when the pump would have turned on.

CRUISE SUMMARY

| Cruise ID#: 2017-42 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 8 July 2017 End: 2 August 2017 |  |  |  |  |  |  |
| Location: West Coast Vancouver Island |  |  |  |  |  |  |
| Party Chief: King J. (events 1001-1030) and Boldt J. (events 1031 – 1039) |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 25 | 0404 | No | Yes |  |

CTD CALIBRATION INFORMATION

ake/Model/Serial#:	SEABIRD/SBE25/0404	Cruise ID#:		2017-42

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2449 | 16Dec2015 | Factory |  |  |
| Conductivity | 1764 | 23Feb2017 | Factory |  |  |
| pH | 0852 | 8Mar2017 | Factory |  |  |
| ECO Fluorometer | 4185 | 7Mar2017 | Factory |  |  |
| SBE43 Oxygen | 1176 | 2Mar2017 | Factory |  |  |
| Pressure | 482 | 28Feb2017 | Factory |  |  |