## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2017-53

Agency: PBS, Ecosystem Sciences Division, Nanaimo, B. C.

Party Chief: Neville C.		Location: Discovery Islands

Project: Discovery Islands Purse Seine

Platform: Sea Crest

Date: 11 September 2017 – 26 September 2017

Processed by: Germaine Gatien

Date of Processing: 29March 2018 – 5 April 2018

Number of original HEX files: 41		Number of CTD files: 30

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (# 0404) was used with temperature sensor #2449, conductivity sensor #1764, Wetlabs ECO Fluorometer #4185, dissolved oxygen sensor #1176, pH sensor #0852 and pressure sensor #0464.

# SUMMARY OF QUALITY AND CONCERNS

A scan of the log book and a spreadsheet giving station names, water depths, positions and times of all casts were provided.

11 files contained only surface data. For 10 this appears to be due to a low battery; after that group of bad files, the battery was changed. Raw file 2017-53-0180 also contained only surface data, but it appears that the download was repeated successfully but saved as 2017-53-0185. The file for event 185 was saved as 2017-53-0190. The last event file was saved with the correct name, but in the confusion the file from event #190 was lost. Maximum sampling depths and the instrument time stamp enabled proper identification of the misnamed files.

The deployment method was to soak the CTD at the surface for from 1.5 minutes to 4.5 minutes with the majority of soaks between 2.5 and 3.5minutes. A wait of at least 2 to 3 minutes is suggested if you are doing a surface soak, and longer in low salinity water or if a dissolved oxygen sensor is being used. The pumps are turned on 45s after deployment only if the minimum conductivity has been experienced. The DO sensor does not register a signal until about 10s later and the conductivity and DO sensors require some further time to equilibrate, so 3 minutes is a minimum in those conditions. The pH sensor appeared to have not fully equilibrated after 2.5minutes during 1 cast.

In order to remove bubbles from the conductivity cell SeaBird recommend lowering the CTD to 10m, waiting a minute then returning it to near the surface. This will also ensure that the pumps turn on so you can get good near-surface data. A further short wait at the surface is a good idea before running the full cast to allow turbulence created by CTD rise to dissipate. This process takes about 2 or 3 minutes, so does not add to the recommended wait time.

There was no salinity or dissolved oxygen calibration sampling. Results of an inter-ship comparison during a previous cruise using this equipment suggests that the salinity calibration is good, but dissolved oxygen values are likely low by up to 4%.

As is generally the case for small boat cruises, the headers do not contain position information or water depths and station names. This slows down processing of the data and the entry of the information is error-prone, though the spreadsheet provided was very helpful in that step. To enter such details in the headers in a way that gets converted properly into IOS Header files, use a format like the following when prompted by the acquisition routine:

Latitude: 49 3.345 N

Longitude: 123 22.372 W

Station: 41

Depth: 245 m

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea. The file names were standard.

##### Preliminary Steps

A scan of the Daily Log was obtained.

The cruise summary sheet was completed.

The configuration file used at sea was checked and the only correction made was to the offset for the fluorometer. The file was saved as 2017-53-ctd.xmlcon.

##### Conversion of Raw Data

All files were converted using configuration file 2017-53-ctd.xmlcon.

Plots were examined and show that all expected channels are present but some problems were found:

- The fluorescence looks bad for some casts; the log indicates that the protective cap was left on the sensor for casts 2 to 41. The same appears to have happened during casts #49 and 55 as well. That channel will be removed later from those casts.

- 11 files (events 93 to 141 and 180) contain only surface data. The battery was changed after the first group of 10 bad files. The other one with only surface data was #180. That will be investigated later.

The deployment method was to soak the CTD at the surface for from 1.5 minutes to 4.5 minutes with the majority of soaks between 2.5 and 3.5minutes. A wait of at least 2 to 3 minutes is suggested if you are doing a surface soak, and longer in low salinity water or if a dissolved oxygen sensor is being used. The pumps are turned on 45s after deployment only if the minimum conductivity has been experienced. The DO sensor does not register a signal until about 10s later and the conductivity and DO sensors require some further time to equilibrate, so 3 minutes is a minimum in those conditions. The pH sensor had not equilibrated fully for a few casts that had soaked for 2.5 minutes.

In order to remove bubbles from the conductivity cell SeaBird recommend lowering the CTD to 10m, waiting a minute then returning it to near the surface. This will also ensure that the pumps turn on so you can get good near-surface data. A further short wait at the surface is a good idea before running the full cast to allow turbulence created by CTD rise to dissipate. This process takes about 2 or 3 minutes, so does not add to the wait time.

The pressures where conductivity just starts to register non-zero values varies from about -0.6 to +1.2db, with most values between -0.2 and +0.5db; ±1db is the accuracy quoted by SeaBird for the pressure sensor.

##### WILDEDIT

Since there are no obvious spikes in the data below the surface, this step was skipped. The near-surface spikes will be removed in the DELETE step or using CTDEDIT.

##### FILTER

WFILTER was run with cosine file width 5. That setting worked well on the pressure, depth, temperature and conductivity during past cruises using this equipment. The resulting salinity looks improved.

##### ALIGNCTD

During 2016-40, 2017-41 and 2017-52 when this sensor was used the DO channel was advanced by 2.5s.

That setting looks appropriate for these data as well.

ALIGNCTD was run to apply a 2.5s advance.

Plots made after this step show some improvement in matching features in downcasts.

##### CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25 and it has proved appropriate for many data sets.

CELLTM was run on all casts using the default setting: (α, 1/β) = (0.04, 8).

Casts were examined after this step and CELLTM worked well in bringing upcast and downcast T-S traces closer.

##### DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included). At this point the 11 casts with only surface data were removed from the processing stream.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

CLEAN was run to add event numbers and replace pad values in the pressure channel with interpolated values.

Spreadsheet “2017-53 CTD-Bongo Data.xlsx” provided information missing from the headers, namely, positions, station names and water depth. The spreadsheet was simplified and formats adjusted. Entries were compared with the log entries and 1 position entry and 2 water depth entries had small errors which were corrected.

Program “Merge CSV file to Headers” was used to add the information to the IOS Header files.

##### Checking Headers

The cross-reference check was run. The times differ from the log times by 7 hours for all events except #185 and #190. Some files appear to be misnamed.

- The file named #185 (it is 186 in the spreadsheet but it is 185 in the log) appears to really be #180. There was a download of 180 with just a few records in it – this looks like a second attempt. The start time is Sep 24 2017 13:34:11 PDT.  That time is close to the log time for #180 – 7 hours and the depth is ok.

- The file named #190 looks like it is really 185. The start time is Sep 24 2017 17:57:55 which again matches the log – 7 hours. The depth is wrong for #190 but right for #185.

- The files named 194 and 196 look ok, with start times and depths that match the log – 7 hours, so that leaves 190 as missing.

- File #190 would have been 110m depth according to the log, and no such file was found. Perhaps it got erased when #185 was saved as #190.

The *.ios files were renamed accordingly, CLEAN and MERGE Headers steps were rerun and a new cross-reference list was produced.

ADD TIME CHANNEL should have been run at this point to add 7 hours to the header time, but was missed. It was run near the end of processing.

CLIP was used to remove initial records from the soak period to ensure that data are only kept from after the pumps came on and the sensors were able to equilibrate. Plots were made to determine how many records to remove from each file. The number of records removed varied from 700 to 2200.

Plots made after this step show that sufficient data were left to cover the complete downcast.

Track plots were produced using event numbers and station names. They were added to the end of this report.

HEADER CHECK was run. Negative values are seen in pressure and fluorescence but may be removed by steps to follow.

The surface check shows an average of 0.6db on files before CLIP was applied. The associated salinity values are mostly extremely low.

Cast #88 is incomplete, stopping at 43db.

##### SHIFT

Conductivity

Tests were run on a few casts using a variety of shifts to align conductivity. A setting was found that improved one cast, but it made all others worse. There was no setting that improved most casts.

Similar results were found for other cruises using these sensors.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen; there is no evidence that further adjustment is needed but it is really impossible to judge due to very low vertical variability.

pH

Tests were run on the alignment of the pH traces with temperature. An advance of 15 records improved the alignment. SHIFT was run to advance the pH signal by 15 records.

##### DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.
Pressure was not filtered as it had been filtered earlier.

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

Plots show that editing will be needed near the surface.

All DEL files were copied to *.EDT.

CTDEDIT was used to remove records near the top and/or bottom of many casts and a few records corrupted by shed wakes. It was also used to clean salinity for a few casts.

Editing was applied for all events except for #13 and 148.

Notes of editing details were made in the headers.

The edited files were copied to *.EDT.

T-S plots were examined. There are some small unstable features but these are casts close to shore where such features are common.

Profile plots look reasonable except that the fluorescence traces show that some slightly negative values remain. These are due to noise in the fluorescence channel when that channel is reading very close to 0.

Filtering of the fluorescence looks like a good idea as there are many large spikes and negative values. Any negative values that survive that step can be fixed later using CALIBRATE.

##### Other calibration checks

Sensor History – The sensors were used on 3 other 2017 cruises. There was no calibration sampling during those cruises.

Comparison of repeat casts –None available.

Historic Ranges – The temperature data were all within the local climatology except for a few casts with higher values near the surface. Salinity values were low around 50m for a few casts on the eastern side of the Strait of Georgia. These excursions are to be expected in near shore data as the climatology is not well suited to near-shore areas and it is somewhat out of date.

Post-cruise calibrations – None were available.

##### CALIBRATE

There was no dissolved oxygen or salinity sampling during this cruise.

An inter-ship comparison of 1 cast from 2017-52 with a cruise that included calibration sampling suggests the salinity values are good but that DO values are low. However, the results were not clear enough to justify recalibration.

The pressure does not need recalibration.

##### Filter

The fluorescence channel was filtered using a median filter, size 5. This removed large spikes but a few negative values remained.

CALIBRATE was run using file 2017-53-recal1.ccf to replace negative fluorescence values with value 0; all other fluorescence values were left unchanged.

##### Bin Average and REMOVE and DO saturation study

The files were bin averaged using 1db bins.

REMOVE was run to remove Scan_Number, Conductivity:Primary, Oxygen:Voltage, Descent Rate and Flag channels. Channel Fluorescence:URU:Wetlabs was removed from casts #2-41 and #49-55.

Dissolved Oxygen was derived in mass units and that was used to calculate DO saturation. Plots of near-surface saturation show values between 50% and 120%. The lowest values were the Strait of Georgia, Gulf Islands and Discovery Passage; the DO profiles show these casts were well-mixed vertically as is typical of these areas. The highest values were in the northern part of the Strait of Georgia where the DO gradients were higher. The values are reasonable but only in a rough sense.

ADD TIME CHANNEL was run to add 7 hours to the start time.

REORDER was used to get the 2 dissolved oxygen channels together.

##### HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and channel names and to add the following note to the headers:

Data Processing Notes:

--------------------------------------------

Dissolved Oxygen and pH:SBE data are nominal and unedited, except that some

records were removed in editing temperature and salinity.

Channel Fluorescence:URU:Wetlabs was removed from casts #2-41 & #49-54 because

the protective cap had not been removed. The data that remain are nominal.

Some small negative values in fluorescence were set to 0 using program

Calibrate.

The raw files from events #93 to #141 contained only surface data and were not

processed. The data from event #190 were lost.

There was no calibration sampling available from this cruise. An inter-ship

comparison of 2 casts from cruise 2017-52 with a cast from a cruise that did

include calibration sampling, suggested that the salinity values from this

CTD were accurate but that dissolved oxygen values were likely low by 2-4%.

The length of the surface soaks varied greatly, with many being long enough

for the pumps to come on and equilibration of pumped channels to be achieved.

For some the soak was too short for equilibration to occur before the CTD had

started the downcast.

For details on the processing see processing report: 2017-53-Processing_Report.doc.

A cross-reference listing was produced.

A header check was run on the CTD files. No problems were found.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

Particulars

2-41. Fluorometer cap left on.

49-54. Fluorometer cap left on.

85. End of upcast missing – probably battery running low.

88. Partial downcast only – probably battery running low. File ends at 42.9db.

93-141. Only surface data acquired – battery problem. Not processed.

180 & 185. Hex file 180 surface data only. Hex file names 185 and 190 have full casts for 180 and 185.

190. Hex file missing – likely copied over in confusion of file names.

CRUISE SUMMARY

| Cruise ID#: 2017-53 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 11 September 2017 End: 26 September 2017 |  |  |  |  |  |  |
| Location: Strait of Georgia |  |  |  |  |  |  |
| Party Chief: Neville C. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 25 | 0404 | No | Yes |  |

CTD CALIBRATION INFORMATION

ake/Model/Serial#:	SEABIRD/SBE25/0404	Cruise ID#:		2017-53

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2449 | 16Dec2015 | Factory |  |  |
| Conductivity | 1764 | 23Feb2017 | Factory |  |  |
| pH | 0852 | 8Mar2017 | Factory |  |  |
| ECO Fluorometer | 4185 | 7Mar2017 | Factory |  |  |
| SBE43 Oxygen | 1176 | 2Mar2017 | Factory |  |  |
| Pressure | 482 | 28Feb2017 | Factory |  |  |