## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2017-56

Agency: Ecosystem Sciences Division

Location: Baynes Sound

Project: Baynes Sound Shellfish Project

Party Chief: Sutherland T.

Platform: Vector

Date: June 27, 2017 – July 2, 2017

Processed by: Germaine Gatien

Date of Processing: 4 December 2017 – 11 December 2017

Number of original HEX files: 33 (No useful data in one file)

Number of CTD files: 32	Number of bottle files: 21

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a SBE 43 DO sensor (#1119), a SeaPoint Fluorometer (#3865) with a 3X cable, a Biospherical QSP-400 PAR sensor (#4565), Surface PAR (#20518) a pH sensor (#0691), an altimeter (#62355)and a Wetlabs CSTAR transmissometer (#1396DR).

The data acquisition program was Seasave 7.26.2.13.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

An SBE21 Thermoslinograph was used. There was no external temperature sensor or flow meter.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log book was missing a personnel list.

There was no duplicate sampling during this cruise.

File 2017-56-0058 contained no useful data, just some surface records with pumps off.

Salinity analysis was done within a month. There were only 3 samples from the bottom bottles, all casts were shallow and this is an areas where flushing of Niskin bottles is expected to be incomplete, so we do not expect good results from the comparison of bottle salinity with CTD salinity. The comparison indicated that the primary salinity was high by ~0.124psu and the secondary high by ~0.126psu; those values are way out of line with the history of these sensors. Fortunately, cruise 2017-06 sampled offshore waters less than a month before this cruise, had extensive bottle sampling in deep waters and used the same sensors, so those results were applied to these data. We do not usually have the luxury of such excellent sampling to guide calibration decisions for inshore cruises.

The calibration of the dissolved oxygen sensor was also based on the results of the bottle comparison during 2017-06. There is some evidence that this is appropriate for this cruise. For details see section 5.

There were serious problems with the primary system, with clearly bad data in primary salinity. The CTD dissolved oxygen sensor was also mounted on the primary pump, and those data frequently looked bad as well. Secondary temperature and salinity channels were selected for archiving, while dissolved oxygen data were removed where there were clear signs of trouble. The quality of data in channel Oxygen:Dissolved:SBE must be considered lower than usual.

No estimate of accuracy has been made due to low confidence in the comparison with bottles.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained; no problems were noted.

Nutrients, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The sensor histories for the pressure, conductivity and DO sensors were found.

The calibration constants were checked for all instruments and the only change made was to factor E in the dissolved oxygen calibration parameters; tests run during cruise 2017-06 found that E= 0.038 worked better than 0.036. This does not affect these data as all sampling was shallow, but it was changed for consistency. One file was saved as 2017-56-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using the configuration files mentioned in the previous section to create CNV files.

A few casts were examined.  All expected channels are present.

- The primary conductivity is bad for many casts, mostly during downcasts. It looks like there was some obstruction that usually cleared during the cast.

- Dissolved oxygen profiles show a similar problem and it was mounted on the primary pump.

- For one cast the secondary conductivity looked poor during a small part of the upcast.

- Cast #58 has only surface data thought the log indicates it sampled to 50m.

- As seen in other recent cruises using this transmissometer there are odd profiles that look as though values got stuck for a half-second or so, then registered a few new values with the usual sampling rate then got stuck again. The profiles look reasonable.

- PAR and pH look normal. Surface PAR sometimes looks odd but sudden changes could be real. At depths very close to 0m the PAR values are generally much higher than the SPAR values.

As usual the descent rate was usually fairly high and steady, but there are some noisier casts.

##### BOTTLE FILE PREPARATION

ROS files were created using files 2017-56-ctd.xmlcon.

No ROS files were produced for casts #29, 31, 32, and 33. The profile shows that the CTD did stop for bottles. There are no BL files for those casts, nor any HDR or XMLCON files. However, there were BL, HDR and XMLCON files for events 8 to 12 and those events were actually GRAB events. Apparently the files were incorrectly named for events 29 – 33; someone later renamed the HEX files, but not the others. The BL, HDR and XMLCON files were renamed and conversion was then successful.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and a few outliers were found in cast #45. CTDEDIT was used to clean channel Salinity:T0:C0 in that file. The output file was copied to *.BOT.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv.

Sample numbers were added to the file based on the rosette log records.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine. Those files were bin averaged on bottle number.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2017-56-bot-hdr.txt which will be updated as needed during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2017-56oxy*.xlsx which includes flags and comments. Draw temperatures are available. There was no duplicate sampling. The spreadsheet page with the final data was simplified and the file was then saved as 2017-56oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in QF2017-56SAL*.xlsx. There were no duplicates. The analysis was done within 27-31 days of collection. There was no duplicate sampling. The files were simplified and saved as 2017-56sal.csv.

That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet NUTS_QF2017-56*.xlsx. There was no duplicate sampling.  The spreadsheet was simplified, reordered on sample numbers and saved as 2017-56nuts.csv. Lines for surface samples without sample numbers were removed.

The file was converted to individual NUT files.

The SAL, OXY and NUT files were merged with CST files in 3 steps.

After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only, with output files named MRGCLN1.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number as the merge channel.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets. A few discrepancies were found.

- The DO labels for Event #3 had been mislabeled as #2.

- The sample numbers for Event #48 were wrong on the analysis log sheet.

- There is no flag for one DO sample that had been padded. It was given a 1 flag.

When those errors were fixed and the merges rerun, the data were exported again and no further problems were found.

CLEAN was run on the MRG files to add 0 flags to empty flag channels and to update header limits.

##### Compare

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

It is not great surprise that the comparison looks very different from the usual fit found for differences versus CTD DO. Calibration drift in DO sensors almost always tends towards low values. But we expect flushing of Niskin bottles to be inefficient in this area of protected waters that don’t shake up the bottles. The water in the Niskin bottles is likely to be representative of water from lower in the water column with lower DO than in ambient waters (except for bottles at the bottom of the cast). This makes the CTD DO values look higher relative to samples than it normally would. The flushing errors mask that drift somewhat. The error associated with poor flushing is most severe in the presence of high DO gradients such as those seen at the surface in this region; that effect is made more complicated by frequent DO gradient reversals near the surface. The bottom bottles differ from the others in that poor flushing leads to errors with the opposite sign so that the CTD DO looks lower than it really is.

A second problem for this cruise is that the primary CTD plumbing showed some severe problems that affected primary conductivity and dissolved oxygen. The problems are very bad during some downcasts but appeared to clear up at about 25m and the upcasts look ok. But there may be some reduction in quality that is not apparent. Another limitation in this area is the small range of DO values making any attempt to find an offset impossible.

When a fit is made through the origin, the slope is 1.0242 using all data.

The slope is 1.0196 if bottom bottles are excluded.

The slope is 1.0759 if only bottom bottles are included.

The slope is 1.0428 if bottom bottles and cases with (CTD-Bottle) >0 are excluded`.

The latter is the closest to observations made during the previous 2 cruises that used this DO sensor in areas where the results are considered more reliable.

Juan de Fuca Strait during 2017-64

CTD DO Corrected = CTD DO * 1.0385 + 0.059    R² = 0.94 (4)

Offshore during 2017-06

CTD DO Corrected = CTD DO * 1.0399 + 0.065    R2 = 0.96 (2017-06)

For more detail see document 2017-56-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No significant outliers were found.

Salinity

Compare was run with pressure as reference channel.

There are only 3 bottle samples, 2 from 50m and 1 from 30m .They come from areas where flushing of bottles is likely to be poor. Since all 3 were fired at the bottom of casts, they likely contain water from somewhat higher in the water column. So it is not a surprise that the CTD primary and secondary salinity were found to be higher than the bottle values by an average of 0.124psu and 0.126psu, respectively. Given that there are few samples, all shallow, all from the bottom and incomplete flushing likely to be a factor, this cannot be considered evidence of calibration drift. Based on standard deviation in the CTD salinity none of these comparisons would normally be selected in such a comparison.

The analysis was done within a month of collection.

For full details for the COMPARE run see file 2017-56-sal-comp1.xls.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

For the past 2 cruises when this sensor was used, the best setting for ALIGNCTD was found to be +2.5s, This setting worked well for this cruise as well. ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. The default setting of was selected and it does improve the data for both conductivity channels for these data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

There were no casts deep enough to make a comparison of channels useful.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book and no errors were found.

The cruise track was plotted and added to the end of this report. Since there were so many visits to some stations, only the plot with station names was produced using a list for one of each site. The usual version with cast numbers is too messy to be useful.

Header Check was run and shows that pressures were occasionally <0db.

- The lowest pressure in the CTD files was about -1.4db during cast #35. That was a single spike, so not significant, but the CTD spent about 12.5min running with negative pressures. During that time the values in most channels gradually change; perhaps they are normal values as the CTD warms up.  The pump turned off shortly after the CTD reached negative pressures. The conductivity values started to fall a few scans after pressure became negative, but once the pump stopped they stopped falling. So the pressure sensor looks like it was reading within ±0.2db.

- Plots were made to see if there are other cases of the CTD passing through the surface with pumps on. The only other case was event #4 during which conductivity began to drop about 2 scans after passing through 0db and fell rapidly within a few more scans.

The Surface Check gave the average surface pressure as 1.7db which is within the normal range for the Vector. The pressure sensor was checked at the factory in February 2017; recalibration does not look appropriate.

The altimeter and bottom depth headers were exported from the SAMAVG and CLN files to spreadsheets. The bottom depths were checked against the log book. The only discrepancies were a matter of 1m difference. The altimetry header entries were checked by calculating:

Check Value = (Bottom depth – maximum depth of sampling + altimetry header)

In 4 cases the check values was >3m so plots were made to examine those cases. The altimetry was often spiky near the bottom but estimates based on the profiles of altimetry confirmed the entries for casts #30 and 36 were reasonable, but that for cast #53 was not. The header entry for #53 was changed from 1.6 to 5m.  Cast #58 contained only surface data so that cast will not be processed further. Neither of the entries that were adjusted were bottle files.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. Many casts were checked to see if the alignment looked ok, and as usual, there is variability with the up and downcast traces sometimes closer than temperature and sometimes further apart. This may be due to varying vertical gradients and descent/ascent speeds. There are few casts with no stops for bottles which does limit the evidence. There is no evidence of poor alignment. The alignment setting used for this cruise has been used for many others during which there was better evidence. No further alignment will be applied.

Conductivity

Given the limitations of the data available, this step was run using the same settings used for cruise 2017-64 which immediately preceded this one.

Tests were run on 4 casts and the best results were with -0.6 records for the primary and -0.5 for the secondary conductivity. SHIFT was run twice on all casts using -0.6 records for the primary and -0.5 for the secondary conductivity.

Comparisons were made of T-S plots using data that had been put through DELETE with and without the shifts in conductivity. The alignment has improved the data, though there remain unstable features that may well be real and/or caused by problems other than alignment. As noted earlier the primary pump appears to have not worked well above 25m for many casts.

pH

A setting of +50 records produced the best results for 2017-64 which immediately preceded this cast.

SHIFT was run to advance the pH channel by 50 records for all casts.

A few casts were examined after this step and the results looked good..

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

At this point a further study was made of the problems in the primary plumbing. We can choose the secondary channels for temperature and salinity, but what to do with the dissolved oxygen is more problematic. There are two issues:

- Which casts are affected? -Because there are so many stops for bottles, finding differences between downcast and upcast profiles is not clear evidence of trouble. Finding differences between downcast primary and secondary salinity is easier. Almost all casts have questionable primary salinity right at the surface. It is harder to judge in high gradient zones, but the behaviour during stops for bottles suggests that the primary salinity is not performing as well as usual. Would removing the top 15m of data be useful? The errors would be smaller below that level, but would also remove what is probably the most useful part of the profiles.

- Are the upcast data reliable? – The largest differences between the primary and secondary salinity channels are in the top 10 to 25m of the downcasts. The problems may continue below that but just be less obvious because of lower gradients. Similarly, the upcasts look somewhat better, but that may be only because the local gradients are small to start with, so trouble doesn’t become obvious until close to the surface where local variability can cause larger differences. Upcast data are not desirable for the other variables, so this does not look like a useful way to produce reliable dissolved oxygen data.

To make judgments on which data are suitable for the archive the following tests were run:

- Profiles of casts at the same station were then compared and variability was very high near the surface. When T-S plots were examined it appears that there is little variability below 20m for most casts, but some casts looked ok below 10m.

- Examination of DO profiles showed odd profiles for those that were found to have high variability to 20m on T-S plots.

- The DO data are probably good enough to keep below 10m for casts 6-7, 30, 34-69 and 53-63.

For all other casts only DO data below 20m looks reliable.

##### Other Comparisons

Previous experience with these sensors –

All sensors were used during 5 earlier cruises in 2017 and 1 after this cruise that has been processed.

2017-63 in early April was a cruise in the Strait of Georgia and Juan de Fuca Strait. The bottles from Juan de Fuca were trusted more than those from the Strait of Georgia because conditions were better for flushing Niskin bottles; they indicated that the primary salinity was close to bottles while the secondary was high by about 0.0045. The dissolved oxygen comparison from 2017-63 was also based on Juan de Fuca bottles.

Cruise 2017-03 was in Baynes Sound and the Strait of Georgia, with shallow sampling and poor flushing of bottles, so the results were not considered reliable enough to use for recalibration.

Cruise 2017-05 was mostly off the west coast of Vancouver Island where flushing was good at depth. The comparison between the bottles and CTD salinity was quite tight; the primary salinity was found to be very close to bottles and the secondary was high by 0.0054psu. The dissolved oxygen comparison was also quite tight and led to a calibration of slope 1.0448 and offset +0.0527.
Cruise 2017-06 was mostly well offshore and showed the primary salinity to be low by ~0.0012 and the secondary high by ~0.0032. The dissolved oxygen was recalibrated using slope 1.0399 and offset 0.0645.

Cruise 2017-64 showed primary and secondary salinity low by 0.0012psu and high by 0.0034psu  in Juan de Fuca Strait. The dissolved oxygen comparison in Juan de Fuca was very close to that of 2017-06, so those results were used for 2017-64

Cruise 2017-23 comparisons were not trusted, so the 2017-06 results were used for recalibration.

Historic ranges – The local climatology is not representative of some of the sites occupied during this study and they are mostly close to shore where a 3-standard deviation climatology is not suitable. The climatology is somewhat out of date so may not reflect changing conditions, while on the other hand it is an area of high variability so the range of values in the climatology is large. All in all, it is not a very useful criterion to apply. However, finding systematic excursions from historic ranges might alert us to trouble. All temperature data fell within the climatology. Salinity was low at the bottom of casts at stations BS2, BS8, BS17 and BS23, but otherwise fell within the climatology. The data from the sites in more open waters to the south of Baynes Sound all fell within the climatology. There is not evidence of calibration problems.

Repeat Casts – There were many repeat casts but they are too shallow to be useful for judging repeatability. During 2017-06 when there were deep casts at a single site within 7 hours, differences along lines of constant σt were almost indistinguishable before bin-averaging and after binning they were <0.002º for the primary temperature and <0.0002psu for the primary salinity at σt= 27.52 (near 1450m). This shows excellent repeatability.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

Examination of primary and secondary T and S profiles and T/S plots show that while noise levels were similar for both channel pairs for most of the profiles, there were a number of casts which had poor primary salinity near the surface, suggesting near-surface plumbing problems. This is an area where biological fouling is quite likely.

The secondary T and S channels were chosen for editing.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes. All files required some editing, but mostly it was removal of records from near the top and bottom of casts and light cleaning of salinity.

Editing was also applied to the DO channel. All casts had DO data removed from the top (down to between 1.2 and 15.7db) except for # 30, 36, 47, 48, 57, 59 and 63.  Choosing the DO data to remove was a subjective matter based on how the primary salinity looked plus the DO profile itself.

Note: Further quality checks are reported in section 19 and more DO data were removed.

##### Initial Recalibration

- There is no evidence that pressure needs recalibration.

- Salinity:T1:C1 was found to read high by 0.0032, 0.0034 and 0.0026 during 3 cruises in June and July 2017. The first is the one considered most reliable with many deep bottles. Recalibration by subtracting 0.0032 looks appropriate.

- For dissolved oxygen the correction from cruise 2017-06 was selected:

CTD DO Corrected = CTD DO * 1.0399 + 0.065

File 2017-56-recal1.ccf was prepared to subtract 0.0032 from the secondary salinity and to recalibrate dissolved oxygen using.

CTD DO Corrected = CTD DO * 1.0399 + 0.065

CALIBRATE was run on the SAM and MRGCLN2 files using 2017-56-recal1.ccf.

COMPARE was run using the recalibrated and confirms that the recalibration was done correctly. The results do not look very good, but the average difference in the secondary salinity is reduced correctly.

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly.

See files 2017-56-sal-comp2.xlsx and 2017-56-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files.

##### Fluorescence Processing and special files for Angelica Peña

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some small unstable features, but those are from sites where active mixing is expected.

##### Further editing of Oxygen:Dissolved:SBE

At this point another study was made of the dissolved oxygen sensor data to ensure that clearly bad data had been removed. First, the bottle files were examined to see how well the CTD and bottle DO agree and they did quite well except in the top 5m a level at which we usually see some discrepancy. So the upcast CTD DO data look as they are reliable, at least when the CTD is stopped. The plots were examined in groups at stations BS2, BS17 and BS23 and there was reasonable consistency in each group.

Using that information, the CTD DO profiles (after editing and bin-averaging) were examined to see if the near-surface values look similar to the bottle values for those 3 groups of casts.

- At BS2 the results look good with no need for further editing of DO

- At BS17 event #31 was out of line down to about 20m, so more data need to be removed.

- At BS23 casts #42 to 44 look suspicious near the surface and maybe 45.

Based on these results the other casts at the same stations but with no DO sampling were examined to see if they seemed out of line. It was very hard to judge as there is considerable variability expected at the levels where the problems are seen.

Plots were then examined that have both salinity channels plus the DO channel. In most cases DO data had been removed at levels at which the two salinity channels have significant differences that look like they are due to plumbing problems. Exceptions found were the following casts: 4, 29, 31, 33, 42, 43, 44, 45, 53, 56 and 60.

A second run of CTDEDIT was made displaying both salinity channels and DO data; DO data were removed as appropriate and the steps in sections 16 to 19 were repeated.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. The differences increase as CTD DO values increase. The CTD DO was higher than the bottles by an average of ~0.24mL/L and standard deviation of 0.30mL/L.

When the same points are displayed in a plot versus pressure, it is seen that the CTD is highest relative to bottles near the surface and the differences are reasonably small below 25m. This makes sense if there is incomplete flushing and vertical DO gradients decrease with depth, which is roughly the case. The largest outliers come from areas where there are some very high gradients.

Most interesting is seeing that the 30m bottles stand out from the others, with the CTD reading low by an average of 0.05mL/L. All of the 30m bottles were fired at the bottom of casts, so flushing errors are expected to have the opposite effect to those fired during the upcast. All of the bottles fired at 40m were fired at least 10m above the bottom bottle, and those show the CTD reading high by an average of 0.10mL/L. This confirms that flushing accounts for a lot of scatter in these data.

A judgment on the calibration quality is limited by plumbing problems in the CTD data and having only shallow casts with high DO gradients combined with incomplete flushing of Niskin bottles, but there is no indication that further recalibration of DO will help.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Transmissivity, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

A note was added to cast #36 about the removal of the pH channel. (This was run on *TR files too.)

Data Processing Notes:

----------------------

Transmissivity, Fluorescence:URU:Seapoint, pH:SBE, PAR and PAR:Reference data are

nominal and unedited except that some records were removed in editing

temperature and salinity.

Oxygen:Dissolved:SBE data are calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small offset in

the fit is allowed. Recalibration was based on the results of cruise

2017-06 during which many samples were collected in areas where Niskin

bottles are believed to flush well.

There were sporadic problems with the primary plumbing that are severe near the

surface. The secondary temperature and salinity channels were selected

for archiving, but there is no alternative for the CTD dissolved oxygen

channel. Profiles of DO sometimes make it obvious where there are problems.

A further assessment was made by examining salinity profiles to identify

the depth at which the two salinity channels are close enough that the

dissolved oxygen sensor data are likely to be reliable.

CTDEDIT was used to remove DO data with quality likely to be poor, but

all the data should be considered of lower quality than usual.

No estimate of DO accuracy has been made due to the many uncertainties

in the comparison with bottles.

For details on the processing see document: 2017-64_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. For many casts near-surface DO data had been removed. Saturation values ranged from 120% to 135%. These are reasonable values for the season but given so little data very close to the surface this is not very useful information.

##### Final Bottle Files

CALIBRATE was run on all files and then the MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

The header check turned up a missing channel in some files so steps from REMOVE onward were repeated.

Plots were made of all casts to look for problems and no problems were found.

The standards check was run on all files. No errors were found.

The track plot looks ok.

Data from the CHE files were exported to a spreadsheet and compared with rosette sheets. No problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars

58. Only surface data in the file. Not processed.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 22Feb2017 | Factory |  |  |
| Conductivity | 1766 | 30Mar2017 | Factory |  |  |
| Secondary Temp. | 4484 | 22Feb2017 | Factory |  |  |
| Secondary Cond. | 3531 | 30Mar2017 | Factory |  |  |
| Transmissometer | 498DR | 21Jun2017 | IOS |  |  |
| Transmissometer | 1396DR | 2Feb2016 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 19Dec2015 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| Surface PAR | 20518 | 21Mar2016 |  |  |  |
| pH | 0691 | 15Jan2015 | Factory |  |  |
| SeaPoint Fluor. | 3685 | n/a |  |  |  |
| Pressure Sensor | 0550 | 24Feb2017 | Factory |  |  |
| Altimeter | 62355 |  |  |  |  |