## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 25 March 2025 | Updated channel names & formats in TOB files. G.G. |
| ### 26 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| ### 3 October 2020 | ### HPLC data added. SH. |
| 26 March 2018 | Corrected times &/or positions in CHE and CTD files for events 1 to 9. See 26end of report for details. |

## PROCESSING NOTES

Cruise: 2017-63

Agency: Ocean Sciences Division

Location: Strait of Georgia and Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Strait Water Properties Survey

Party Chief: Romaine S.

Platform: Vector

Date: April 8, 2017 – April 13, 2017

Processed by: Germaine Gatien

Date of Processing: 17 May 2017 – 20 July 2017

Number of original HEX files: 88 (includes 3 split casts)

Number of CTD files: 85	Number of bottle files: 29

Number of TSG files: 1

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#11185DR), a SBE 43 DO sensor (#1119), a SeaPoint Fluorometer (#3865) with a 3X cable, a Biospherical QSP-400 PAR sensor (#4565), Surface PAR (#20518) a pH sensor (#0691) and an altimeter.

The data acquisition program was Seasave v7.26.2.13.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

An SBE21 Thermoslinograph was used. There was no external temperature sensor or flow meter.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log book was in good order with notes of a few problems encountered. The sampling/rosette log sheets were generally in good order although two casts had the wrong event numbers and in a few cases duplicates were taken from different bottles than those indicated.

The biggest problem concerned time and positions: there were errors in both the NMEA data, computer time and time and positions recorded in the Daily Science Log book. Those issues were resolved by Event #10. Header times and positions were adjusted based on TSG and ship track records for events #4 to #9. The positions were correct for events #1 and #2; times were corrected but some uncertainty remains for the time of those events.

Salinity analysis was done promptly. Based on all casts the primary salinity was low by ~0.002psu and the secondary high by ~0.002psu. However, when only casts are included that were taken in Juan de Fuca Strait the primary CTD salinity was high by 0.0008psu and the secondary high by 0.0045. Since better flushing of Niskin bottles is expected in Juan de Fuca due to much more vertical motion during stops for bottles, that comparison is considered more reliable. The very calm conditions in the Strait of Georgia would lead to Niskin bottles containing water from lower in the water column with salinity higher than ambient waters, making the CTD values look lower than they really are. The primary sensors were selected for archiving and the primary salinity calibration was judged to be good.

The calibration of the dissolved oxygen sensor was also based on the results of the bottle comparison in Juan de Fuca Strait. As well as flushing problems in the Strait of Georgia, there are also very complex DO profiles with frequent reversals and great variation in vertical gradient. Even a cast-by-by recalibration would not work well there.

Channel Oxygen:Dissolved:SBE was recalibrated. The downcast Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.3 mL/L from 0 to 150db

±0.08 mL/L from 150db to 300db

±0.03 mL/L below 300db

As some of these errors may be due to incomplete flushing of Niskin bottles, this may underestimate the accuracy of the dissolved oxygen sensor during downcasts.

The fit of CTD fluorescence versus extracted chlorophyll showed the usual pattern of fluorescence reading lower relative to extracted CHL as CHL values rise. CTD Fluorescence was ~50% of the Extracted CHL values overall and slightly higher at ~60% in Juan de Fuca Strait where CHL values were lower.

Transmissivity data looked odd with sudden shifts in values that usually persisted through the remainder of a cast. Upcast and downcast data were often further apart than usual. In this region of great variability it is not completely clear that this is due to malfunction of the sensor. However, similar problems were seen during 2017-01 in offshore waters where the poor quality of the data was more obvious. As was done for 2017-01, the transmissivity channel will be removed from the final files, but the data will be available upon request, likely in the HISTORY folder.

Channel pH:SBE was removed from event #36 because there was a sudden shift to low values that differed from nearby casts and stood out in plots from the whole region. It is likely that something got deposited on the sensor. Transmissivity was extremely low for that cast as well, though the latter values are not trusted as noted above.

There is 1 TSG file. The Vector TSG set-up has no loop sampling, flow rate meter or intake temperature sensor. Comparisons with CTD data are the only means of assessing accuracy, but there was great variability in the results due to high near-surface vertical gradients. TSG salinity appears to be low by about 0.2psu but this may be due to small bubbles in the loop water and/or the TSG drawing water from above 2m. The TSG temperature is higher than CTD temperatures by about 0.2Cº, but this may only partly be due to heating in the loop and partly to water being drawn from above 2m.

For about 15 hours there was no flow in the loop so the temperature and salinity were removed from that section.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. Notes from the log book indicate a few problems that were investigated:

- Events 62/63 refer to the same cast; a new file started at the bottom so the upcast was given a different event # from the downcast.  This will not affect processing.

- For event 110/110a one Niskin bottle didn’t close so the CTD was returned to the water to collect surface samples. Unfortunately the pumps were not turned on so the CTD pumped channels will not be usable for those bottles; but CTD data from the first firing of the bottle are available.

- According to the log book there were problems with the NMEA time in events 1 to 7. One note suggests it was off by approximately 15 minutes. Not noted in the log is that there are varying discrepancies between the log times and computer clock times that persist until cast #10. The thermosalinograph file was converted to see if it would shed any light on which times are most reliable. The TSG NMEA appears to agree with the CTD NMEA time

To determine which times are best to use for casts 1 to 9, the following table was prepared:

| Event | Log | NMEA | Comp clock |
| --- | --- | --- | --- |
| 1 | 316 | 310 | 353 PST |
| 2 | 346 | 345 | 400 PST |
| 4 | 1440 | 1434 | 1510 PST |
| 5 | 1456 | 1520 | 1555 PST |
| 6 | 1601 | 1622 | 1701 PST |
| 7 | 1704 | 1724 | 1804 PST |
| 8 | 1741 | 1801 | 1840 PST |
| 9 | 1909 | 1908 | 1947 PST |
| 10 | 2207 | 2205 | 2207 PDT |
| 12 | 2306 | 2303 | 2306 PDT |
| 23 | 1249 | 1248 | 1249 PDT |

For the first 3 events it looks as though the Log time was taken from the NMEA time. When it was discovered that NMEA time was wrong, the log time was likely taken from the computer clock after 1 hour was subtracted due to a time zone error in the computer clock time. The NMEA time was supposedly fixed after cast #9, yet during cast #9 NMEA and log agree. This could be accidental as NMEA time is recorded when acquisition starts and there may have been a delay before the actual cast started. There are some erasures in the log that suggest something was going on that could have slowed down start of the cast. After cast #9 the log and NMEA are in good agreement and from cast #10 onwards the computer clock was in fixed and all 3 times are in good agreement.

For casts 4 to 7 the NMEA time is approximately 20 to 25 minutes fast compared to computer clock time adjusted to the right time zone. The computer time looks reliable once that adjustment is made. For casts 1 and 2 it is not clear what time is correct. If we assume that NMEA was 20min fast then they should be about 256 and 326. The first agrees pretty well with the computer clock, but event #2 does not. But event #2 had acquisition start at the bottom and the time recorded in the log appears to be when the CTD was at the bottom. There is so much room for confusion that picking the NMEA time minus 20minutes is the best choice. For cast #9 using NMEA minus 20 minutes matches the adjusted computer time quite well.

CONCLUSION: Use NMEA – 20 minutes. Not a perfect answer but close enough. And since the NMEA time is when the position was acquired it is the best time to use.

One mystery is whether the NMEA time to the TSG got corrected. There is no obvious sign of it, so when the TSG vs CTD comparison is done this may be a problem.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

There was no relevant sensor history for the pressure, conductivity and DO sensors because they have not been used since their factory recalibration in February or March 2017.

The only change in the configuration file concerned how the NMEA device was connected from “to the deck unit” to “to the PC”.

The calibration constants were checked for all instruments and no problems were found. One file was saved as 2017-63-ctd.xmlcon.

##### Conversion of Full Files from Raw Data

All hex files were converted using 2017-63-ctd.xmlcon to create CNV files.

A few casts were examined.  All expected channels are present. The primary and secondary temperature and conductivity channels are close during downcasts, but as usual the upcasts differ more due to noise in both channels. Dissolved oxygen, fluorescence, PAR, SPAR, pH and altimetry profiles all look normal.

Transmissivity looks odd with upcast and downcast further apart than usual for some casts, but not all.

There was a problem with this sensor on a previous cruise, so the data may not be reliable. This will be investigated further later.

The descent rate was mostly fairly high. It was quite noisy in Juan de Fuca Strait with some complete reversals of direction near the mouth. It was fairly steady in the Strait of Georgia.

##### BOTTLE FILE PREPARATION

ROS files were created using file 2017-63-ctd.xmlcon.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and no significant outliers were found.

There was a problem at the bottom of Event #62, forcing the start of a new file for the upcast. The upcast file name was called Event #63.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv.

Sample numbers were added to the file based on the rosette log records. There was no sample #105.

During cast #110 the surface Niskin bottle (#17) failed to close so the CTD was returned to the water to collect a particles sample in a cast identified as event #110a. The pumps were not on for #110a so pumped CTD channels are not useful. Sample #300 which was actually collected during event #110a, was treated as coming from cast #110, Niskin #17, in order to provide reasonable CTD data to accompany the sample.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine. Those files were bin averaged on bottle number.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2017-63-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF 2017-63CHL*.xlsx. The file included comments, flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2017-63chl.csv, event numbers were added, and the file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2017-63oxy*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2017-63oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in QF2017-63SAL*.xlsx. There were no duplicates. The analysis was done within 14-19 days of collection. The files were simplified and saved as 2017-63sal.csv.

That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet NUTS_QF2017-63nuts.xlsx. This includes a precision study. The spreadsheet was simplified, reordered on sample numbers and saved as 2017-63nuts.csv.

The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only, with output files named MRGCLN1.

There is no ROS file for cast #4. Though some samples are indicated on the sampling log and there are samples with labelled as #4 at station 102, but event #4 was at station 101 and there was no sampling. The Daily Log indicates sampling from event #5 at station 102. The samples were renamed as from #5.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number as the merge channel.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. There were a few other discrepancies between the Sampling Log and the samples in the analysts’ spreadsheets:

- Cast #8 - There was supposed to be a salinity sample 27, but the only salinity sample reported is labelled as #28. This will be investigated in COMPARE.

- Cast #34 - There was supposed to be a salinity sample 107. The salinity sample was labelled as #106. This will be investigated in COMPARE.

- Cast #57 – there is a sample that is not on the sampling log, but the values looks appropriate, so no action is needed.

- Cast #107 – There was supposed to be a sample #274 but the bottle was empty.

CLEAN was run on the MRG files to add 0 flags to empty flag channels and to update header limits.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There were many outliers with most looking like they were the result of poor flushing of Niskin bottles in the presence of large vertical salinity gradients. Significant outliers and the cases of questionable labels were investigated.

- Cast #8. There was supposed to be a sample 27 at 200m. The sample reported as #28 (175m) is likely mislabelled since it looks more like the CTD salinity at 200m and this is a cast in an area where we expect reasonably good flushing of Niskin bottles. So the sample number was changed but flag 3 was added with a comment of explanation.

- Cast #26 – Significant outlier, but there is a high vertical gradient and it is in an area where flushing is likely to be poor. There is no justification for flagging it.

- Cast #34 - There was supposed to be a sample 107 at ~200m. The sample reported as #106 (212m) is likely mislabelled since it looks more like the CTD salinity at 201db. While the sample looks more believable as #107 it is not completely clear from the comparison since salinity is noisy and the local gradient is not very high. Nonetheless, it is clear enough to doubt it was 106. The sample number was changed to 107 with flag 3 and a comment “Sample labelled #106 but sample # believed to be 107 based on sampling log and comparison with CTD salinity.”

- Cast #46 – the sample is an outlier but poor flushing can explain it as the local vertical salinity gradient is very high. No action needed.

When the sample numbers were corrected a further investigation was made of less significant outliers:

- Cast #85- A large shed wake passed through during the stop and the descent rate was very steady so flushing is likely to be incomplete; these could easily explain the bottle value.

- Cast #97- A large shed wake passed through during the stop and could easily explain the bottle value.

- Cast #101- There are large shed wakes during the stops at 100m and 125m and smaller ones at 75m and 150m. The standard deviations in the CTD salinity are also relatively high. The descent rate is very quiet so inefficient flushing is likely. These factors can explain the differences for this cast.

- Cast #107- There are large shed wakes during the stops at 100m, 150 and 175m, especially large at 150m. The descent rate was very steady so flushing is likely to be incomplete. These factors can explain the outliers.

- Cast #116- The CTD is higher than the bottle in this case, but this is in a zone with high variability and many gradient reversals and very steady descent rate. Incomplete flushing easily explains the difference and the standard deviation of the CTD salinity is high which may also explain part of the difference.

No further quality flags are appropriate as the bottle values probably do represent the contents of the bottles.

The analysis was done promptly after collection so we do not expect significant evaporation or adsorption.

When outliers were excluded based on differences >0.01psu, the primary salinity was found to be low by an average of 0.002psu and the secondary was high by 0.002psu, and the standard deviations were 0.003 for both.  It is expected that the CTD should read lower due to flushing problems in the Strait of Georgia. Data from Juan de Fuca Strait were examined since that area had the noisiest descent rate, so the Niskin bottles were likely better flushed. Using casts 5 to 18 only the primary salinity was found to be high by 0.0008 and the secondary high by 0.0045 with standard deviations 0.0014 and 0.0018 respectively.

This suggests that the primary sensors are providing the best salinity and does not require recalibration.

Secondary salinity will be recalibrated by subtracting 0.004psu to bring the two channels close in case there is a need to use the secondary channels.

For full details for the COMPARE run see file 2017-63-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There were 3 significant outliers in the plot of differences versus CTD DO:

Event 63 – The DO value from the CTD is very low – the pumps were turned off at the end of the downcast (event 62), a new file was started but pumps did not come on until the CTD had started to rise, so it took a while to get up to the value found at the end of the downcast. So the CTD data is not reliable and there is no reason to flag the bottle value.

Event 76 – In this cast the standard deviation in the CTD DO was extremely high; the vertical DO gradient was very high. Values did not appear to have equilibrated before the CTD started to move upwards again which may be due to high local variability. Poor flushing would explain at least part of the difference. There is no justification for flagging the DO value as it likely reflects the contents of the bottle, just not what the CTD was experiencing.

Event 110 – The standard deviation in the CTD DO is fairly high and the profile shows reversals in both CTD temperature and dissolved oxygen just as the bottle was closed –the Niskin bottle would not likely have captured that lower DO water. The sample value probably reflects what was in the bottle, but not the ambient value. No quality flag is justified.

When all casts were included and the 3 outliers above excluded as well as other outliers based on residuals, the fit was:

CTD DO Corrected = CTD DO * 1.0342 + 0.0712    R² = 0.80 (1)

Using only the 6 casts from the central Strait of Georgia, the fit was:

CTD DO Corrected = CTD DO * 1.0287 + 0.0896    R² = 0.68 (2)

Using only the 5 casts from Juan de Fuca Strait, the fit was:

CTD DO Corrected = CTD DO * 1.0478 + 0.0150    R² = 0.88 (3)

If some near-surface points are excluded first, then more outliers based on residuals the fit for Juan de Fuca looks like:

CTD DO Corrected = CTD DO * 1.0422 + 0.0358    R² = 0.91 (4)

These fits are always a compromise with the correction that suits high-gradient zones, not working as well in other areas. Most of the values above 10m were outliers, likely because the gradients were low so even if flushing is incomplete, the bottle values are likely closer to ambient values. This may indicate that the corrections should be even higher than what is indicated.

Fit (4) looks best overall for Juan de Fuca, suiting both the bottom and mid-depths.

These variations are consistent with other cruises during which there was a comparison between waters where flushing is likely to be good in one part and poor in another. The noisy descent rate in Juan de Fuca Strait indicates that flushing was likely to be good there whereas it is very quiet in the Strait of Georgia. It is likely the most appropriate correction to apply to all the data, though it may still be an underestimate.

For more detail see document 2017-63-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further problems were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor. There were 84 samples with a range of ~0.4ug/L to ~32.4ug/L. The fit of Fluorescence against extracted CHL is:

SBE Fluorescence = 0.51 * CHL

This is fairly typical for these fluorometers when there are high CHL values.

In Juan de Fuca Strait where CHL is lower, the fluorescence is about 60% of CHL, on average.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were run on ALIGNCTD to determine the best setting to advance the DO signal. The best results were with 1.5s.

ALIGNCTD was run on all casts using +1.5s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. The default setting of was selected and it does improve the data for both conductivity channels for these data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 4 of the deeper casts to calculate differences between sensor pairs. None of the casts are very deep.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2017-63-0074 | 400 | -0.0001 | +0.00037 | +0.0040 | F.High, Steady |
| 2017-63-0083 | 400 | 0 | +0.00038 | +0.0039 | F.High, Steady |
| 2017-63-0107 | 320 | -0.0006 N | +0.00040 N | +0.0045 N | F.High, Steady |
| 2017-63-0129 | 300 | -0.00045 N | +0.00048 N | +0.0050 N | F.High, F.Steady |

The differences in conductivity and temperature are not large, though they have cumulative effects on salinity so those differences are larger than we would expect from newly calibrated sensors. Any suggestion of temporal drift may be due more to the depth and the noisiness of the temperature and conductivity during casts #107 and #129. The salinity differences for casts #74 and #83 are consistent with the differences inferred from the comparison of CTD and bottles.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book. As noted in section 3 there are problems with times in both the log book and the file headers. It was determined that the most reliable time is to subtract 20 minutes from the NMEA (header) times for Events #1-9. A note will be added to the headers to indicate that time is estimated for those casts.

There were a few other problems noted:

- Event #97 – the NMEA entries were all on a single line so the positions did not get converted properly.

- Event #100 – had wrong station name – should be station 8.

These corrections, including the time corrections were made to the CNV files after DERIVE. Then the Conversion to IOS headers and CLEAN were rerun.

Events #2 and #63 contain only upcast data, so will not be processed further.

The cross-reference list was rerun after that step and all corrections had been done successfully.

The cruise track was plotted and added to the end of this report.

The Surface Check gave the average surface pressure as 1.6db which is within the normal range for the Vector.

Header Check was run and shows that pressures were occasionally <0db.

- The lowest pressure in the CTD files was about -0.6db during cast #5 and pumps were on then. Plots show that the CTD was moving up and down between positive and negative pressures for about 15 seconds.

- Another case of negative pressure was during cast #126 and that lasted for only 3 seconds.

- During cast #17 the CTD had negative pressures for about 1 second at the beginning of the cast and as soon as the pressure went below 0, salinity spiked to low values. It did not get below 26psu but there was likely not enough time for that to happen. It looks like the CTD was out of water when pressure was <0db.

While the pressure could be reading slightly low, it is more likely that the CTD was moving in and out of water occasionally. The sensor was recalibrated shortly before the cruise. Recalibration does not look appropriate.

The altimeter and bottom depth headers were exported from the SAMAVG and CLN files to spreadsheets. The bottom depths were checked against the log book. There were many minor discrepancies but a few were large enough to need investigation. The altimetry header entries were checked by calculating

Check Value = (Bottom depth – maximum depth of sampling + altimetry header)

The following cases were examined because either the Check Value >4m or bottom depths differed significantly from log entries:

- Casts 4, 76 & 89 had depths in the log that produced better Check Values than the header entries so the latter were changed.

- Casts 3, 74 & 108 had check values > 4m and headers that matched the log; plots showed that the altimetry headers contain reasonable estimates. There was likely some drift during the casts; no changes were made.

- Cast #134 had no altimetry header as it did not get within 15m of the bottom, yet the maximum depth is just 1m less than the header entry for the bottom depth. The log entry had been erased and 231m entered, but in 2016 at this site a value of 213m was found. That value makes much more sense, so it was entered into the header.

- Cast #90 had no altimeter header. While it looks like a value of 8m is reasonable, it is noisy and was not entered.

- Casts 118 and 123 also had no altimeter headers. The altimetry has large spikes but a value of 8m looks good for both and produces good check values. So an altimetry header was entered for those casts.

- Cast #119 was similar to the 2 previous cases except that entering a header value did not improve the check value. But it didn’t make it worse either, so this is assumed to be a case where either the log depth is slightly off or some drift occurred between recording depth in the log and the CTD nearing the bottom. An altimetry header was added.

The only bottle file affected was #76 (water depth changed). No altimetry header is appropriate for #123 because there was no deep sampling.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and as usual, there is variability with the up and downcast traces sometimes closer than temperature and sometimes further apart. This is likely due to varying vertical gradients and descent/ascent speeds. Overall the choice made earlier looks appropriate, so no further alignment will be applied.

Conductivity

Tests were run on 4 casts and the best results were with -0.6 records for the primary and -0.5 for the secondary conductivity.

SHIFT was run twice on all casts using -0.6 records for the primary and -0.5 for the secondary conductivity.

pH

Tests were run on 3 casts to determine the best shift to make the difference between upcast and downcast pH profiles similar to that for temperature. A setting of +50 records produced the best results.

SHIFT was run to advance the pH channel by 50 records for all casts.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

There is no history for the pressure, temperature, conductivity and dissolved oxygen sensors since they have not been used on any other cruises since they were last serviced and recalibrated. The transmissometer malfunctioned during cruise 2017-01.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed.

- The casts to the south at the mouth of Juan de Fuca Strait had temperatures a little above the climatological maximum in areas where salinity was low. All other temperature data were within the climatology.

- Salinity was mostly within the climatology except for low values above 50m and below 100m in many casts in Juan de Fuca Strait. There were just a few other cases of slightly low salinity elsewhere.

Repeat Casts – The only repeat casts were in areas of active mixing, so not useful for judging the repeatability of the sensors.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

There primary T and S channels were chosen for editing based on the bottle comparison; there appears to be a little less noise in the profiles for the primary, as well, though the difference is not great

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes. All but 1 file (#117) required some editing, but mostly it was removal of records from near the top and bottom of casts and light cleaning of salinity.  Casts #18 to 25 and #134 had noisy descent rates with some complete reversals of direction so editing was heavier for those.

##### Initial Recalibration

- There is no evidence that pressure needs recalibration.

- Salinity:T1:C1 is reading higher than the primary by ~0.004psu, so subtracting -0.004 will bring it into agreement with the primary which is considered more realistic

- Dissolved oxygen will be recalibrated based on the Juan de Fuca comparison:

CTD DO Corrected = CTD DO * 1.0422 + 0.0358

File 2017-63-recal1.ccf was prepared to apply the above corrections to salinity and dissolved oxygen.

COMPARE was run using the recalibrated salinity data to check that the correction was applied correctly; the average differences for primary and secondary salinity are within 0.0001 of each other.

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly.

See files 2017-63-sal-comp2.xlsx and 2017-63-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When bottles from above 60m were removed the CTD DO was higher than the bottles by an average of ~0.05mL/L and standard deviation of 0.09mL/L.  When examined according to event number it was found that the differences were smaller in the Strait of Georgia despite the fact that flushing problems are likely more serious there. This is explained by looking at the individual casts. The casts from the western part of Juan de Fuca Strait have large and fairly steady vertical DO gradients. The casts from the central part of SoG have reversals in DO at mid-depths so that both flushing and response errors would lead to some cases of CTD reading lower than bottles. Where the DO vertical gradients are high with DO decreasing with depth, the CTD is reading higher than bottles. The last two casts with DO sampling were well-mixed casts so that errors due to poor flushing and slow DO sensor response are both negligible.

As is usual in this region, the profiles are very complex. A judgment on the calibration quality is not clear, but there is no indication that further recalibration of DO is needed.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

The CTD files from rosette casts were clipped to 50m; sigma-T was derived and the data were exported to a single file, 2017-63-SOG.csv.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some small unstable features, but those are from sites where active mixing is expected.

Transmissivity looks unreliable. This is not obvious when only downcast profiles are examined, but looking at full casts there are some sudden shifts in values and there are significant differences between downcast and upcast profiles for many casts. This area has complex profiles so this might not be a damning observation had the same thing not been noticed during 2017-01 in February. At that time there was some suspicion that the problem might have been cable-related, but these data are from another ship so this does not seem likely. In any case the data do not look sufficiently reliable to go into the OSD_DATA_Archive. Two sets of files will be created, one with transmissivity and one without, so that users who really want to see it can find it.

The pH data from cast #36 looks odd starting with values that look like nearby casts but suddenly shifting to lower values and staying there through both downcast and upcast. This does not correspond to a shift in transmissivity although transmissivity is low throughout the cast. The shift only occurs during the one cast, so the causes are assumed to be different. However, the very low transmissivity may indicate that there was something that the CTD went through waters with biological material that may have draped over one or both sensors. The pH channel should be removed from cast #36 only.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Transmissivity, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

For cast #36 the pH:SBE channel was also removed.

REMOVE was run a second time without removing transmissivity so that the data will be available for those who request it. The letters TR were entered after the regular extensions.

A second SBE DO channel (with umol/kg units) was added. (This was run separately on *TR casts)

REORDER was run to get the two DO channels together. (This was run separately on *TR files.)

HEADER EDIT was used to fix formats and channel names and to add the following comments:

A note was added to cast #36 about the removal of the pH channel. (This was run on *TR files too.)

Data Processing Notes:

----------------------

Fluorescence:URU:Seapoint, pH:SBE, PAR and PAR:Reference data are nominal and unedited

except that some records were removed in editing temperature and salinity.

The Transmissivity channel was removed because of frequent shifts in values during

casts and large variations between values during repeat casts.

The data are available, by request, from the chief scientist.

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.3 mL/L from 0 to 150db

±0.08 mL/L from 150db to 300db

±0.03 mL/L below 300db

The Transmissivity channel was removed because of frequent shifts in values during

casts and large variations between values during repeat casts.

The data are available, by request, from the chief scientist.

For details on the processing see document: 2017-63_Processing_Report.doc.

Warning: For casts #1 to 9 there were problems with NMEA times. A correction was made

by subtracting 20 minutes from those times, but this should be considered an estimate.

For details see section 3 of document 2017-63_Processing_Report.doc.

For details on the processing see document: 2017-63_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values ranged from 85% to 130%. All casts near the mouth of Juan de Fuca Strait were at ~103% which is what we expect in open waters. As usual values were lowest in the southern Gulf Islands region and at a few casts in narrow channels where vertical mixing is strong. The highest values were in the central part of the Strait of Georgia. There is no evidence of a DO sensor calibration problem.

##### Final Bottle Files

CALIBRATE was run on all files and then the MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, pH:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the pairs of DO channels together.

When processing was almost finished a change was made to how CHL flagging is done. The flags and comments were affected as follows: Event/sample number 43/130, 63/164, 94/236&237, 110/296 & 297, 116/309. The CHL and MRGREO files were updated and the HEAD EDIT step rerun with updated comments.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

The steps from REMOVE to HEADEDIT were repeated for casts that contain transmissivity data and have extensions with *TR. (The corrections to the CHL were made in the MRGREOTR files before Head Edit was rerun.)

Plots were made of all casts to look for problems and none were found.

Standards check and a header check were run on all files. No errors were found.

The track plot looks ok.

Data from the CHE files were exported to a spreadsheet and compared with rosette sheets. No problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

##### Thermosalinograph Data

There were no loop samples, flow meter or intake thermistor. The intake is at about 2m. The only method to check calibration is to compare with the CTD casts. There are many more CTD data available at 3m than 2m, so data from 3m will be used as well as 2m.

a.) Checking calibrations

The configuration file, 2017-63-tsg.xmlcon, had no errors in the calibration parameters.

b.) Conversion of Files

The hex files were converted using file 2017-63-tsg.xmlcon.

They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots were produced. The traces look ok to start, but then temperature starts to rise steadily at about 22:30 on April 9th, then spikes to about 20ºC at 13:30 on April 10th. then suddenly drops back to the sort of values seen earlier in the cruise. The spike is at about the time of a net cast at station 63. It looks as though the pump might have been off for about 14 hours. This may be clearer when the comparison with the CTD is done.

The track plot looks fine.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.3db of 3db, and then again using 2db. These were exported to 2 spreadsheets which was saved as 2017-63-ctd*-tsg-comp.xls. There were 72 and 80 casts, at 2db and 3db respectively, which overlapped with TSG files.

The TSG file was opened in EXCEL, median and standard deviations (over 5 records) were calculated for lab temperature and salinity and the files were reduced to the times of CTD files. Those data were added to as 2017-63-ctd3-tsg-comp.xls.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. As expected the results are not as good as we would expect early in the cruise, with differences particularly high for casts 4, 7 and 8, and all a little high for casts 4 to 9. After cast #10 all looks well. The CTD times are estimates and whether the position comparison is good or poor depends on whether the ship was stopped or not during the time difference. It was not obvious if the times were also wrong in the TSG file, but there is no evidence of that. There is no sudden jump in time. A comparison was made between the CTD data for casts 4-9 using TSG data from 20 minutes later and the position differences are sometimes better, sometimes worse, and sometimes longitude is better but latitude worse, or vice versa. So it is likely that the TSG time is correct and the CTD times are likely slightly off. So those casts will not be included in further comparisons.

There were no position differences >0.0016º and both the average and median differences were 0.0001º and 0.0000º.

d.) Comparison of T and S from TSG and CTD data

In the table below T3 and T2 refer to CTD data from 3db and 2db, respectively.

When the casts with estimated time were excluded there were still some large differences between the CTD and the TSG. A plot of temperature differences versus cast number offers an explanation.

Casts #10 and 12 look ok, but for casts #13 to 23 the differences are extremely large. This result agrees with the observations of odd temperature rises in the time plots. It looks as though the flow stopped or was severely reduced. Probably the pump was accidently turned off and then restarted before or during cast #25. Examination of the TSG file shows the time when the pump must have restarted as 13:34 on April 10th, but normal flow took a little while to be re-established.  It is harder to determine when the problem started, but it must have been between 22:00 on April 9th and 0:40 on April 10th. Even during cast #12 there are hints that the temperature was beginning to rise. A plot of salinity suggests a lack of small variability started shortly before 2300. So temperature and salinity should be removed from the TSG between about 22:45 on the 9th and 13:35 on the 10th. Casts during that time will be removed from the comparisons (casts #12 to 25).

Tmed and SALmed refer to median TSG temperature and salinity, with the median taken over 5 records.

|  | Tmed-Tctd3 | Smed-Sctd3 | Tmed-Tctd2 | Smed-Sctd2 |
| --- | --- | --- | --- | --- |
| Average using all available casts* | +0.2542 | -0.4013 | 0.2207 | -0.2989 |
| Median using all available casts | 0.2238 | -0.1120 | 0.2055 | -0.0718 |
| Std Dev using all available casts | 0.1751 | 0.6801 | 0.0775 | 0.5245 |
| Average of casts with outliers excluded | 0.2275 | -0.0877 | 0.2146 | -0.0748 |
| Median of casts with outliers excluded | 0.2182 | -0.0485 | 0.2046 | -0.0437 |
| Standard dev with outliers excluded | 0.0389 | 0.0799 | 0.0661 | 0.0756 |
| # of casts after outliers excluded | 48 | 41 | 46 | 39 |

*Casts with TSG pumps off or unreliable CTD positions were not included.

The outliers come mostly from the Central Strait of Georgia where near-surface vertical gradients are relatively high. We expect the 2m comparison to be better, but we have fewer reliable CTD measurements at that level.

Usually we get better estimates by looking at the casts that occurred when the standard deviation in the TSG data was low, but for this cruise those casts were frequently ones with high vertical gradients so that the standard deviations in the differences were very high.

The TSG temperature is found to be higher than the CTD temperature by between ~0.22C° and ~0.25C° at 3m and~ 0.20C° to ~0.22C° at 2m depending on which data were selected. The median and average get closer after removing outliers. The best results are likely from the comparison of 2m data excluding outliers, with temperature high by a median value of 0.205C° with a standard deviation of 0.066C°.

The TSG salinity is consistently lower than the TSG salinity but there is an extremely large scatter in the results. When all available casts were used the average difference was 4 times the median difference at both depths. When outliers were removed the TSG salinity was low by a median value of 0.05psu using the 3m data and 0.044 at 2m. Even in the reduced sets the standard deviations are about 0.08psu.

The lower TSG salinity may be due small bubbles in the loop water and/or be due to water being drawn into the loop from above 2m. Calibration drift may explain some of the difference but that error is expected to be <<0.01psu.

There is too little evidence to support recalibrating the TSG salinity. It may well reflect actual conditions closer to the surface than that measured by CTD data. This would also explain some of the temperature difference, so no attempt will be made to calculate a proxy for intake temperature, though assuming it is slightly high is a valid conclusion given that the water must warm up a little in the loop.

(See 2017-63-ctd-tsg-comp.xls.)

Calibration History

The TSG temperature and conductivity were recalibrated in December 2015 and have been used for 3 other cruises: 2016-07, 2016-10 and 2016-71 in the same region. For the cruise in June the lab temperature was higher than the CTD temperature by between 0.1C° and 0.2C°.  Salinity was found to be low by 0.62 using all casts, but the standard deviation was very high at 3.0psu. The median difference was 0.03psu when only the 10 casts with the lowest standard deviation in the TSG salinity were included. For 2016-10 in September the CTD temperature was higher than the CTD temperature by a median of 0.127C°. The TSG salinity was lower than the CTD salinity by a median value of 0.93psu using only the 20 casts with the lowest standard deviation in the TSG salinity. The differences were slightly lower when the 2db data were used but there were fewer data available. During 2016-71 in November the TSG lab temperature was higher than the CTD temperature by about 0.13Cº. The TSG Salinity was lower than that from the 3m CTD by about 0.05psu. Calibration drift is unlikely to be that large. It is more likely that the difference is due to small bubbles in the loop and or that water is being drawn from above 2m. Bubbles are not obvious in the TSG traces, but small ones could be present.

Conclusions

1. The TSG clock worked well after cast #10, and was probably fine before then but we cannot confirm this from the CTD data.

2. The TSG lab temperature was higher than the CTD temperature by about 0.2Cº. We have little experience with the Vector loop system and no knowledge of the flow rate during the cruise. From the Tully we might expect heating of about this magnitude given similar intake temperatures, but the Vector loop is shorter so that would reduce the amount of heating. The intake temperatures were lower than the last time the TSG was used on the Vector, so it is not surprising that the heating in the loop is higher for this cruise. It is possible that the flow rate was lower which would also lead to more heating.

3. The TSG Salinity is lower than that from the 2m CTD data by about 0.04psu. Calibration drift is unlikely to be that large. It is more likely that the difference is due to small bubbles in the loop water and/or due to water being drawn into the loop from above 2m. Bubbles are not obvious in the TSG traces, but small ones could be present. The near-surface vertical gradients are very high for some casts.

4. Overall, the quality of the data was good except for the section during which the pump must have been turned off accidentally.

5. There is insufficient evidence to justify making an estimate of intake temperature or to recalibrate salinity.

f.) Editing

CTDEDIT was used to remove temperature and salinity data from scans 2018 to 3804 because it appears that the flow in the loop stopped or was severely reduced. The TSG pump may have been turned off accidentally.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Conductivity:Primary, Flag and Position:New channels.

HEADER EDIT was used to change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats and to add the following comment:

TSG Data Processing

-------------------

There was no loop sampling, intake temperature sensor or flow rate meter.

A comparison of temperature and salinity was made with co-incident CTD casts

but high near-surface vertical gradients resulted in a lot of variability.

The TSG salinity is lower than that from the CTD salinity with a median difference

of about 0.04psu compared to CTD data from 2m. This may be due to bubbles

in the loop water or the loop may be drawing water from above 2m. There could

also be some error due to calibration drift but that is likely to be negligble.

Similarly, there was considerable variability in the differences between the

TSG temperature and that from the CTD. The TSG temperature appears to be high

by about 0.2 C degrees. However, this may be partly due to the loop drawing

water from higher in the water column.

Temperature and Salinity data were removed from a 15-hour section of the record

because temperature climbed steadily and salinity was near-constant, indicating

that flow in the loop had stopped.

The Temperature:Primary was renamed as Temperature:Lab to make it clear that it is not the intake temperature.

The file was saved as 2017-63-0001.TOB.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series and all look fine.

The cruise plot was added to the end of this report.

##### Particulars (mostly notes from log)

1. test cast – no bottles fired

2. Upcast only logged.

4. High current

5. NMEA time was off again. Computer time is later – need to adjust.

6. NMEA time is fast.

8. Starting with this cast NMEA feed was set to “from deck unit” to fix time.

10. First cast with computer time/NMEA time and log time very close.

31. Alarm sounded on way down. Scott re-terminated – just electrical not mechanical portion. Cast rerun as event 34. No bottle file but downcast data may be useful.

62. Downcast only – no bottles

63. Upcast and bottles only.

108. Muddy bottom hit.

110. Bottle 17 didn’t trip. File 110a just for recollecting water at surface.

##### Corrections to times and positions in events 1 to 9 CTD and CHE casts.

Problems were mentioned in the log book about times of some early casts. However, it was not noticed that positions were also wrong until cast #10.  A track plot shows station names misplaced from normal.

Positions were obtained from the GPS records of the vessel track. They are in line with the normal positions for those stations. The thermosalinograph record shows no problems with NMEA download. At the time when the feed to the CTD system was corrected there is no jump in time in the TSG record, so it is assumed that it was not affected. So those records were also trusted in establishing time and position.

Using the GPS track of the vessel and the TSG record when there was little ship movement, times and positions were found for events #4 through #10. For Event #10 the times from the log, TSG and NMEA download are within 3 minutes each other. In most cases the time found this way was close to the “SYSTEM UTC” time entered in the SeaBird headers, but there are large differences for events #6, 8 and 9. This likely reflects the efforts being made by the science crew to fix the time problems, resulting in what seems to be a random pattern. See document “2017-63-Times-positions-study.xlsx”.

For events #1 and #2 at station SI it is assumed that the positions recorded are correct since they are very close to the usual site for that station. Event #3, a drifter deployment, is also very close to the usual position for station 58 and that position is found right at the beginning of the TSG records, just a few minutes after the log time entry for the drifter deployment. That makes sense, so it appears that the position problem arose after event #3. The time of events #1 and #2 are impossible to establish beyond doubt since the TSG was not operating early in the cruise, but since the System UTC entry time in the SeaBird headers are generally in line with the results of the TSG data study, those values will be used to correct the first 2 events.

The problems with the positions are very odd with what looks like each cast having an initial position close to the ones expected for the previous cast. There is also a large drift during the casts as the position seems to catch up to where we expected it to be. For example the cast at station 75 shows a drift of 7km between start and end positions. The depths of the casts confirm that they really did occur at the sites given in the Station entry. So the NMEA positions are clearly bad. The data appear to have been cached. When the computer was started it has info from the last cast and catches up while on station. The order of stations matches the TSG map, but the positions are wrong.

Conclusions:

For Events 1, 2 and 3 the positions are correct.  Time is correct for event #3. Time for events 1 and 2 are impossible to determine but since the times for events 3, 4 and 5 appear to be close to the System UTC time, that time will be used for events 1 and 2.

The positions for events 4 through 9 will be those from the ship track record.

For events 4, 5, 6, 7, 8 and 9, the time based on TSG records at the corrected positions will be used.

A note will was added to each header to indicate that changes were made and why.

A new cross-reference list was prepared.

Cross-reference list before corrections – events 1 to 12

2017-63-0001.ctd   0001  SI          48 39.27 N  123 27.57 W  UTC 2017/04/09 02:49

2017-63-0004.ctd   0004  101      48 25.43 N  124 41.11 W  UTC 2017/04/09 14:14

2017-63-0005.ctd   0005  102      48 25.32 N  124 45.19 W  UTC 2017/04/09 15:00

2017-63-0006.ctd   0006  103      48 29.97 N  124 44.05 W  UTC 2017/04/09 16:03

2017-63-0007.ctd   0007  76        48 32.58 N  124 39.15 W  UTC 2017/04/09 17:20

2017-63-0008.ctd   0008  75        48 31.38 N  124 29.97 W  UTC 2017/04/09 17:41

2017-63-0009.ctd   0009  74        48 28.05 N  124 32.94 W  UTC 2017/04/09 18:48

Cross-reference list after corrections – events 1 to 9

2017-63-0001.ctd   0001  SI          48 39.27 N  123 27.57 W  UTC 2017/04/09 03:53

2017-63-0004.ctd   0004  101      48 25.36 N  124 45.19 W  UTC 2017/04/09 15:13

2017-63-0005.ctd   0005  102      48 29.96 N  124 44.05 W  UTC 2017/04/09 15:48

2017-63-0006.ctd   0006  103      48 32.98 N  124 43.36 W  UTC 2017/04/09 16:49

2017-63-0007.ctd   0007  76        48 31.32 N  124 29.77 W  UTC 2017/04/09 18:08

2017-63-0008.ctd   0008  75        48 28.21 N  124 32.75 W  UTC 2017/04/09 18:34

2017-63-0009.ctd   0009  74        48 25.21 N  124 35.95 W  UTC 2017/04/09 19:31

For more information see spreadsheet 2017-63-Times-positions-study.xlsx.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 22Feb2017 | Factory |  |  |
| Conductivity | 1766 | 30Mar2017 | Factory |  |  |
| Secondary Temp. | 4484 | 22Feb2017 | Factory |  |  |
| Secondary Cond. | 3531 | 30Mar2017 | Factory |  |  |
| Transmissometer | 1185DR | 27Apr2016 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 19Dec2015 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| Surface PAR | 20518 | 21Mar2016 |  |  |  |
| pH | 0691 | 15Jan2015 | Factory |  |  |
| SeaPoint Fluor. | 3865 | n/a |  |  |  |
| Pressure Sensor | 0550 | 24Feb2017 | Factory |  |  |
| Altimeter | ? |  |  |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3363       Cruise ID#:	2017-63

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3363 | 17Dec15 | Factory |  |  |
| Conductivity | 3363 | 17Dec15 | Factory |  |  |