## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 20 Mar 2025 | ### Updated channel names & formats in TOB files. GG |
| ### 6 December 2021 | ### Corrected Salinity:Bottle precision lost during the addition of HPLC. S.H. |
| 3 October 2020 | HPLC data added. SH. |

## PROCESSING NOTES

Cruise: 2017-64

Agency: Ocean Sciences Division

Location: Strait of Georgia and Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Strait Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: June 21, 2017 – June 27, 2017

Processed by: Germaine Gatien

Date of Processing: 28 August 2017 – 23 October 2017

Number of original HEX files: 89

Number of CTD files: 89	Number of bottle files: 27

Number of TSG files: 2

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0550) was used for this cruise. It was mounted in a rosette and attached were a SBE 43 DO sensor (#1119), a SeaPoint Fluorometer (#3865) with a 3X cable, a Biospherical QSP-400 PAR sensor (#4565), Surface PAR (#20518) a pH sensor (#0691), an altimeter and a Wetlabs CSTAR transmissometer (#4985DR for casts 1-100 and #1396DR for casts 101-110).

The data acquisition program was Seasave v7.26.2.13.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

An SBE21 Thermoslinograph was used. There was no external temperature sensor or flow meter.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log book was missing an equipment list and there was no mention of the change of transmissometer. In one case there was no indication that bottles were sampled for a cast.  The sampling/rosette log sheets had errors in sample numbers that led to duplicate numbers and great confusion in labels, as some matched the log sheet and some had the numbers that should have been entered. It took considerable work to ensure that consistent sample numbers were used for all samples from a single Niskin firing. There were errors in some event numbers on rosette sheets.

Salinity analysis was done promptly. The comparison of bottle salinity with CTD salinity indicated that the primary salinity was low by ~0.0018psu and the secondary high by ~0.0012psu. However, many of those casts were likely affected by poor flushing of Niskin bottles which leads to the CTD appearing to read lower than it actually does. Conditions in the western part of Juan de Fuca Strait are usually rough enough to ensure better flushing of bottles, but few bottles are collected there. For this cruise only 1 of 4 bottles from that area could be usefully compared with CTD salinity. The comparison for that bottle showed the primary CTD salinity to be low by 0.0012psu and the secondary high by 0.0034psu; that is very close to the findings from 2017-06 which sampled offshore waters just before this cruise, had extensive bottle sampling in deep waters and used the same sensors. We do not usually have the luxury of such excellent sampling to guide calibration decisions for inshore cruises.

For salinity-calibration sampling during Strait of Georgia/Juan de Fuca Strait Water Properties cruises the following sampling schemes are recommended:

- Take more salinity samples in the western part of Juan de Fuca Strait (stations 102, 75, 72 and 69) where the sea state is usually high enough to help flush the Niskin bottles.

- Multiple bottles per cast would be especially useful from stations 102 and 75 where the sea state is usually very high.

- When taking only 1 bottle per cast, make sure it is from at least 10m above the bottom. The usual scheme used for this project is to sample Niskin #2 which is generally a good choice.

The calibration of the dissolved oxygen sensor was also based on the results of the bottle comparison in Juan de Fuca Strait. As well as flushing problems in the Strait of Georgia, there are often complex DO profiles there, with frequent reversals and great variation in vertical gradients.

Channel Oxygen:Dissolved:SBE was recalibrated. The downcast Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.2 mL/L from 0 to 150db

±0.08 mL/L from 150db to 300db

±0.01 mL/L below 300db

The fit of CTD fluorescence versus extracted chlorophyll showed the usual pattern of fluorescence reading higher than CHL for low CHL and then dropping relative to extracted CHL as CHL values rise. CTD Fluorescence was ~77% of the Extracted CHL values overall. CHL ranged from 0 to 11.3ug/L with only 5 samples with CHL>4ug/L.

Two transmissometers were used during this cruise and comparisons were possible since some stations were occupied before and after the switch. Sensor #498DR had been serviced recently. Sensor #1396 was used in late 2016 without a problem, but in recent cruises it shows an odd profile as though values get stuck for a half-second or so, then a few new values are recorded with the usual sampling rate then values get stuck again. There were systematic differences in values with sensor #1396 having a smaller range of values, reading lower near the surface and often higher at greater depth; however, the two sensor readings were close at the bottom. Since temperatures also changed between repeat occupations and there is evidence of active mixing in the area, the differences could be due to changes in water properties.

Channel pH:SBE was removed from event #19 because very low values suggest that the buffer bottle was not removed.

There are 2 TSG files. The Vector TSG set-up has no loop sampling, flow rate meter or intake temperature sensor. Comparisons with CTD data are the only means of assessing accuracy, but there was great variability in the results which may be due to large vertical and horizontal gradients. TSG salinity appears to be low by about 0.5psu which may be due to the TSG drawing water from above 2m and/or small bubbles in the loop water. The TSG temperature is higher than CTD temperatures by about 0.13Cº; this may be partly due to heating in the loop and partly to water being drawn from above 2m.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. Notes from the log book indicate a few problems that were investigated.

- The station name was wrong for event 9. This was corrected.

- Station 101 was skipped due to rough conditions.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The sensor histories for the pressure, conductivity and DO sensors were found.

There were 2 changes in the configuration file; the first concerned how the NMEA device was connected from “to the deck unit” to “to the PC” and the second to a change in transmissometer.

The calibration constants were checked for all instruments and no problems were found. Based on the files used at sea, 3 versions were saved as:

- 2017-64-ctd1.xmlcon – Casts 1-65 and 67

- 2017-64-ctd2.xmlcon – Casts 66 and 68-100

- 2017-64-ctd3.xmlcon – Casts 101-110

Note that in the log it was said of cast 66 “NMEA feed to Seasave not showing input data – time and position were taken from the bridge feed for the log”. The cast could not be converted.with the con file used for casts 1 to 63 and 68. It could be converted using the con file with the connection to the PC, 2017-64-ctd2.xmlcon. After conversion the position looked fine, but the * NMEA UTC (Time) needed adjusting to match the log entry and System Upload Time.

NOTE: After the data had been partially processed, further information became available from another cruise and showed that the hysteresis factor E in the dissolved oxygen sensor calibration should be changed from 0.36 to 0.38. Since considerable processing had already been done and there was no sampling below 425m, so hysteresis is not a problem.

##### Conversion of Full Files from Raw Data

All hex files were converted using the configuration files mentioned in the previous section to create CNV files.

An error was found in the file name for the first CTD cast at station 38. Filename 2017-64-0051 was changed to 2017-64-0052. Cast 351 was a NET cast.

A few casts were examined.  All expected channels are present. The primary and secondary temperature and conductivity channels are close during downcasts, but as usual the upcasts differ more due to noise in both channels. Dissolved oxygen, fluorescence, PAR, SPAR and altimetry profiles all look normal. The pH traces look ok except for cast #19 which has very low values.  There is no note about this but it is likely that the soaker bottle was left on. The pH channel should be removed for that cast; all casts were checked and the pH looks ok in all the others.

Two transmissometers were used during this cruise, but there is no note about why. It may have been changed in order to do a comparison since there was a return to many sites that were visited earlier in the cruise. Sensor #498DR had been serviced recently and the traces look ok. Sensor #1396 was used in late 2016 without a problem, but in recent cruises it showed an odd profile as though values got stuck for a half-second or so, then registered a few new values with the usual sampling rate then got stuck again. That same pattern is seen for this cruise. There were systematic differences in values with sensor #1396 having a smaller range of values, reading lower near the surface and mostly higher at greater depth; however, the two sensor readings were close at the bottom. There are also some differences in temperature, so while it is clear that sensor #1396 has a problem, it is not clear if the values are poor.

As usual the descent rate was usually high and steady, except in the western part of Juan de Fuca Strait, where it was very noisy with many complete reversals of direction. It was a little noisy towards the end of the cruise.

##### BOTTLE FILE PREPARATION

ROS files were created using files 2017-64-ctd1.xmlcon and 2017-64-ctd2.xmlcon. There were no rosette casts corresponding to the 3rd con file.

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and a few outliers were found in casts #10 and 95. CTDEDIT was used to clean channel Salinity:T0:C0. The output files were copied to *.BOT.

There were some errors in event numbers on rosette logs. The sheet for station 72 indicates it was event #13 but should be #14. The sheet for station 3 indicates it was event #71 but should be #73.

There was some duplication of sample numbers in the rosette logs. This was complicated by some labels not matching the log entries.

- Event #27 had the right sample numbers on the rosette sheet but the labels were wrong. This only affects CHL and the analyst dealt with that by changing the numbers to match the log.

- Event #68 is the very bizarre one. On the rosette log sheet sample #219 is followed by 210, 211, 212, 213 and 214 which had already been used for that same event. It appears that the oxygen and nutrient samples were given labels 220, 221, 212, 223, 224 as expected, which would have been ok EXCEPT that those numbers got used again during event 73. And since sample 214 was the last one listed on the log for event 68, samples 215-219 got repeated on the casts that followed which affect CHL, SAL and NUTS. So to remove duplication and to match what the CHL analyst did, a leading 9 should be added to the sample numbers actually recorded on the rosette sheet, as follows:

  - Entry in spreadsheet 215 – leave – Tamara fixed 2nd incidence in event 69 and there was only CHL sampling

  - Entry in spreadsheet 216 – leave – Tamara fixed 2nd incidence in event 69 and there was only CHL sampling

  - Entry in spreadsheet 217 - leave – Tamara fixed 2nd incidence in event 69 and there was only CHL sampling

  - Entry in spreadsheet 218 – change to 9218

  - Entry in spreadsheet 219 – change to 9219

  - Entry in spreadsheet 220 – change to 9210

  - Entry in spreadsheet 221 – change to 9211

  - Entry in spreadsheet 222 -  change to 9212

  - Entry in spreadsheet 223 -  change to 9213

- Event 73 will be ok after step 2 is taken. It has numbers that currently repeat only because of how the DO and OXY samples were labelled for event #68. (Note: event #73 may be entered as #71 due to the error mentioned above.)

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv.

Sample numbers were added to the file based on the rosette log records. There were a few problems beyond the ones mentioned above.

- For cast #7 16 bottles were fired but only 15 are listed on the sampling log. The daily science log notes that an extra surface bottle was fired because of concerns about Niskin #15.

- For cast #10 bottle #15 was fired but not sampled and has no sample number.

- For cast #27 the sampling log indicates that Niskin #1 was not fired; it was, but was not sampled and no sample number was assigned to it.

- Event #61 – Niskin #18 fired but no sample number assigned, no sampling done.

- Event #68 – Niskin #17 fired but no sample number assigned, no sampling done.

- Event #89 – test trips – no sampling – dropped from list.

- Event #94 – Niskin #17 fired but no sample number assigned, no sampling done.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine. Those files were bin averaged on bottle number.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2017-64-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF 2017-64CHL*.xlsx. The file included comments, flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2017-64chl.csv, event numbers were added, and the file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2017-64oxy*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2017-64oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in QF2017-64SAL*.xlsx. There were no duplicates. The analysis was done within 14-19 days of collection. The files were simplified and saved as 2017-64sal.csv.

That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet NUTS_QF2017-64nuts.xlsx. This includes a precision study. The spreadsheet was simplified, reordered on sample numbers and saved as 2017-64nuts.csv.

The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only, with output files named MRGCLN1.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number as the merge channel.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for problems. None were found.

CLEAN was run on the MRG files to add 0 flags to empty flag channels and to update header limits.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There were many outliers with most coming from the top 100m where poor flushing of Niskin bottles is likely a factor. There are 3 outliers that have CTD salinity values higher than the bottle salinity, and poor flushing would not explain those.

- Cast #6, sample 84. The standard deviation in the CTD salinity is high and the bottle was fired near the bottom, so any error due to incomplete flushing is likely to have the opposite sign to the usual.

- Cast #34 – There are 4 bottles that look out of line; for all, the standard deviation in the CTD salinity is high.

- Sample 122 – 200m – right at the bottom so having the CTD salinity higher makes sense.

- Sample 127 – 75m. Very high standard deviation and a local reversal in salinity probably account for the large difference and the CTD reading higher than the bottle sample.

- Sample 130 – 30m. Very high standard deviation in CTD salinity plus poor flushing likely account for the difference.

- Sample 133 - 5m. Very high standard deviation. There is a salinity reversal around 5m, so the CTD being higher can be explained easily.

There is no justification for changing any flags.

The analysis was done promptly after collection except for 2 samples.

When outliers were excluded based on differences >0.01psu and pressure <100db, the primary salinity was found to be low by an average of 0.0018psu and the secondary was high by 0.0012psu; the standard deviations were 0.005 and 0.006, respectively.  It is expected that the CTD should read lower due to flushing problems in the Strait of Georgia. Data from Juan de Fuca Strait were examined since that area had the noisiest descent rate, so the Niskin bottles were likely better flushed. The 4 bottles from the most westerly casts were examined. One had a very high standard deviation in the CTD salinity and another had a fairly high standard deviation. A third was sampled right at the bottom which tends to produce poor results. The only sample that looks reliable shows the CTD primary salinity to be low by 0.0012psu and the secondary salinity to be high by 0.0034. This is very little evidence, but it is in line with the results of 2017-06, an offshore cruise using the same equipment that occurred just before this cruise.

For full details for the COMPARE run see file 2017-64-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

First outliers were investigated for which the CTD DO was lower than bottles by more than 0.04 or were higher than bottles. Most of the cases with CTD reading higher than bottles came from close to the surface and poor flushing of Niskin bottles likely explains this. Some other outliers are associated with large standard deviations in the CTD DO. There were 3 outliers that were investigated:

- Event 34, Sample 127 – The bottle from 75m has a fairly high standard deviation in the CTD DO but likely not enough noise to explain this result. The bottle DO profile looks odd, with either the 75m or100m bottle being out of line. The salinity and nutrients (except silicate) from the 75m bottle also look out of line with values that appear to be from around 20m rather than 75m. The silicate profile is too complex to be helpful in this judgment. The draw temperature also looks a little out of line, but that is not very reliable evidence. So this might suggest that there was a misfire of Niskin #7. However, examination of the temperature profile suggests that this may not be a misfire, but rather the effect of having a section from about 60m to 80m that is barely stable and includes a reversal in temperature, salinity and dissolved oxygen. There would appear to be active mixing and this likely explains the outlier at 75m.

- Event 42, Sample 139 – The DO sample looks out of line in profile and the value looks higher than expected compared to CTD DO. No other samples are out of line. It doesn’t look close to any other samples, so doesn’t look like a miss-sample. There are signs in the T-S plot of mixing going on at about 110m, and during the downcast there was a quite a large reversal in DO at that level. On the upcast there are similar features but not as large, but they may be affected by shed wakes. This looks as likely to be real as not, with some incomplete flushing of bottles and slow response of the CTD DO sensor complicating any interpretation.

- Event 61, Sample 206 – The DO was flagged due to replicate outliers. The standard deviation in the CTD DO is very high with a reading of ~7 during the window around firing time, but during the stops values ranged from ~7 to ~8, so either  bottle value could be correct and quite possibly both are ok.

When all casts were included and the 3 outliers above excluded as well as other outliers based on residuals, the fit was:

CTD DO Corrected = CTD DO * 1.0267 + 0.086    R² = 0.70 (1)

Using only the 6 casts from the central Strait of Georgia, the fit was:

CTD DO Corrected = CTD DO * 1.0210 + 0.086    R² = 0.63 (2)

Using only the 4 casts from the western Juan de Fuca Strait, the fit was:

CTD DO Corrected = CTD DO * 1.0437 + 0.048    R² = 0.82 (3)

If more outliers were excluded based on residuals the fit for Juan de Fuca looks like:

CTD DO Corrected = CTD DO * 1.0385 + 0.059    R² = 0.94 (4)

This is close to the result when the sensor was last used during 2017-06 with many offshore casts where flushing is likely to be good:

CTD DO Corrected = CTD DO * 1.0399 + 0.065    R2 = 0.96 (2017-06)

The reduced correction in the inshore parts of the cruise is likely due to poor flushing of bottles so they contain water from deeper in the water column than the CTD data with which they are being compared. Lower bottle DO would reduce the correction. These variations are consistent with other cruises during which there was a comparison between waters where flushing is likely to be good in one part and poor in another. The noisy descent rate in Juan de Fuca Strait indicates that flushing was likely to be good there whereas it is very quiet in the Strait of Georgia. So the Juan de Fuca result is considered the most appropriate correction to apply to all the data.

For more detail see document 2017-64-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further problems were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor. There were 74 samples with a range of ~0.1ug/L to ~11.3ug/L. The fit of Fluorescence against extracted CHL is:

SBE Fluorescence = 0.77 * CHL

This is fairly typical for these fluorometers when there are NO high CHL values.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

For the past 2 cruises when this sensor was used, the best setting for ALIGNCTD was found to be +2.5s, This setting worked well for this cruise as well. ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. The default setting of was selected and it does improve the data for both conductivity channels for these data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on 4 of the deeper casts to calculate differences between sensor pairs. None of the casts are very deep.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2017-63-0074 | 400 | -0.0001 | +0.00037 | +0.0040 | F.High, Steady |
| 2017-63-0107 | 320 | -0.0006 N | +0.00040 N | +0.0045 N | F.High, Steady |
| 2017-05-0038 | 1000 | +0.0001 | +0.0004 | +0.0046 | High, X Noisy |
|  | 1900 | +0.0003 | +0.0004 | +0.0040 |  |
| 2017-05-0101 | 1000 | +0.0006 | +0.0005 | +0.0055 | High, Moderate |
|  | 1900 | +0.0004 | +0.0005 | +0.0052 | “ |
| 2017-05-0139 | 1000 | +0.0002 | +0.0005 | +0.0051 | High, X Noisy |
|  | 1900 | +0.0004 | +0.0005 | +0.0055 | “ |
| 2017-06-0036 | 400 | +0.0002 | +0.0004 | +0.0045 |  |
|  | 1000 | +0.0002 | +0.0004 | +0.0047 | High, X.Noisy |
|  | 1900 | +0.0004 | +0.0004 | +0.0045 |  |
| 2017-06-0064 | 400 | +0.0002 | +0.0004 | +0.0047 | High, Noisy |
|  | 1000 | ~0 | +0.0004 | +0.0047 | “ |
|  | 1900 | +0.0003 | +0.0004 | +0.0042 | “ |
| 2017-06-0085 | 400 | +0.0001 | +0.0004 | +0.0043 |  |
|  | 1000 | +0.0002 | +0.0004 | +0.0043 | High, V Noisy |
|  | 1900 | +0.0002 | +0.0003 | +0.0040 | “ |
| 2017-64-0056 | 400 | -0.0005 | +0.0003 | +0.0045 | High, F.Steady |
| 2017-64-0060 | 400 | -0.0005 | +0.0005 | +0.0047 | High, F.Steady |
| 2017-64-0066 | 400 | -0.0004 | +0.0004 | +0.0047 | High, Steady |
| 2017-64-0104 | 400 | -0.0005 | +0.0004 | +0.0042 | High, Noisy |

The temperature differences have the opposite sign to those during the most recent cruise, but are similar to those of 2017-63. This is likely due to the temperature increasing around 400m during 2017-63 and 2017-64 whereas it was decreasing for the other cruises, so slight misalignments could lead to errors of opposite sign. During bottle stops the differences were often very noisy due to a lot of vertical motion; in cases where the vertical movement of the CTD was slight, the differences were small and of variable sign. The differences in salinity and conductivity are similar to those during other recent cruises.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A cross-reference list was checked against the log book and no errors were found. (The only error had been fixed earlier based on a log note.)

The cruise track was plotted and added to the end of this report.

Header Check was run and shows that pressures were occasionally <0db.

- The lowest pressure in the CTD files was about -0.3db during cast #104 and pumps were on then. Plots show that the CTD was moving up and down for ~6s, with pressure <0db for ~1s. It is possible that the CTD was briefly out of the water, but not long enough to go to very low conductivity. There was only one other case of negative pressure and it involved about 6 scans at the end of the upcast.

- The highest fluorescence came from spikes during cast #100 but these were only during upcast stops, and were not during the 10s-window for the bottle file.

The Surface Check gave the average surface pressure as 1.9db which is within the normal range for the Vector. The pressure sensor was checked at the factory in March; recalibration does not look appropriate.

The altimeter and bottom depth headers were exported from the SAMAVG and CLN files to spreadsheets. The bottom depths were checked against the log book. There were many minor discrepancies but a few were large enough to need investigation. The altimetry header entries were checked by calculating:

Check Value = (Bottom depth – maximum depth of sampling + altimetry header)

The following cases were examined because either the Check Value ≥4m or bottom depths differed significantly from log entries:

- Cast 1- CTD did not get near the bottom. The altimetry header was caused by a spike and should be removed.

- Casts 5, 29, 30, 33, 34, 43, 46, 53, 55, 87, 96, 103 and 104 – Check Values ≥4m and making water depth entry match log improved the Check Value so the changes were made.

- Cast 97 went deeper than the bottom depth entered in log and file; changed based on bottom depth + altimetry.

- Casts 16, 18, 19, 105 – Water depths differ from log by at least 2m and improve check values, so these were changed

- Casts 79, 85 - No altimetry entry due to spike and depth wrong for 79. Enter estimates based on plot and fix depth to match log for 79.

- Cast #78 – No altimetry due to spikes but plot also unclear, so no estimate.

- A few casts had Check value ≥4m but the altimetry and water depth entries seem ok – likely the depth changed during the casts, but no change was made to entry in files.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and as usual, there is variability with the up and downcast traces sometimes closer than temperature and sometimes further apart. This is likely due to varying vertical gradients and descent/ascent speeds. Overall the choice made earlier looks appropriate, so no further alignment will be applied.

Conductivity

Tests were run on 4 casts and the best results were with -0.6 records for the primary and -0.5 for the secondary conductivity. SHIFT was run twice on all casts using -0.6 records for the primary and -0.5 for the secondary conductivity.

pH

Tests were run on 3 casts to determine the best shift to make the difference between upcast and downcast pH profiles similar to that for temperature. A setting of +50 records produced the best results.

SHIFT was run to advance the pH channel by 50 records for all casts.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warnings concerned the upcast sections of 2 files, so are of no concern.

##### Other Comparisons

Previous experience with these sensors –

All sensors were used during 4 earlier cruises in 2017.

2017-63 in early April was a cruise in the Strait of Georgia and Juan de Fuca Strait. The bottles from Juan de Fuca were trusted more than those from the Strait of Georgia because conditions were better for flushing Niskin bottles; they indicated that the primary salinity was close to bottles while the secondary was high by about 0.0045. The dissolved oxygen comparison from 2017-63 was also based on Juan de Fuca bottles.

Cruise 2017-03 was in Baynes Sound and the Strait of Georgia, with shallow sampling and poor flushing of bottles, so the results were not considered reliable enough to use for recalibration.

Cruise 2017-05 was mostly off the west coast of Vancouver Island where flushing was good at depth. The comparison between the bottles and CTD salinity was quite tight; the primary salinity was found to be very close to bottles and the secondary was high by 0.0054psu. The dissolved oxygen comparison was also quite tight and led to a calibration of slope 1.0448 and offset +0.0527.
Cruise 2017-06 was mostly well offshore and showed the primary salinity to be low by ~0.0012 and the secondary high by ~0.0032. The dissolved oxygen was recalibrated using slope 1.0399 and offset 0.0645.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Almost all temperature data fell within the climatology, with just a few brief shallow values that were a little high. Salinity was low in parts of some profiles in Juan de Fuca Strait and Haro Strait, mostly between about 10 and 100m. In the Strait of Georgia there were also some low values but closer to the surface, and none were seen for the later casts in the central part of the Strait. These excursions are not likely to be due to calibration drift, but rather reflect real variability. There were some instances of low salinity in the same areas in April 2017.

Repeat Casts – There were repeat casts but they are too shallow and too far apart in time to be useful for judging repeatability. During the previous cruise when there were deep casts at the same site within 7 hours differences along lines of constant σt were almost indistinguishable before bin-averaging and after binning they were <0.002º for the primary temperature and <0.0002psu for the primary salinity at σt= 27.52 (near 1450m). This shows excellent repeatability.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

Examination of primary and secondary T and S profiles and T/S plots show that while noise levels were similar for both channel pairs for most of the profiles, there were a number of casts which had poor primary salinity near the surface, suggesting near-surface pump problems.

The secondary T and S channels were chosen for editing.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes. All files required some editing, but mostly it was removal of records from near the top and bottom of casts and light cleaning of salinity.

##### Initial Recalibration

- There is no evidence that pressure needs recalibration.

- Salinity:T0:C0 is reading lower than bottles by an average of 0.0018 using bottles from all regions and by 0.0012 using the bottle in Juan de Fuca Strait with the most reliable comparison. During the previous cruise in deep offshore waters, the primary was found to be low by an average of 0.0012psu based on many bottles

- Salinity:T1:C1 is reading higher than bottles by ~0.0012psu using bottles from all regions, but is higher by 0.0034 than the 1 bottle in Juan de Fuca Strait with the most reliable comparison. During the previous cruise in deep offshore waters, the secondary was found to be high by an average of 0.0032psu based on many bottles.

- Salinity was often below the climatology minimum, but that may reflect the age of the climatology and the limited applicability to near-shore values. The same sensors performed well during the previous cruise and the bottle comparison does not suggest low secondary salinity.

- For dissolved oxygen the fit based on Juan de Fuca Strait is:

CTD DO Corrected = CTD DO * 1.0385 + 0.059    R² = 0.94 (4)

This is close to the result when the sensor was last used during 2017-06 with many bottles from offshore casts where flushing is likely to be good:

CTD DO Corrected = CTD DO * 1.0399 + 0.065    R2 = 0.96 (2017-06)

File 2017-64-recal1.ccf was prepared to add 0.0012psu to the primary salinity, subtract 0.0032 from the secondary salinity and to recalibrate dissolved oxygen using.

CTD DO Corrected = CTD DO * 1.0399 + 0.065

CALIBRATE was run on the SAM and MRGCLN2 files using 2017-64-recal1.ccf.

COMPARE was run using the recalibrated salinity data and confirm that the recalibration was done correctly; the average differences for primary and secondary salinity are within 0.0001 of each other.

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly.

See files 2017-64-sal-comp2.xlsx and 2017-64-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed based on residuals the CTD DO was higher than the bottles by an average of ~0.06mL/L and standard deviation of 0.05mL/L. When the same points are displayed in a plot versus pressure, it is seen that the CTD is highest relative to bottles near the surface and the differences disappear at about 300m. This makes sense if there is incomplete flushing and vertical DO gradients decrease with depth, which is roughly the case. The largest outliers come from areas where there are some very high gradients. There are many DO reversals to complicate the comparison.

A judgment on the calibration quality is not clear, but there is no indication that further recalibration of DO is needed.

##### Fluorescence Processing and special files for Angelica Peña

The COR1 files were clipped to 150db and processed in 2 ways, with a filter and without a filter, followed by 0.5m-bin averaging in both cases.

The CTD files from rosette casts were clipped to 50m; sigma-T was derived and the data were exported to a single file, 2017-64-SOG.csv.

Those files were set aside for Dr. Peña.

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some small unstable features, but those are from sites where active mixing is expected.

The pH signal looks bad for event #19 and odd for event #66. For the latter there are large sections of data full of pad values. The problem with #66 arose because that cast was rerun late in processing, and the pH channel was included in the channels put through the WILDEDIT stage; the file was rerun through all steps again correctly. This removed the problem for #66.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Transmissivity, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

For cast #19 the pH:SBE channel was also removed.

The transmissivity from casts 101 to 110 was left in the files.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

A note was added to cast #36 about the removal of the pH channel. (This was run on *TR files too.)

Data Processing Notes:

----------------------

Transmissivity, Fluorescence:URU:Seapoint, pH:SBE, PAR and PAR:Reference data are

nominal and unedited except that some records were removed in editing temperature

and salinity.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

The pH:SBE channel was removed from cast #19 because very low values suggest that

the soaker bottle was left on.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.2 mL/L from 0 to 150db

±0.08 mL/L from 150db to 300db

±0.01 mL/L below 300db

Warning: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

Warning: The transmissometer used for casts #101 to #108 alternated between sampling

at the normal rate and getting stuck on a single value for about a half second.

These data are considered of lower quality than usual, though the general

shape of profiles looks reasonable.

For details on the processing see document: 2017-64_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values ranged from 60% to 140% with lower values in Juan de Fuca and Haro Straits, and higher values in the Strait of Georgia. There is no evidence of a DO sensor calibration problem, though with so a large range of values it does not indicate that the calibration is good either. During the previous cruise the results of this test indicated that calibration looks good.

##### Final Bottle Files

CALIBRATE was run on all files and then the MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Plots were made of all casts to look for problems and it was found that some nutrients and chlorophyll samples had the wrong event number, so were missed in the merge process. Those steps were rerun.

Cast #89 had no bottle sampling, so the CHE file was removed.

Cast #55 had particle sampling only; a file was created in case it is needed later to go with those samples.

Standards check and a header check were run on all files. No errors were found.

The track plot looks ok.

Data from the CHE files were exported to a spreadsheet and compared with rosette sheets. No problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

##### Thermosalinograph Data

There were no loop samples, flow meter or intake thermistor. The intake is at about 2m. The only method to check calibration is to compare with the CTD casts. There are many more CTD data available at 3m than 2m, so data from 3m will be used as well as 2m.

a.) Checking calibrations

The 2 configuration files were identical and had no errors; one was saved as 2017-64-tsg.xmlcon.

b.) Conversion of Files

The 2 hex files were converted using file 2017-64-tsg.xmlcon.

They were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

Time-series plots were produced. The traces look fine.

The track plot looks fine.

c.)  Checking Time Channel

The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.3db of 3db, and then again using 2db. These were exported to 2 spreadsheets which were saved as 2017-64-ctd*-tsg-comp.xls. There were 65 and 81 casts, at 2db and 3db respectively, which overlapped with TSG files.

The TSG file was opened in EXCEL, median and standard deviations (over 5 records) were calculated for lab temperature and salinity and the files were reduced to the times of CTD files. Those data were added to as 2017-64-ctd3-tsg-comp.xls.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. There were no position differences >0.0004º and the median differences were both 0.0001º. So the TSG clock worked well.

d.) Comparison of T and S from TSG and CTD data

The differences between the TSG and CTD were large especially in the Strait of Georgia where the surface gradients were generally higher. The plot below is based on the differences at 3m. Using all casts the temperature is high by a median of 0.57C°and the salinity is low by a median of 1.06psu, with standard deviations that are about 2 times the median values.

The 2m comparison is very similar to that at 3m.

When differences are picked out from the sections with fairly well-mixed surface waters, as judged from CTD profiles, the mean and average values are much closer than when the whole data set is used.

|  | Tmed-Tctd3 | Smed-Sctd3 | Tmed-Tctd2 | Smed-Sctd2 |
| --- | --- | --- | --- | --- |
| average | 0.1496 | -0.5461 | 0.1407 | -0.5676 |
| median | 0.1333 | -0.5080 | 0.1233 | -0.5216 |
| Std deviation | 0.0878 | 0.1262 | 0.0892 | 0.1431 |
| # of casts | 11 | 11 | 8 | 8 |

For these well-mixed comparisons, the salinity is low by a median of about 0.5psu while the temperature is high by a median value of 0.12C°.  We expect the lab temperature to be higher than the CTD, though we do not have sufficient experience with a TSG on the Vector to predict by how much, though similar results were seen in some other cruises. As well as the heating effect, some of the difference may be due to the TSG drawing water from a little shallower than 2m, or due to the CTD temperature being of lower quality near the surface. The salinity is reading much lower than we expect which could also be due to the TSG drawing from higher in the water column or poor CTD data. The presence of small bubbles in the water could also add to the differences, but there is no evidence of that in the time-series plots so large bubbles are not likely.

There is too little evidence to support recalibrating the TSG salinity. It may well reflect actual conditions closer to the surface than that measured by CTD data. This would also explain some of the temperature difference, so no attempt will be made to calculate a proxy for intake temperature, though assuming it is slightly high is a valid conclusion given that the water must warm up a little in the loop.

(See 2017-64-ctd-tsg-comp.xls.)

Calibration History

The TSG temperature and conductivity were recalibrated in December 2015 and have been used for 3 other cruises: 2016-07, 2016-10 and 2016-71, 2017-63 in the same region.

For 2016-07 (a June cruise) the lab temperature was higher than the CTD temperature by between 0.1C° and 0.2C°.  Salinity was found to be low by 0.62 using all casts, but the standard deviation was very high at 3.0psu. The median difference was 0.03psu when only the 10 casts with the lowest standard deviation in the TSG salinity were included.

For 2016-10 in September the CTD temperature was higher than the CTD temperature by a median of 0.127C°. The TSG salinity was lower than the CTD salinity by a median value of 0.93psu using only the 20 casts with the lowest standard deviation in the TSG salinity. The differences were slightly lower when the 2db data were used but there were fewer data available.

During 2016-71 in November the TSG lab temperature was higher than the CTD temperature by about 0.13Cº. The TSG Salinity was lower than that from the 3m CTD by about 0.05psu. Calibration drift is unlikely to be that large. It is more likely that the difference is due to small bubbles in the loop and or that water is being drawn from above 2m. Bubbles are not obvious in the TSG traces, but small ones could be present.

During 2017-63 the TSG lab temperature was higher than the CTD temperature by about 0.2Cº and the TSG Salinity was lower than that from the 2m CTD data by about 0.04psu.

Conclusions

1. The TSG clock worked well.

2. The usefulness of comparisons with CTD data are limited by highly variable vertical and horizontal gradients.

2. The TSG lab temperature was higher than the CTD temperature by about 0.12º. We have little experience with the Vector loop system and no knowledge of the flow rate during the cruise. Finding less heating in the loop than during the April 2017 cruise is consistent with the intake temperatures being higher in June. It is possible that some of the apparent heating in the loop is really due to water being drawn from a little higher in the water column than 2m.

3. The TSG Salinity is lower than that from the 2m CTD data by about 0.5psu. This is a much larger difference than in the previous cruise, but even larger differences were seen in September 2016 even using only the 20 casts with the lowest standard deviation in the TSG salinity. A high surface salinity gradient together with the TSG drawing water from a little above 2m may account for the large median difference. Calibration drift is unlikely to be that large. It is also possible that small bubbles in the loop water have lowered salinity slightly. Bubbles are not obvious in the TSG traces, but small ones could be present.

4. Given the great variability in comparisons over the past year, it appears that near-surface gradients dominate these comparisons. Salinity gradients are frequently large in spring.

5. The quality of the data was good with the only spikes looking real and there is no evidence of problems in the flow rate.

6. There is insufficient evidence to justify making an estimate of intake temperature or to recalibrate salinity.

f.) Editing

No editing was required.

g.) Recalibration

No recalibration was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Conductivity:Primary, Flag and Position:New channels.

HEADER EDIT was used to change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats and to add the following comment:

TSG Data Processing

-------------------

There was no loop sampling, intake temperature sensor or flow rate meter.

A comparison of temperature and salinity was made with co-incident CTD casts

but high near-surface vertical gradients resulted in a lot of variability.

The TSG salinity is lower than that from the CTD salinity with a median difference

of about 0.5psu compared to CTD data from 2m. This may be due to bubbles

in the loop water or the loop may be drawing water from above 2m. There could

also be some error due to calibration drift but that is likely to be negligble.

Similarly, there was considerable variability in the differences between the

TSG temperature and that from the CTD. The TSG temperature appears to be high

by about 0.12 C degrees. This is presumed to be due primarily to heating in the

loop, but may be partly due to the loop drawing water from above 2m.

The Temperature:Primary was renamed as Temperature:Lab to make it clear that it is not the intake temperature.

The files were saved as 2017-64-000*.TOB.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series and all look fine.

The cruise plot was added to the end of this report.

Particulars (mostly notes from log)

7-11. Rough seas

9. Possible error in station name in file

16. Fluorescence spiking during upcast.

21. Conductivity spikes – possibly plumbing blockage.

29. Ship drifted with current during cast; repairs to trip mechanism, many spigot o-rings.

66. NMEA error in Seasave – use log positions

89. Bottles fired as test – no bottle file needed.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 22Feb2017 | Factory |  |  |
| Conductivity | 1766 | 30Mar2017 | Factory |  |  |
| Secondary Temp. | 4484 | 22Feb2017 | Factory |  |  |
| Secondary Cond. | 3531 | 30Mar2017 | Factory |  |  |
| Transmissometer | 498DR | 21Jun2017 | IOS |  |  |
| Transmissometer | 1396DR | 2Feb2016 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 19Dec2015 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| Surface PAR | 20518 | 21Mar2016 |  |  |  |
| pH | 0691 | 15Jan2015 | Factory |  |  |
| SeaPoint Fluor. | 3685 | n/a |  |  |  |
| Pressure Sensor | 0550 | 24Feb2017 | Factory |  |  |
| Altimeter | 62355 |  |  |  |  |

# CRUISE SUMMARY     TSG

Make/Model/Serial#:	SEABIRD/21/3363       Cruise ID#:	2017-64

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3363 | 17Dec15 | Factory |  |  |
| Conductivity | 3363 | 17Dec15 | Factory |  |  |