## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 26 Nov 2021 | Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| 3 October 2020 | HPLC data added. |

## PROCESSING NOTES

Cruise: 2017-65

Agency: Ocean Sciences Division

Location: Strait of Georgia and Juan de Fuca Strait

Project: Strait of Georgia / Juan de Fuca Strait Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: October 3, 2017 – October 9, 2017

Processed by: Germaine Gatien

Date of Processing: 22 February 2017 – 6 March 2017

Number of original HEX files from SBE911: 87 (1 contains surface data only)

Number of XML files from SBE25: 5

Number of CTD files: 91	Number of bottle files:  27

# INSTRUMENT SUMMARY

Two SeaBird CTDs were used during this cruise:

- A SeaBird Model SBE 911+ CTD (#0550) was mounted in a rosette and attached were a SBE 43 DO sensor (#1119), a SeaPoint Fluorometer (#3865) with a 3X cable, a Biospherical QSP-400 PAR sensor (#4565), Surface PAR (#20518) a pH sensor (#0691), an altimeter (#62355) and a Wetlabs CSTAR transmissometer (#1185DR). The data acquisition program was Seasave v7.26.2.13.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 68572.

There were 24 10L bottles mounted on an IOS Rosette.

- A SeaBird Model SBE25+ was used for 5 casts; attached were a SBE 43 DO sensor (#3234) and a WET labs ECO Fluorometer (#2214).

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log book was in good order with many useful comments.

An SBE911+ CTD was used for most casts. An SBE25 was used for a few casts with no bottle sampling as it was not mounted in the rosette.

Salinity analysis was done very promptly. The results were out of step with other recent cruises, with the primary salinity reading high by about 0.002psu and the secondary by about 0.010psu. Examination of data during stops shows that the CTD was falling during many of the stops and there were some salinity reversals. These features would explain a small part of the difference but probably no more than 0.001psu. Poor flushing usually makes the CTD look low but having the CTD drop during casts and the complexity of the T and C profiles may reduce that effect. The secondary salinity calibration was very different from earlier cruises and suggests a problem with the conductivity sensor, so the primary channels were selected for archiving.

The calibration of the dissolved oxygen sensor was based on the results of the bottle comparison in Juan de Fuca Strait as there is evidence of incomplete flushing of Niskin bottes in the Strait of Georgia.

There has been heavy use of these sensors in 2017 and there is evidence of significant calibration drift in some of them.

Channel Oxygen:Dissolved:SBE was recalibrated.  The Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.6 mL/L from 0 to 20db

±0.3 mL/L from 20 to 100db

±0.1 mL/L from 100db to 150db

±0.03 mL/L below 150db

The fit of CTD fluorescence versus extracted chlorophyll showed the usual pattern of fluorescence reading higher than CHL for low CHL and then dropping relative to extracted CHL as CHL values rose. CHL values ranged from ~0.2ug/L to ~22.3ug/L.

File 2017-65-0077 contained upcast data only and was created just to fire a bottle that was missed. The bottle data are included in file 2017-0065-0076.CHE and there is no CTD data file. The pumps were not turned on when the bottle was fired so the CTD temperature, salinity, fluorescence and dissolved oxygen from the surface record may not be as reliable as other CTD data in this file.

The thermosalinograph could not be used because the water pump was unserviceable.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.  There are many notes about problems encountered and the CTD casts using the SBE25 are easy to identify.

Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

The cruise summary sheet was completed.

The sensor histories for the pressure, conductivity and DO sensors were found.

There was a change in the configuration file before cast #8; factor E in the dissolved oxygen calibration was updated from 0.036 to 0.038 based on a study during 2017-06. However, a subsequent study from 2017-08 provided a further refinement of that factor so it was changed to 0.0375 at the processing stage. This was done just for consistency with other records; this cruise has no sampling deep enough for this to matter.

The calibration constants were checked for all instruments and the only other correction needed was to add the day and month to the date for the transmissometer calibration.

No configuration file was provided for the SBE25. One for an earlier cruise was obtained.

There were 8 XML files, 5 with non-standard names. 3 of those had been copied and renamed in standard format except that the cruise number was wrong. The remaining 2 appear to correspond to the last 2 SBE25 casts but the start times differ from the log time by 1 hour each. Those 2 were copied and renamed in standard format.

##### Conversion of Full Files from Raw Data

All HEX files and XML were converted using the configuration files mentioned in the previous section to create CNV files.

A few casts were examined. As mentioned in the log there are some large pressure spikes. A test of WILDEDIT shows that it will greatly reduce the spikes though it won’t completely eliminate them. All expected channels are present. The primary and secondary temperature and conductivity channels are close during downcasts, but as usual the upcasts differ more due to noise in both channels. Dissolved oxygen, fluorescence, pH, PAR, SPAR and altimetry profiles all look normal.

The descent rate was usually high and steady, except in the western part of Juan de Fuca Strait, where it was high but noisy. As noted in the log book the rate was unusually high for some casts.

For the SBE25 casts all expected channels are present. The fluorescence traces are very noisy but there is little variation in the profiles which exaggerates the spiky appearance and not being pumped probably reduces mixing so that small scale variations are visible.

##### BOTTLE FILE PREPARATION

There was no sampling during casts with the SBE25.

ROS files were created using files 2017-65-ctd.xmlcon. The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers and a few outliers were found in cast #98. CTDEDIT was used to clean channel Salinity:T0:C0. The output file was copied to *.BOT.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv.

Sample numbers were added to the file based on the rosette log records. Two problems were encountered in this process

- The rosette sheet for event #76 indicates that 3 bottles were fired and samples were analyzed with reasonable results. But the CTD file for #76 contains no data from the surface sample from Niskin #3. A second file, #77, was started to fire the surface bottle. So the two BOT files were named 76.bot1 and 76.bot2 and then combined using the JOIN Files routine. The resulting BOT file was opened in Ultraedit to change the Bottle firing order number for Niskin #3 and the header for that channel limit was changed. The ADDSAMP file was adjusted to match.

- For event 101 the sample numbers 287 and 288 were skipped. The sample labels did not skip those numbers, so the addsamp file was adjusted to match the labels.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

SAM files were created using the Add Sample Number routine. Those files were bin averaged on bottle number.

NOTE: Late in the processing it was noticed that pH and PAR:Reference channels had not been converted for the bottle files. Files were converted with sample numbers and the missing channels and they were converted to SAMAVG2 files. The original SAMAVG files were renamed SAMAVG1. The SAMAVG1 and SAMAVG2 files were merged to create new SAMAVG files. The pH and PAR:Reference values from file #77 were entered into file #76 using a text editor.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2017-65-bot-hdr.txt which will be updated as needed during processing.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF 2017-65CHL*.xlsx. The file included comments, flags and a precision study. A simplified version of the spreadsheet was prepared in which some columns were removed and the file was saved as 2017-65chl.csv, event numbers were added, and the file was then converted to individual CHL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2017-65oxy*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and the file was then saved as 2017-65oxy.csv.

That file was converted into individual *.OXY files.

SALINITY

Salinity analysis was obtained in QF2017-65SAL*.xlsx. There were no duplicates. The analysis was done within 3-8 days of collection. The files were simplified and saved as 2017-65sal.csv.

That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet NUTS_QF2017-65nuts.xlsx. This includes a precision study. The spreadsheet was simplified, reordered on sample numbers and saved as 2017-65nuts.csv.

The file was converted to individual NUT files.

The SAL, CHL, OXY and NUT files were merged with CST files in 4 steps.

After the 4th step the files were put through CLEAN to reduce the headers to File and Comment sections only, with output files named MRGCLN1.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number as the merge channel.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for problems. The only problem noted is a discrepancy between the sample numbers on the rosette sheet and sample labels for event #101. The numbers on the rosette log sheet skipped 2 numbers. Since this affects both oxygen and chlorophyll it is easier to correct the rosette log rather than the analysis sheets. Samples 289 through 293 will be renamed as 287 to 291. The HPLC analyst was informed and already had the correct sample number.

CLEAN was run on the MRG files to add 0 flags to empty flag channels and to update header limits.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There is a lot of scatter in the comparison for both sensor pairs, but the most significant outliers are associated with noisy CTD salinity during the bottle stops. A quick check was made to see if pressure spikes had affected bottle stops and the only case where that could be a problem is for the 150m bottle during cast #9. Even there, the spike appears to have occurred at the end of the stop, so should not be a problem.

When all cases where the standard deviation in the CTD salinity was >0.0008psu are excluded, the primary salinity was found to be high by an average of 0.0014 and the secondary high by 0.0089, with standard deviations of 0.0020 and 0.0022psu. Compared with other recent cruises the primary CTD salinity is higher by 0.002 or 0.004 but the secondary is much higher, by 0.008psu or more.

The near-surface bottles that do not have high standard deviations in the CTD salinity show slightly smaller differences which could be due to incomplete flushing of Niskin bottles. There is one more surface bottle included in the primary than in the secondary. A comparison excluding both standard deviation >0.0008psu plus pressure <100m indicates that the primary salinity is high by 0.0023psu and the secondary by 0.0092psu.

Finally, only cast 9 was included since the CTD appears to have been moving a lot as judged by the descent rate of the CTD, so flushing of the Niskin bottles may be more efficient. There are only 5 bottles in the comparison (the bottle from the bottom and those above 100m were excluded). This shows the primary salinity to be high by 0.002psu and the secondary to be high by 0.010psu.

A close examination of the data during bottle stops shows that in several cases the CTD was falling during the stop and for cast #9 Niskin #1 was fired at the bottom. In both situations it is quite possible that the Niskin bottle could contain water from higher in the water column, thus making the CTD look like it was reading higher than it really is. For several bottles there is a lot of variability in salinity during the stops even if there was not during the 10s window. So the comparison results may be a little high, but probably by no more than ~0.001psu. The bottles from Juan de Fuca are considered less prone to poor flushing, though the descent rate was steadier than usual.

The small number of bottles fired and lack of duplicates make such a comparison somewhat unreliable, especially given the presence of reversals in the vertical salinity gradients. The extremely steady descent rate for some of the bottle casts might even result in some bottles fired during the upcast containing water from the downcast.

There are no significant outliers that are not explained by noisy CTD data.

For full details for the COMPARE run see file 2017-65-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There were no severe outliers. There is a lot of scatter due mostly to bottles from the top 20m where there is a lot of variability, some reversals and high standard deviations in the CTD dissolved oxygen data. The largest outlier below 20m had been flagged by the analyst.

When all casts were included and outliers were excluded based on residuals, the fit was:

CTD DO Corrected = CTD DO * 1.0799 + 0.0505    R² = 0.90 (1)

When only casts from the Strait of Juan de Fuca are included the fit is:

CTD DO Corrected = CTD DO * 1.0826 + 0.0447    R² = 0.90 (2)

When only casts from the central part of Georgia Strait are included, the fit is:

CTD DO Corrected = CTD DO * 1.0782 + 0.0568    R² = 0.89 (3)

It makes sense that the correction is higher for Juan de Fuca as there is probably better flushing of Niskin bottles there. In Georgia Strait the Niskin bottles likely contain water from a little deeper in the water column and the DO values would generally be lower there so the CTD might look higher relative to bottles. The Juan de Fuca result is considered the most appropriate correction to apply to all the data.

For more detail see document 2017-65-dox-comp1.xlsx.

Plots of Titrated DO and CTD DO against CTD salinity were examined. No further problems were found.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor. There were 57 samples with a range of ~0.2ug/L to ~22.3ug/L. The fit of Fluorescence against extracted CHL is:

SBE Fluorescence = 0.65 * CHL

This is fairly typical for these fluorometers when there are no very low CHL values. For the lowest values the CTD fluorescence is closer to the CHla and it reads lower as CHL increases, reading about 40% of CHL when CHL=22ug/L.

##### WILDEDIT

- For the SBE911 data program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

For cast #115, 2 runs through this step were required to remove bad data around 9db.

- A test run on the SBE25 casts showed no improvement, so the WILDEDIT step was skipped.

##### WINDOWS FILTER

This routine is sometimes applied to SBE25 data, but does not look necessary for this CTD, perhaps due to the higher sampling rate, 16Hz. Pressure can be filtered in the DELETE process. Temperature and conductivity are noisy but no more so than the SBE911 data, so it is hard to judge whether this is real or instrumental. So the Windows Filter was not run.

##### ALIGN DO

- For the SBE911: During the past 3 cruises when this sensor was used, the best setting for ALIGNCTD was found to be +2.5s for 2 and +3.5s for another. The temperature data is so noisy that it is difficult to judge but +3.5s looks better than +2.5s. ALIGNCTD was run on all casts using +3.5s.

- For the SBE25 data most of the casts are not suitable for judging alignment because there is either little vertical variation or the cast is too shallow. Two casts suggest that an advance is needed and a setting of +1s looks appropriate. ALIGNCTD was run on all casts using +1.0s.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. The default setting of was selected for both CTDs and it does improve the data for both conductivity channels for these data.

- For the SBE911 data CELLTM was run using (α = 0.0245, β=9.5) for both conductivity channels.

- For the SBE25 data CELLTM was run using (α = 0.3, β=8) for conductivity.

Comparisons of a few files processed with and without step CELLTM showed an improvement for both CTDs.

##### DERIVE and Channel Comparisons

For SBE911

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration. It is usually run a second time to calculate differences between channel pairs but since there are no deep casts, this step was skipped. However, for a few casts from recent cruises, differences at 350m were calculated for comparison with this cruise; they are in chronological order.

| Cast # | T1-T0 | C1-C0 | S1-S0 |
| --- | --- | --- | --- |
| 2017-36 | -0.0001 | +0.00044 | +0.0048 |
|  | -0.0004 | +0.00036 | +0.0051 |
|  | -0.0001 | +0.00046 | +0.0053 |
| 2017-08 | -0.0004 | +0.00009 | +0.0014 |
|  | +0.0008 | +0.00001 | -0.0007 |
|  | -0.0003 | +0.00003 | +0.0006 |
|  | -0.0003 | +0.00001 | +0.0005 |
| 2017-09 | -0.0003 | +0.00017 | +0.0022 |
|  | -0.0002 | +0.00006 | +0.0010 |
|  | -0.0006 | +0.00005 | +0.0012 |
| 2017-65-0060 | -0.0006 | +0.00070 | +0.0075 |
| 2017-65-0065 | -0.0004 | +0.00072 | +0.0069 |
| 2017-65-0076 | -0.0010 | +0.00060 | +0.0069 |
| 2017-65-0112 | -0.0006 | +0.00052 | +0.0060 |

- The temperature differences are small though slightly higher for 2017-65.

- The conductivity and salinity differences were lower for 2017-08 than for earlier cruises; this was thought likely to be due to drifts in opposite directions for the two conductivity sensors. Both CTD salinity channels were closer to bottles, so this did not seem to be just a change in one sensor pair. It looked like the primary conductivity was drifting upwards and the secondary downwards.

- For 2017-09 the temperature differences are similar to 2017-08, while the conductivity differences are a little higher than for 2017-08. So the slightly higher salinity differences look reasonable. The CTD/bottle comparison was harder to understand since the primary salinity was lower than bottles by about 0.002 while the secondary was close to bottles implying that the primary salinity was drifting downwards while the secondary had not changed. However, when the data are examined at about 350m both salinity channels appear lower by the same amount compared to 2017-08. So it is likely that different sampling patterns are relevant with more very deep and very shallow samples from 2017-08.

- Between 2017-09 and 2017-65 there is a very large change in conductivity and salinity channels. The bottle comparison around 350m shows little change in the primary salinity but a very large change in the secondary. So it is assumed that there was a sudden change to the secondary conductivity.

- These sensors have been used a lot and should be serviced.

For SBE25

- Program DERIVE was run on all casts to calculate primary salinity and dissolved oxygen concentration.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert both SBE 911+ and SBE25+ CNV files to IOS Headers.

The SBE25 files lack positions and station names. Those were added using a text editor.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

From this point on file #77 will not be processed.

At this point the 10m soak data were removed from the SBE25 casts. Plots were used to determine how many scans to remove; removing 1800 for each file using CLIP removed the soak data from all 5 casts.

##### Checking Headers

A cross-reference list was checked against the log book a few discrepancies were found:

- Two station names were entered wrong; they were fixed.

- One time (event 5) differed, but the log looks wrong and the header looks right. No change was made.

The cruise track was plotted and added to the end of this report.

Header Check was run and other than evidence of pressure spikes, no problems were found.

The Surface Check for the SBE911 data gave the average surface pressure as 1.6db which is within the normal range for the Vector. A few casts were found with low pressures at the end of the files. At the end of cast #105 the transmissivity and conductivity drop suddenly at pressure = -0.1db. At the end of cast #123 the transmissivity and conductivity have in-water values with the pumps on at pressure = +0.2db. These values suggest that the pressure is quite accurate, at least at the surface.

The Surface Check for the SBE25 was run on files before the soak data were clipped, and shows pressure and salinity values very close to 0, as expected.

The altimeter and bottom depth headers were exported from the SBE911 SAMAVG and CLN files to spreadsheets. The bottom depths were checked against the log book. There were many minor discrepancies but a few were large enough to warrant investigation. The altimetry header entries were checked by calculating:

Check Value = (Bottom depth – maximum depth of sampling + altimetry header)

The same calculation was done using the log entry for bottom depth to see if it produced better results.

The following cases were examined because either the Check Value ≥4m or bottom depths differed significantly from log entries:

- Casts 15, 35, 39, 67, 88, 95, 99, 111, 112, 115 and 116 –  Check Values ≥4m. Making water depth entry match log entries improved the Check Value so those changes were made.

- Casts 91, 107, 108 and 112 - No altimetry entry due to spike. Entered estimates based on plots.

- A few casts had Check value ≥4m but the altimetry and water depth entries seem ok – likely the depth changed during the casts, but no change was made to the entry in files.

The changes were made to the CLN files. Surprisingly, none of the affected casts were rosette casts so no changes were needed to the SAM files. Perhaps the crew had more time during rosette casts and entered the depth in the log a little later in the process when the CTD was in the water.

##### Shift

Fluorescence

- SHIFT was run on the SeaPoint fluorescence channel in the SBE911 casts using the usual advance of +24 records. Examination of plots after this step shows that the fluorescence offset is reasonably close to the temperature offset.

- The Eco fluorometer on the SBE25 does not usually require alignment since it is not pumped. A few profiles were examined and no alignment appears to be necessary, though the spikiness of the fluorescence makes it hard to judge.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and as usual, there is variability with the up and downcast traces sometimes closer than temperature and sometimes further apart. This is likely due to varying vertical gradients and descent/ascent speeds. Overall the choice made earlier looks appropriate, so no further alignment will be applied to either CTD DO sensor.

Conductivity

- For the SBE 911

For other recent cruises using these CTD sensors the best alignment setting for the primary conductivity sensor was always -0.6 records. For the secondary it was either -0.6 or -0.5 records. A test run on 3 casts showed that -0.6 worked well for the primary while -0.5 looks slightly better for the secondary.

SHIFT was run twice on all casts using -0.6 records for the primary and -0.5 for the secondary conductivity. Salinity was recalculated.

- For the SBE25

Tests were run on 2 casts and a shift of +1.3 records to the conductivity reduced noise in the salinity as seen on a T-S plot. SHIFT was run on all casts advancing the primary conductivity by 1.3 records. Salinity was recalculated.

pH

Tests were run on 2 casts to determine the best shift to make the difference between upcast and downcast pH profiles similar to that for temperature. A setting of +50 records produced the best results overall.

SHIFT was run to advance the pH channel by 50 records for all casts.

##### DELETE

- For the SBE 911

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: As expected the many pressure spikes resulted in many warnings, but all but 4 occur in the upcasts so are of no concern for the profiles. The others were investigated.

-Casts 12, 14 and 17 had some reversals around 150m – large jumps in pressure between records and then a skip back to expected values. Some are associated with off-scale pressures in the raw data. DELETE made the best of these malfunctions, choosing reasonable data and leaving only small gaps in pressure, ~0.4db.

-Cast 48 – There was a spike at 118db; DELETE performed properly and left only a small gap in pressure.

There was no warning for cast #115, but at the editing stage it was found that cast 115 was missing data above 7db due to bad pressure data. A return was made to the WILDEDIT stage to deal with bad pressure data that caused the problem. A second run of WILDEDIT removed the bad data.  All steps between WILDEDIT and DELETE were repeated. DELETE worked well after these steps.

- For the SBE25

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 11 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 9 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.0625 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

All sensors on the SBE911 were used during 6 earlier cruises in 2017.

- 2017-63 in early April was a cruise in the Strait of Georgia and Juan de Fuca Strait. The bottles from Juan de Fuca were trusted more than those from the Strait of Georgia because conditions were better for flushing Niskin bottles; they indicated that the primary salinity was close to bottles while the secondary was high by about 0.0045. The dissolved oxygen comparison from 2017-63 was also based on Juan de Fuca bottles.

- Cruise 2017-03 was in Baynes Sound and the Strait of Georgia, with shallow sampling and poor flushing of bottles, so the results were not considered reliable enough to use for recalibration.

- Cruise 2017-05 was mostly off the west coast of Vancouver Island where flushing was good at depth. The comparison between the bottles and CTD salinity was quite tight; the primary salinity was found to be very close to bottles and the secondary was high by 0.0054psu. The dissolved oxygen comparison was also quite tight and led to a calibration of slope 1.0448 and offset +0.0527.

- Cruise 2017-06 was mostly well offshore and showed the primary salinity to be low by ~0.0012 and the secondary high by ~0.0032. The dissolved oxygen was recalibrated using slope 1.0399 and offset 0.0645.

- Cruise 2017-64 showed primary and secondary salinity low by 0.0012psu and high by 0.0034psu in Juan de Fuca Strait. The dissolved oxygen comparison in Juan de Fuca was very close to that of 2017-06, so those results were used for 2017-64.

- Cruise 2017-56 included only near-shore shallow casts so 2017-06 results were used for recalibration.

- Cruise 2017-23 comparisons were not trusted as much as the 2017-06 results, so the latter were used for recalibration.

- Cruise 2017-36 had lots of bottles; DO comparison was very close to 2017-08. Primary salinity was low by ~0.0027 and secondary was high by 0.0033psu.

- Cruise 2017-08 was mostly well offshore and showed the primary salinity to be low by 0.0002 and the secondary high by 0.0004. The dissolved oxygen was recalibrated using slope 1.0514 and offset 0.0533. Fluorescence was recalibrated.

- Cruise 2017-09 included offshore and inland waters. The salinity channels were low by 0.0029 and 0.0006 though part of that was considered to be due to incomplete flushing. The dissolved oxygen was recalibrated using slope 1.0620 and offset 0.0606. Fluorescence was recalibrated.

All sensors on the SBE25 were used for a number of cruises during 2017, but oxygen calibration sampling was only available from 2017-01, an off-shore cast with good sampling. Salinity sampling was available from a number of small-boat cruises in the Strait of Georgia. There was a lot of scatter in the CTD/Bottle comparison but CTD salinity was thought to be high by about 0.0025psu of bottles, but the comparisons were not considered strong enough to justify recalibration.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All temperature data and most salinity data fell within the local climatology, though for a few casts no local climatology was available. Salinity was sometimes above the historic maximum and sometimes below so the excursions are not systematic and thus don’t suggest calibration problems. Most of the excursions from the climatology were in areas like the Gulf Islands and Discovery Islands that are likely not well represented in the climatology or possibly completely absent. Salinity was a little high at depth at stations 56, 47 and 48 in the Strait of Georgia. Such excursions are not unusual in this area which has very high variability. The climatology is old and a 3-standard-deviation climatology is not considered suitable for near-shore sampling.

Repeat Casts – There was one pair of casts at the same station and about 2 hours apart. From 1400db downward the differences along σt-lines are ≤ 0.002C° and ≤ 0.0003psu and usually much closer than that. This is excellent repeatability.

# Post-Cruise Calibration

None was available.

##### DETAILED EDITING

Before deciding which channel pair to edit and archive, a study was made of the salinity calibration.

- The bottle comparison shows the primary salinity to be high by ~0.002psu and the secondary to be high by ~0.01psu. Since the CTD was sometimes falling during casts and there was a lot of variability in salinity during stops, the CTD may be reading a little closer to bottles than the comparison suggests.

- The results for the primary sensor are a little different from other 2017 cruises using these sensors, while the secondary salinity results are very different.

- Pressure dependence in the fits could possibly explain some of the difference from results in previous cruises, so the 2017-08 and 2017-09 bottle comparisons were re-examined with only bottles included from between 10db and 400db and with CTD standard deviation <0.0008psu. There were very few such bottles for 2017-08. For 2017-09 the results were similar to the full comparison but that conclusion was based on few bottles.

- There was no obvious evidence of significant calibration drift from these sensors up until 2017-09. The differences between sensor pairs were consistent with the bottle comparison results. For 2017-08 the sensors were in better agreement than for the previous cruises but that was considered to be due to a lot of very deep sampling combined with pressure dependence of opposite signs in the conductivity channels so that they came closer together at depth. The secondary salinity showed more pressure dependence than the secondary during 2017-08. During 2017-09 the pressure dependence had opposite signs but the same magnitude.

- During 2017-09 the bottle comparison showed the primary salinity lower than bottles by about 0.0029 which is similar to 2017-23, 2017-36 and 2017-09. It was different from 2017-08 but 2017-08 had little near-surface sampling included in the fit so there would be less effect from inefficient flushing. The secondary salinity comparison was similar to previous results, but perhaps that is because the drift noted during 2017-65 had started during 2017-09, thus offsetting flushing errors.

- For this cruise the primary sensors are higher than bottles by 0.002, whereas they were lower or close to bottles in all previous cruises. If inefficient flushing were an issue we would expect them to be lower than bottles, though the frequent reversals in the salinity profiles do confuse that issue and an extremely steady descent rate for some casts could result in bottles fired above the bottom containing water from the downcast. There is some indication that having the CTD drop a metre or two during bottle stops, combined with high variability, may have led to some bottles containing fresher water than the ambient salinity at the CTD sensor level. Nonetheless, it looks like the salinity has drifted upwards somewhat.

- The secondary salinity is higher than bottles by about 0.01psu, much higher than in all previous cruises. That may be a little higher than the actual drift for the reasons given for the primary.

- So the primary sensors should be chosen for editing and archiving since the secondary is clearly drifting significantly. Recalibration is not recommended at this point since the salinity is likely within 0.002 and a factory post-cruise calibration may be available soon.

- Examination of traces showed less noise in the primary salinity than the secondary.

- There are too few data for a convincing conclusion except to say that the secondary data are likely reading significantly high. The primary salinity data are likely within 0.002psu.

The primary T and S channels were chosen for editing.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes. All files required some editing, but mostly it was removal of records from near the top and bottom of casts and light cleaning of salinity.

##### Initial Recalibration

For the SBE911 files:

- There pressure does not appear to need recalibration.

- The Salinity:T0:C0 channel appears to be reading too high by ~0.002psu but the evidence is weak, so no recalibration will be applied. If the post-cruise calibration shows clear evidence of drift the data can be recalibrated later.

- For dissolved oxygen the fit based on Juan de Fuca Strait is the best option.

File 2017-65-recal1.ccf was prepared to recalibrate dissolved oxygen using.

CTD DO Corrected = CTD DO * 1.0826 + 0.0447

CALIBRATE was run on the SAM and MRGCLN2 files using 2017-65-recal1.ccf.

COMPARE was rerun for dissolved oxygen using the recalibrated values. The results confirm that the recalibration was applied properly. See file 2017-65-DO-comp2.xlsx for details.

CALIBRATE was then run on the EDT files for the SBE911.

For the SBE25 files:

- Sampling from other cruises suggests the salinity is a little high, but the evidence is too weak to justify recalibration.

- The pressure appears to be ok.

- The only calibration information available for dissolved oxygen comes from cruise 2017-01 in February 2017. That was in open waters with good sampling so will be used for these data.

File 2017-65-sbe25-recal1.ccf was prepared to apply the following correction to the SBE25 EDT files:

CTD Dissolved Oxygen Corrected = 1.0484* CTD Dissolved Oxygen + 0.0413

##### Final Calibration of DO – SBE911 only

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed based on residuals the CTD DO was higher than the bottles by an average of ~0.02mL/L and standard deviation of 0.04mL/L. When the same points are displayed in a plot versus pressure, it is seen that the CTD is noisiest relative to bottles near the surface with more outliers on the high side than the low side, but it is not as one-sided as often seen which suggests that bottle flushing may be slightly better than usual despite the steady descent rate. By about 150m the differences are very small. The calibration quality looks good with the largest differences in regions of high DO vertical gradient.

##### Fluorescence Processing

A median filter, size 11, was applied to the SBE 911 fluorescence channel in the COR1 files.

A median filter, size 7, was applied to the SBE25 fluorescence channel in the COR1 files.

Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files for both CTDs (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some small unstable features, but those are from sites where such features are expected. There was one large spike that had been missed, so that file was edited further.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all SBE911 casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Transmissivity, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

REMOVE was run on all SBE25 casts to remove the following channels:

Scan_Number, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

There were separate runs for the two CTD systems.

SBE911

Data Processing Notes:

----------------------

Transmissivity, Fluorescence:URU:Seapoint, pH:SBE, PAR and PAR:Reference data are

nominal and unedited except that some records were removed in editing temperature

and salinity.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.6 mL/L from 0 to 20db

±0.3 mL/L from 20 to 100db

±0.1 mL/L from 100db to 150db

±0.03 mL/L below 150db

Warning: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

For details on the processing see document: 2017-65_Processing_Report.doc.

SBE25

Data Processing Notes:

----------------------

Most casts during this cruise were taken using an SBE911 CTD mounted in a rosette.

However, the data in this file were collected using an SBE25 CTD without

the use of a rosette.

No calibration sampling was available for the SBE25 for this cruise.

Recalibration of dissolved oxygen was based on sampling during cruise 2017-01 in

February 2017.

There was salinity calibration sampling during a series of cruises in the Strait of

Georgia in 2017 (2017-07, 2017-48, 2017-49). There was a lot of scatter in

the comparison but the CTD salinity appeared to be within 0.003psu of bottles.

No recalibration has been applied to the CTD salinity in this file.

Fluorescence:URU:Wetlabs data are nominal and unedited except that some records were

removed in editing temperature and salinity.

Warning: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples.

For details on the processing see document: 2017-65_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values ranged from 45% to 135% with lower values in Juan de Fuca and Haro Straits. In the central part of the Strait of Georgia saturation rates ranged from 90% to 110%.  There is no evidence of a DO sensor calibration problem, though with so large a range of values it does not indicate that the calibration is good either.

##### Final Bottle Files

CALIBRATE was run on all files and then the MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Transmissivity, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added for both the CTD DO and bottle DO, with mass units and REORDER was run to get the pairs of DO channels together.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

Plots were made of all casts to look for problems and no problems were found.

Cast 76 –A note was added to the header that the pumps were not on for bottle #3 so the CTD temperature, salinity, fluorescence and DO may not be as reliable as usual.

Standards check and a header check were run on all files. No errors were found.

The track plot looks ok.

Data from the CHE files were exported to a spreadsheet and compared with rosette sheets. No problems were found.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Cast #115 had to be recalibrated to correct an error in the transmissivity units caused by a problem in the update to IOS SHELL.

Particulars (mostly notes from log)

TSG – water pump unserviceable

SBE25+ deployment method – 10m dip, soak 2 min. Data from dip will need to be removed.

SBE911+ deployment method - 2 minute soak at 12m, acquisition started after return to surface.

3. SBE25

5. SBE25

7. SBE25

8. Con file changed – factor E in DO calibration

30-33. Descent rate may have been too fast.

33. Surface bottle wasn’t tripped

34. Spikes at 250db

37. Spikes ~150db. Spikes only in pressure.

43. New top O-ring was installed in bottle 1, bottle 2 was re-tensioned

65. Spikes ~150m

66. Spikes ~150m

76. Stopped winch at 410 to adjust spooling. This solved spiking problem.

77. Cast run just to get a missed bottle.

79. Keyboard bumped at ~150m so no 30s wait time for bottle 5.

82. Spigot O-rings changed on 15, 14, 12

88. Depth sounder bouncing from 323 to 349.

128. SBE25

130. SBE25

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |
| 2 | SEABIRD | 25 | 1123 | No | Yes |

| Calibration Information SBE911+ |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 22Feb2017 | Factory |  |  |
| Conductivity | 1766 | 30Mar2017 | Factory |  |  |
| Secondary Temp. | 4484 | 22Feb2017 | Factory |  |  |
| Secondary Cond. | 3531 | 30Mar2017 | Factory |  |  |
| Transmissometer | 1185DR | 9Aug2017 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 28Mar2017 | Factory |  |  |
| PAR | 4565 | 16Mar2011 | IOS |  |  |
| Surface PAR | 20518 | 21Mar2016 |  |  |  |
| pH | 0691 | 15Jan2015 | Factory |  |  |
| SeaPoint Fluor. | 3685 | 13Mar2017 | Factory |  |  |
| Pressure Sensor | 0550 | 24Feb2017 | Factory |  |  |
| Altimeter | 62355 |  |  |  |  |

| Calibration Information SBE25+ |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4888 | 7Nov2015 | Factory |  |  |
| Conductivity | 3396 | 22Dec2015 | Factory |  |  |
| ECO Fluorometer | 2214 | 02Jun2017 | ? |  |  |
| SBE43 Oxygen | 3234 | 11Nov2015 | Factory |  |  |
| Pressure | 1123 | 23May2014 | Factory |  |  |